## Main
# Channels= | 3 |  |  |  | COMPARE Version 3.6.1 | Run at 2023/03/23 13:26:15 | Compare Excel Add-in Version 2.2
| Primary | Secondary |  | Comparison Name
Ref | Pressure | Pressure |  | Pressure - Pressure
Comp 1 | Sal_Bot | Sal0 |  | Sal0 - Sal_Bot
Comp 2 | Sal_Bot | Sal1 |  | Sal1 - Sal_Bot
Average Size: | 1
|  |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary
Primary | Secondary | Secondary | Pressure | Pressure | Pressure | (S-P) | Sal_Bot | Sal0 | Sal0 | (S-P) | Sal_Bot | Sal1 | Sal1 | (S-P)
File Name | File Name | Count | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff
2022-075-0008.mrg | 2022-075-0008.sam | 361 | 140.164 | 140.104 | 0.40052 | -0.060318 | 33.6654 | 33.6583 | 0.0045588 | -0.0070915 | 33.6654 | 33.6651 | 0.0044282 | -0.00029755
2022-075-0008.mrg | 2022-075-0008.sam | 241 | 101.646 | 101.646 | 0.20795 | 0.00042725 | 32.9072 | 32.9244 | 0.003603 | 0.017223 | 32.9072 | 32.9293 | 0.0062695 | 0.022076
2022-075-0008.mrg | 2022-075-0008.sam | 241 | 50.414 | 50.414 | 0.25479 | 1.5259e-05 | 32.0346 | 32.0218 | 0.0076446 | -0.012764 | 32.0346 | 32.0313 | 0.0079106 | -0.0033417
2022-075-0008.mrg | 2022-075-0008.sam | 241 | 20.314 | 20.3142 | 0.14954 | 0.00024033 | 31.4048 | 31.4235 | 0.0054459 | 0.018726 | 31.4048 | 31.4278 | 0.0051295 | 0.023018
2022-075-0008.mrg | 2022-075-0008.sam | 241 | 9.94 | 9.9402 | 0.26963 | 0.00023651 | 31.2761 | 31.2196 | 0.015861 | -0.056549 | 31.2761 | 31.2281 | 0.012788 | -0.048019
2022-075-0008.mrg | 2022-075-0008.sam | 241 | 4.92 | 4.9201 | 0.19723 | 9.5367e-05 | 31.1213 | 31.088 | 0.023121 | -0.033329 | 31.1213 | 31.0847 | 0.019831 | -0.03665
2022-075-0008.mrg | 2022-075-0008.sam | 290 | 2.062 | 2.0698 | 0.11198 | 0.0078275 | 30.9714 | 30.9181 | 0.020837 | -0.053278 | 30.9714 | 30.9249 | 0.018099 | -0.046549
2022-075-0015.mrg | 2022-075-0015.sam | 323 | 165.741 | 165.742 | 0.023007 | 0.00071716 | 32.3197 | 32.312 | 0.00030883 | -0.00774 | 32.3197 | 32.3179 | 0.0002274 | -0.0018043
2022-075-0015.mrg | 2022-075-0015.sam | 241 | 99.758 | 99.7583 | 0.028394 | 0.00032806 | 32.3121 | 32.3022 | 0.00031241 | -0.0099487 | 32.3121 | 32.3079 | 0.00022213 | -0.0042076
2022-075-0015.mrg | 2022-075-0015.sam | 325 | 49.735 | 49.7348 | 0.017645 | -0.00022125 | 32.0882 | 32.0196 | 0.0011668 | -0.068642 | 32.0882 | 32.029 | 0.0038889 | -0.059185
2022-075-0015.mrg | 2022-075-0015.sam | 241 | 19.597 | 19.5967 | 0.022493 | -0.00033188 | 31.3044 | 31.1793 | 0.00091162 | -0.12506 | 31.3044 | 31.1847 | 0.0017084 | -0.11972
2022-075-0015.mrg | 2022-075-0015.sam | 241 | 10.129 | 10.1288 | 0.025977 | -0.00017834 | 30.6347 | 30.3544 | 0.085225 | -0.28028 | 30.6347 | 30.4127 | 0.028532 | -0.22195
2022-075-0015.mrg | 2022-075-0015.sam | 241 | 5.125 | 5.125 | 0.02347 | -2.0981e-05 | 29.8126 | 29.895 | 0.077612 | 0.082355 | 29.8126 | 29.8711 | 0.060408 | 0.058479
2022-075-0015.mrg | 2022-075-0015.sam | 1864 | 1.084 | 1.0524 | 0.03088 | -0.03159 | 23.198 | 25.2094 | 0.26061 | 2.0114 | 23.198 | 25.1706 | 0.27005 | 1.9726
2022-075-0022.mrg | 2022-075-0022.sam | 241 | 400.574 | 400.574 | 0.023737 | -0.00039673 | 33.6698 | 33.6612 | 0.00028395 | -0.0085907 | 33.6698 | 33.6662 | 0.00021143 | -0.003624
2022-075-0022.mrg | 2022-075-0022.sam | 241 | 300.472 | 300.472 | 0.026687 | -0.00039673 | 33.6576 | 33.6483 | 0.00029008 | -0.0093193 | 33.6576 | 33.6533 | 0.00022966 | -0.0043373
2022-075-0022.mrg | 2022-075-0022.sam | 241 | 200.315 | 200.315 | 0.026394 | 0.00021362 | 33.6102 | 33.5923 | 0.00042303 | -0.017887 | 33.6102 | 33.5973 | 0.00021376 | -0.012943
2022-075-0022.mrg | 2022-075-0022.sam | 241 | 100.742 | 100.742 | 0.024624 | -2.2888e-05 | 33.3148 | 33.2637 | 0.0012357 | -0.051113 | 33.3148 | 33.2696 | 0.0023026 | -0.045166
2022-075-0022.mrg | 2022-075-0022.sam | 293 | 50.459 | 50.4597 | 0.023018 | 0.00074387 | 32.2411 | 32.1227 | 0.0075451 | -0.11838 | 32.2411 | 32.1242 | 0.0067382 | -0.11686
2022-075-0022.mrg | 2022-075-0022.sam | 241 | 20.878 | 20.8778 | 0.020249 | -0.00021934 | 31.4477 | 31.3457 | 0.026631 | -0.10202 | 31.4477 | 31.366 | 0.032925 | -0.081732
2022-075-0022.mrg | 2022-075-0022.sam | 241 | 4.996 | 4.9961 | 0.026438 | 0.00014114 | 30.4995 | 30.2433 | 0.019079 | -0.25619 | 30.4995 | 30.2447 | 0.022111 | -0.25477
2022-075-0022.mrg | 2022-075-0022.sam | 374 | 0.607 | 0.60851 | 0.020449 | 0.001508 | 25.8269 | 27.67 | 0.3077 | 1.8431 | 25.8269 | 27.6062 | 0.22802 | 1.7793
2022-075-0029.mrg | 2022-075-0029.sam | 295 | 408.21 | 408.204 | 0.02982 | -0.006134 | 33.0496 | 33.043 | 0.00025733 | -0.0066147 | 33.0496 | 33.0486 | 0.00022454 | -0.0010109
2022-075-0029.mrg | 2022-075-0029.sam | 241 | 49.876 | 49.8762 | 0.023 | 0.00021362 | 32.4164 | 32.3624 | 0.0043851 | -0.054031 | 32.4164 | 32.3735 | 0.0048893 | -0.0429
2022-075-0029.mrg | 2022-075-0029.sam | 287 | 0.783 | 0.7822 | 0.024358 | -0.00080138 | 29.5229 | 29.3222 | 0.092595 | -0.20071 | 29.5229 | 29.4263 | 0.066218 | -0.096649
2022-075-0030.mrg | 2022-075-0030.sam | 320 | 614.27 | 614.318 | 0.047014 | 0.047485 | 33.0777 | 33.0676 | 0.00024755 | -0.010132 | 33.0777 | 33.0731 | 0.000223 | -0.0046425
2022-075-0030.mrg | 2022-075-0030.sam | 402 | 49.976 | 49.9757 | 0.026992 | -0.00030136 | 32.303 | 32.1782 | 0.0013148 | -0.12478 | 32.303 | 32.1869 | 0.0037931 | -0.11613
2022-075-0031.mrg | 2022-075-0031.sam | 326 | 763.48 | 763.484 | 0.030456 | 0.0039673 | 33.6723 | 33.6617 | 0.00028165 | -0.010635 | 33.6723 | 33.6671 | 0.00020928 | -0.0052376
2022-075-0031.mrg | 2022-075-0031.sam | 241 | 699.995 | 699.995 | 0.033726 | 0.00036621 | 33.6715 | 33.6616 | 0.00029148 | -0.0098534 | 33.6715 | 33.6671 | 0.0002098 | -0.0044289
2022-075-0031.mrg | 2022-075-0031.sam | 241 | 600.34 | 600.34 | 0.028178 | 0.00036621 | 33.6693 | 33.66 | 0.00027003 | -0.0092888 | 33.6693 | 33.6654 | 0.0002151 | -0.0038719
2022-075-0031.mrg | 2022-075-0031.sam | 241 | 499.977 | 499.977 | 0.051366 | 9.1553e-05 | 33.6655 | 33.6551 | 0.0002673 | -0.010429 | 33.6655 | 33.6604 | 0.00044776 | -0.0050812
2022-075-0031.mrg | 2022-075-0031.sam | 241 | 400.338 | 400.338 | 0.039425 | -0.00039673 | 33.6564 | 33.6468 | 0.00027851 | -0.0095749 | 33.6564 | 33.6522 | 0.00021008 | -0.0042381
2022-075-0031.mrg | 2022-075-0031.sam | 241 | 300.733 | 300.733 | 0.052801 | 0.00030518 | 33.6499 | 33.6401 | 0.00030403 | -0.009758 | 33.6499 | 33.6455 | 0.00022587 | -0.0044022
2022-075-0031.mrg | 2022-075-0031.sam | 241 | 200.943 | 200.943 | 0.036852 | 0.00035095 | 33.628 | 33.6192 | 0.00030047 | -0.008812 | 33.628 | 33.6245 | 0.0002199 | -0.0034828
2022-075-0031.mrg | 2022-075-0031.sam | 241 | 100.177 | 100.177 | 0.034315 | 0.00026703 | 33.1582 | 33.1735 | 0.014541 | 0.015305 | 33.1582 | 33.1768 | 0.011935 | 0.01865
2022-075-0031.mrg | 2022-075-0031.sam | 293 | 50.427 | 50.4141 | 0.037896 | -0.01289 | 32.1343 | 32.0349 | 0.00061061 | -0.099445 | 32.1343 | 32.0403 | 0.00048319 | -0.093956
2022-075-0031.mrg | 2022-075-0031.sam | 241 | 20.375 | 20.3752 | 0.031919 | 0.00021934 | 31.5039 | 31.4093 | 0.0036862 | -0.094595 | 31.5039 | 31.4149 | 0.0024565 | -0.088989
2022-075-0031.mrg | 2022-075-0031.sam | 241 | 10.284 | 10.2843 | 0.031046 | 0.0002861 | 31.2093 | 31.0757 | 0.00072816 | -0.13362 | 31.2093 | 31.0811 | 0.00095705 | -0.1282
2022-075-0031.mrg | 2022-075-0031.sam | 241 | 5.359 | 5.3594 | 0.034278 | 0.00041056 | 31.0721 | 31.0269 | 0.00049115 | -0.045202 | 31.0721 | 31.0318 | 0.0003757 | -0.040302
2022-075-0031.mrg | 2022-075-0031.sam | 326 | 0.473 | 0.48351 | 0.028801 | 0.010506 | 33.3466 | 31.0147 | 0.00031885 | -2.3319 | 33.3466 | 31.0201 | 0.00029634 | -2.3265
2022-075-0031.mrg | 2022-075-0031.sam | 326 | 0.456 | 0.48351 | 0.028801 | 0.027506 | 31.0264 | 31.0147 | 0.00031885 | -0.011656 | 31.0264 | 31.0201 | 0.00029634 | -0.0063209
2022-075-0031.mrg | 2022-075-0031.sam | 326 | 0.49 | 0.48351 | 0.028801 | -0.0064939 | 30.9388 | 31.0147 | 0.00031885 | 0.075943 | 30.9388 | 31.0201 | 0.00029634 | 0.081278
2022-075-0044.mrg | 2022-075-0044.sam | 345 | 517.803 | 517.809 | 0.024307 | 0.0059815 | 33.2737 | 33.2652 | 0.0002937 | -0.0085335 | 33.2737 | 33.2707 | 0.00026625 | -0.0029717
2022-075-0044.mrg | 2022-075-0044.sam | 293 | 200.578 | 200.577 | 0.02079 | -0.001358 | 33.2132 | 33.2012 | 0.00037732 | -0.011974 | 33.2132 | 33.2066 | 0.00033544 | -0.006588
2022-075-0044.mrg | 2022-075-0044.sam | 278 | 100.239 | 100.244 | 0.020831 | 0.0055008 | 33.0206 | 32.9854 | 0.00074611 | -0.035214 | 33.0206 | 32.9917 | 0.00081813 | -0.028942
2022-075-0044.mrg | 2022-075-0044.sam | 289 | 50.335 | 50.3362 | 0.025135 | 0.0012131 | 32.1431 | 32.1131 | 0.0095104 | -0.02998 | 32.1431 | 32.1226 | 0.010156 | -0.020523
2022-075-0044.mrg | 2022-075-0044.sam | 281 | 20.267 | 20.2675 | 0.026564 | 0.00051308 | 31.5687 | 31.4616 | 0.0056697 | -0.10709 | 31.5687 | 31.4409 | 0.0096277 | -0.1278
2022-075-0044.mrg | 2022-075-0044.sam | 308 | 9.926 | 9.9188 | 0.029318 | -0.0072107 | 31.1787 | 30.8264 | 0.020065 | -0.35234 | 31.1787 | 30.807 | 0.0089025 | -0.37171
2022-075-0044.mrg | 2022-075-0044.sam | 282 | 5.12 | 5.1223 | 0.021302 | 0.0022483 | 27.7246 | 29.0417 | 0.17793 | 1.3171 | 27.7246 | 29.089 | 0.15724 | 1.3644
2022-075-0044.mrg | 2022-075-0044.sam | 344 | 0.774 | 0.77146 | 0.021958 | -0.0025349 | 25.3444 | 24.6655 | 0.022731 | -0.67887 | 25.3444 | 24.8025 | 0.1441 | -0.54194
2022-075-0073.mrg | 2022-075-0073.sam | 383 | 365.859 | 365.861 | 0.022843 | 0.00177 | 32.9633 | 32.9539 | 0.00028597 | -0.0094109 | 32.9633 | 32.9594 | 0.00023941 | -0.0038948
2022-075-0073.mrg | 2022-075-0073.sam | 281 | 299.931 | 299.935 | 0.024233 | 0.004303 | 32.9581 | 32.9502 | 0.00029136 | -0.0078812 | 32.9581 | 32.9557 | 0.00020033 | -0.0023956
2022-075-0073.mrg | 2022-075-0073.sam | 282 | 200 | 199.991 | 0.020838 | -0.0094757 | 32.9208 | 32.9094 | 0.00081142 | -0.011402 | 32.9208 | 32.9143 | 0.0004072 | -0.0064735
2022-075-0073.mrg | 2022-075-0073.sam | 297 | 100.929 | 100.927 | 0.027104 | -0.0020828 | 32.7751 | 32.7632 | 0.0017323 | -0.011864 | 32.7751 | 32.7653 | 0.0030382 | -0.0098381
2022-075-0073.mrg | 2022-075-0073.sam | 304 | 50.33 | 50.3263 | 0.026904 | -0.0037498 | 32.1694 | 32.1093 | 0.0018971 | -0.060097 | 32.1694 | 32.1112 | 0.0067474 | -0.058174
2022-075-0073.mrg | 2022-075-0073.sam | 298 | 20.795 | 20.7948 | 0.020562 | -0.00022507 | 31.9417 | 31.8283 | 0.00039716 | -0.11335 | 31.9417 | 31.8343 | 0.00061997 | -0.10745
2022-075-0073.mrg | 2022-075-0073.sam | 308 | 10.265 | 10.2701 | 0.023464 | 0.0050974 | 31.7634 | 31.6885 | 0.0019151 | -0.074911 | 31.7634 | 31.6912 | 0.0029368 | -0.07218
2022-075-0073.mrg | 2022-075-0073.sam | 285 | 5.346 | 5.3449 | 0.023945 | -0.0011125 | 31.6451 | 31.3613 | 0.0035346 | -0.28382 | 31.6451 | 31.3717 | 0.0033182 | -0.27337
2022-075-0073.mrg | 2022-075-0073.sam | 377 | 1.044 | 1.0399 | 0.021843 | -0.0040796 | 28.2791 | 30.7979 | 0.23754 | 2.5188 | 28.2791 | 31.0294 | 0.24394 | 2.7503
2022-075-0084.mrg | 2022-075-0084.sam | 294 | 5.198 | 5.2115 | 0.024662 | 0.013463 | 23.2924 | 23.6972 | 0.10552 | 0.40481 | 23.2924 | 23.7651 | 0.10977 | 0.47266
2022-075-0084.mrg | 2022-075-0084.sam | 359 | 0.982 | 0.99503 | 0.027365 | 0.013031 | 16.3297 | 21.3267 | 0.4505 | 4.997 | 16.3297 | 21.2485 | 0.38476 | 4.9188
Create Fit Sheet
Plot
Plot
Plot

## File List
Number of Files= |  | 9
Pair # | Primary File | Secondary File | First Row in Main | Last Row in Main | Comparison
1 | 2022-075-0008.mrg | 2022-075-0008.sam | 12 | 18 | 2022-075-0008.mrg & 2022-075-0008.sam
2 | 2022-075-0015.mrg | 2022-075-0015.sam | 20 | 26 | 2022-075-0015.mrg & 2022-075-0015.sam
3 | 2022-075-0022.mrg | 2022-075-0022.sam | 28 | 35 | 2022-075-0022.mrg & 2022-075-0022.sam
4 | 2022-075-0029.mrg | 2022-075-0029.sam | 37 | 39 | 2022-075-0029.mrg & 2022-075-0029.sam
5 | 2022-075-0030.mrg | 2022-075-0030.sam | 41 | 42 | 2022-075-0030.mrg & 2022-075-0030.sam
6 | 2022-075-0031.mrg | 2022-075-0031.sam | 44 | 58 | 2022-075-0031.mrg & 2022-075-0031.sam
7 | 2022-075-0044.mrg | 2022-075-0044.sam | 60 | 67 | 2022-075-0044.mrg & 2022-075-0044.sam
8 | 2022-075-0073.mrg | 2022-075-0073.sam | 69 | 77 | 2022-075-0073.mrg & 2022-075-0073.sam
9 | 2022-075-0084.mrg | 2022-075-0084.sam | 79 | 80 | 2022-075-0084.mrg & 2022-075-0084.sam

## Fit Index
Index to Fit Worksheets
Fit# | Description | Versus | Number of File Pairs
1 | (Sal0 - Sal_Bot) | Pressure | 9
2 | (Sal1 - Sal_Bot) | Pressure | 9
3 | (Sal0 - Sal_Bot) | File Pair # | 9
4 | (Sal1 - Sal_Bot) | File Pair # | 9

## Fit_1
Calculation of Linear Least Squares Fit For
(Sal0 - Sal_Bot)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal0 - Sal_Bot), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.028139 |  | X | CALC Y
m= | 2.1704e-06 | 3.1889e-06 |  | Std Dev of Resid |  | 0.0024223 |  | 165.742 | -0.010381
b= | -0.010741 | 0.0013877 |  | Resid Sum of Squares |  | 9.388e-05 |  | 763.484 | -0.0090836
# of Values Extracted = |  |  | 61
# of Values Used in Fit = |  |  | 18
File | Use | Pressure | Sal0 | Sal0 | Diff (Sal0 - Sal_Bot) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 140.104 | 33.6583 | 0.0045588 | -0.0070915 | -0.010437 | -0.0033451 |  | 165.742 | -0.00774
1 | NO | 101.646 | 32.9244 | 0.003603 | 0.017223 | -0.01052 | -0.027744 |  | 400.574 | -0.0085907
1 | NO | 50.414 | 32.0218 | 0.0076446 | -0.012764 | -0.010631 | 0.0021327 |  | 300.472 | -0.0093193
1 | NO | 20.3142 | 31.4235 | 0.0054459 | 0.018726 | -0.010697 | -0.029423 |  | 200.315 | -0.017887
1 | NO | 9.9402 | 31.2196 | 0.015861 | -0.056549 | -0.010719 | 0.04583 |  | 408.204 | -0.0066147
1 | NO | 4.9201 | 31.088 | 0.023121 | -0.033329 | -0.01073 | 0.022599 |  | 614.318 | -0.010132
1 | NO | 2.0698 | 30.9181 | 0.020837 | -0.053278 | -0.010736 | 0.042542 |  | 763.484 | -0.010635
2 | YES | 165.742 | 32.312 | 0.00030883 | -0.00774 | -0.010381 | -0.002641 |  | 699.995 | -0.0098534
2 | NO | 99.7583 | 32.3022 | 0.00031241 | -0.0099487 | -0.010524 | -0.00057549 |  | 600.34 | -0.0092888
2 | NO | 49.7348 | 32.0196 | 0.0011668 | -0.068642 | -0.010633 | 0.058009 |  | 499.977 | -0.010429
2 | NO | 19.5967 | 31.1793 | 0.00091162 | -0.12506 | -0.010698 | 0.11436 |  | 400.338 | -0.0095749
2 | NO | 10.1288 | 30.3544 | 0.085225 | -0.28028 | -0.010719 | 0.26956 |  | 300.733 | -0.009758
2 | NO | 5.125 | 29.895 | 0.077612 | 0.082355 | -0.01073 | -0.093085 |  | 200.943 | -0.008812
2 | NO | 1.0524 | 25.2094 | 0.26061 | 2.0114 | -0.010738 | -2.0222 |  | 517.809 | -0.0085335
3 | YES | 400.574 | 33.6612 | 0.00028395 | -0.0085907 | -0.0098713 | -0.0012806 |  | 200.577 | -0.011974
3 | YES | 300.472 | 33.6483 | 0.00029008 | -0.0093193 | -0.010089 | -0.00076928 |  | 365.861 | -0.0094109
3 | YES | 200.315 | 33.5923 | 0.00042303 | -0.017887 | -0.010306 | 0.0075811 |  | 299.935 | -0.0078812
3 | NO | 100.742 | 33.2637 | 0.0012357 | -0.051113 | -0.010522 | 0.040591 |  | 199.991 | -0.011402
3 | NO | 50.4597 | 32.1227 | 0.0075451 | -0.11838 | -0.010631 | 0.10775
3 | NO | 20.8778 | 31.3457 | 0.026631 | -0.10202 | -0.010695 | 0.091325
3 | NO | 4.9961 | 30.2433 | 0.019079 | -0.25619 | -0.01073 | 0.24546
3 | NO | 0.60851 | 27.67 | 0.3077 | 1.8431 | -0.010739 | -1.8538
4 | YES | 408.204 | 33.043 | 0.00025733 | -0.0066147 | -0.0098548 | -0.0032401
4 | NO | 49.8762 | 32.3624 | 0.0043851 | -0.054031 | -0.010632 | 0.043399
4 | NO | 0.7822 | 29.3222 | 0.092595 | -0.20071 | -0.010739 | 0.18997
5 | YES | 614.318 | 33.0676 | 0.00024755 | -0.010132 | -0.0094074 | 0.0007244
5 | NO | 49.9757 | 32.1782 | 0.0013148 | -0.12478 | -0.010632 | 0.11415
6 | YES | 763.484 | 33.6617 | 0.00028165 | -0.010635 | -0.0090836 | 0.0015518
6 | YES | 699.995 | 33.6616 | 0.00029148 | -0.0098534 | -0.0092214 | 0.00063191
6 | YES | 600.34 | 33.66 | 0.00027003 | -0.0092888 | -0.0094377 | -0.00014895
6 | YES | 499.977 | 33.6551 | 0.0002673 | -0.010429 | -0.0096556 | 0.00077383
6 | YES | 400.338 | 33.6468 | 0.00027851 | -0.0095749 | -0.0098718 | -0.00029694
6 | YES | 300.733 | 33.6401 | 0.00030403 | -0.009758 | -0.010088 | -0.00033002
6 | YES | 200.943 | 33.6192 | 0.00030047 | -0.008812 | -0.010305 | -0.0014927
6 | NO | 100.177 | 33.1735 | 0.014541 | 0.015305 | -0.010523 | -0.025828
6 | NO | 50.4141 | 32.0349 | 0.00061061 | -0.099445 | -0.010631 | 0.088814
6 | NO | 20.3752 | 31.4093 | 0.0036862 | -0.094595 | -0.010697 | 0.083898
6 | NO | 10.2843 | 31.0757 | 0.00072816 | -0.13362 | -0.010718 | 0.1229
6 | NO | 5.3594 | 31.0269 | 0.00049115 | -0.045202 | -0.010729 | 0.034473
6 | NO | 0.48351 | 31.0147 | 0.00031885 | -2.3319 | -0.01074 | 2.3211
6 | NO | 0.48351 | 31.0147 | 0.00031885 | -0.011656 | -0.01074 | 0.00091611
6 | NO | 0.48351 | 31.0147 | 0.00031885 | 0.075943 | -0.01074 | -0.086683 |  |  |  | -0.0098799 | 0.0023837
7 | YES | 517.809 | 33.2652 | 0.0002937 | -0.0085335 | -0.0096169 | -0.0010834
7 | YES | 200.577 | 33.2012 | 0.00037732 | -0.011974 | -0.010305 | 0.0016689
7 | NO | 100.244 | 32.9854 | 0.00074611 | -0.035214 | -0.010523 | 0.02469
7 | NO | 50.3362 | 32.1131 | 0.0095104 | -0.02998 | -0.010631 | 0.019348
7 | NO | 20.2675 | 31.4616 | 0.0056697 | -0.10709 | -0.010697 | 0.096391
7 | NO | 9.9188 | 30.8264 | 0.020065 | -0.35234 | -0.010719 | 0.34162
7 | NO | 5.1223 | 29.0417 | 0.17793 | 1.3171 | -0.01073 | -1.3278
7 | NO | 0.77146 | 24.6655 | 0.022731 | -0.67887 | -0.010739 | 0.66813
8 | YES | 365.861 | 32.9539 | 0.00028597 | -0.0094109 | -0.0099467 | -0.00053581
8 | YES | 299.935 | 32.9502 | 0.00029136 | -0.0078812 | -0.01009 | -0.0022086
8 | YES | 199.991 | 32.9094 | 0.00081142 | -0.011402 | -0.010307 | 0.0010954
8 | NO | 100.927 | 32.7632 | 0.0017323 | -0.011864 | -0.010522 | 0.001342
8 | NO | 50.3263 | 32.1093 | 0.0018971 | -0.060097 | -0.010632 | 0.049465
8 | NO | 20.7948 | 31.8283 | 0.00039716 | -0.11335 | -0.010696 | 0.10266
8 | NO | 10.2701 | 31.6885 | 0.0019151 | -0.074911 | -0.010718 | 0.064193
8 | NO | 5.3449 | 31.3613 | 0.0035346 | -0.28382 | -0.010729 | 0.27309
8 | NO | 1.0399 | 30.7979 | 0.23754 | 2.5188 | -0.010738 | -2.5295
9 | NO | 5.2115 | 23.6972 | 0.10552 | 0.40481 | -0.010729 | -0.41554
9 | NO | 0.99503 | 21.3267 | 0.4505 | 4.997 | -0.010739 | -5.0078
[11 empty] | count-total |  | 61
[11 empty] | count in fit |  | 18
NO
NO
NO
NO
NO
NO
NO
YES
NO
NO
NO
NO
NO
NO
YES
YES
YES
NO
NO
NO
NO
NO
YES
NO
NO
YES
NO
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
NO
NO
NO
NO
NO
NO
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
FIT vs. Pressure
FIT vs. Sal0
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_2
Calculation of Linear Least Squares Fit For
(Sal1 - Sal_Bot)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal1 - Sal_Bot), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.031265 |  | X | CALC Y
m= | 2.4337e-06 | 3.3868e-06 |  | Std Dev of Resid |  | 0.0025726 |  | 165.742 | -0.0050858
b= | -0.0054891 | 0.0014739 |  | Resid Sum of Squares |  | 0.0001059 |  | 763.484 | -0.003631
# of Values Extracted = |  |  | 61
# of Values Used in Fit = |  |  | 18
File | Use | Pressure | Sal1 | Sal1 | Diff (Sal1 - Sal_Bot) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 140.104 | 33.6651 | 0.0044282 | -0.00029755 | -0.0051482 | -0.0048506 |  | 165.742 | -0.0018043
1 | NO | 101.646 | 32.9293 | 0.0062695 | 0.022076 | -0.0052418 | -0.027317 |  | 400.574 | -0.003624
1 | NO | 50.414 | 32.0313 | 0.0079106 | -0.0033417 | -0.0053664 | -0.0020248 |  | 300.472 | -0.0043373
1 | NO | 20.3142 | 31.4278 | 0.0051295 | 0.023018 | -0.0054397 | -0.028458 |  | 200.315 | -0.012943
1 | NO | 9.9402 | 31.2281 | 0.012788 | -0.048019 | -0.0054649 | 0.042554 |  | 408.204 | -0.0010109
1 | NO | 4.9201 | 31.0847 | 0.019831 | -0.03665 | -0.0054772 | 0.031173 |  | 614.318 | -0.0046425
1 | NO | 2.0698 | 30.9249 | 0.018099 | -0.046549 | -0.0054841 | 0.041065 |  | 763.484 | -0.0052376
2 | YES | 165.742 | 32.3179 | 0.0002274 | -0.0018043 | -0.0050858 | -0.0032814 |  | 699.995 | -0.0044289
2 | NO | 99.7583 | 32.3079 | 0.00022213 | -0.0042076 | -0.0052463 | -0.0010387 |  | 600.34 | -0.0038719
2 | NO | 49.7348 | 32.029 | 0.0038889 | -0.059185 | -0.0053681 | 0.053817 |  | 499.977 | -0.0050812
2 | NO | 19.5967 | 31.1847 | 0.0017084 | -0.11972 | -0.0054414 | 0.11428 |  | 400.338 | -0.0042381
2 | NO | 10.1288 | 30.4127 | 0.028532 | -0.22195 | -0.0054645 | 0.21649 |  | 300.733 | -0.0044022
2 | NO | 5.125 | 29.8711 | 0.060408 | 0.058479 | -0.0054767 | -0.063956 |  | 200.943 | -0.0034828
2 | NO | 1.0524 | 25.1706 | 0.27005 | 1.9726 | -0.0054866 | -1.9781 |  | 517.809 | -0.0029717
3 | YES | 400.574 | 33.6662 | 0.00021143 | -0.003624 | -0.0045142 | -0.00089029 |  | 200.577 | -0.006588
3 | YES | 300.472 | 33.6533 | 0.00022966 | -0.0043373 | -0.0047579 | -0.00042056 |  | 365.861 | -0.0038948
3 | YES | 200.315 | 33.5973 | 0.00021376 | -0.012943 | -0.0050016 | 0.0079417 |  | 299.935 | -0.0023956
3 | NO | 100.742 | 33.2696 | 0.0023026 | -0.045166 | -0.005244 | 0.039922 |  | 199.991 | -0.0064735
3 | NO | 50.4597 | 32.1242 | 0.0067382 | -0.11686 | -0.0053663 | 0.11149 |  |  |  | -0.0045238 | 0.0025358
3 | NO | 20.8778 | 31.366 | 0.032925 | -0.081732 | -0.0054383 | 0.076293
3 | NO | 4.9961 | 30.2447 | 0.022111 | -0.25477 | -0.005477 | 0.24929
3 | NO | 0.60851 | 27.6062 | 0.22802 | 1.7793 | -0.0054877 | -1.7847
4 | YES | 408.204 | 33.0486 | 0.00022454 | -0.0010109 | -0.0044957 | -0.0034848
4 | NO | 49.8762 | 32.3735 | 0.0048893 | -0.0429 | -0.0053677 | 0.037532
4 | NO | 0.7822 | 29.4263 | 0.066218 | -0.096649 | -0.0054872 | 0.091162
5 | YES | 614.318 | 33.0731 | 0.000223 | -0.0046425 | -0.0039941 | 0.00064844
5 | NO | 49.9757 | 32.1869 | 0.0037931 | -0.11613 | -0.0053675 | 0.11076
6 | YES | 763.484 | 33.6671 | 0.00020928 | -0.0052376 | -0.003631 | 0.0016066
6 | YES | 699.995 | 33.6671 | 0.0002098 | -0.0044289 | -0.0037855 | 0.00064332
6 | YES | 600.34 | 33.6654 | 0.0002151 | -0.0038719 | -0.0040281 | -0.00015615
6 | YES | 499.977 | 33.6604 | 0.00044776 | -0.0050812 | -0.0042723 | 0.00080885
6 | YES | 400.338 | 33.6522 | 0.00021008 | -0.0042381 | -0.0045148 | -0.00027669
6 | YES | 300.733 | 33.6455 | 0.00022587 | -0.0044022 | -0.0047572 | -0.00035507
6 | YES | 200.943 | 33.6245 | 0.0002199 | -0.0034828 | -0.0050001 | -0.0015173
6 | NO | 100.177 | 33.1768 | 0.011935 | 0.01865 | -0.0052453 | -0.023895
6 | NO | 50.4141 | 32.0403 | 0.00048319 | -0.093956 | -0.0053664 | 0.08859
6 | NO | 20.3752 | 31.4149 | 0.0024565 | -0.088989 | -0.0054395 | 0.08355
6 | NO | 10.2843 | 31.0811 | 0.00095705 | -0.1282 | -0.0054641 | 0.12273
6 | NO | 5.3594 | 31.0318 | 0.0003757 | -0.040302 | -0.0054761 | 0.034826
6 | NO | 0.48351 | 31.0201 | 0.00029634 | -2.3265 | -0.005488 | 2.321
6 | NO | 0.48351 | 31.0201 | 0.00029634 | -0.0063209 | -0.005488 | 0.00083299
6 | NO | 0.48351 | 31.0201 | 0.00029634 | 0.081278 | -0.005488 | -0.086766
7 | YES | 517.809 | 33.2707 | 0.00026625 | -0.0029717 | -0.0042289 | -0.0012573
7 | YES | 200.577 | 33.2066 | 0.00033544 | -0.006588 | -0.005001 | 0.001587
7 | NO | 100.244 | 32.9917 | 0.00081813 | -0.028942 | -0.0052452 | 0.023697
7 | NO | 50.3362 | 32.1226 | 0.010156 | -0.020523 | -0.0053666 | 0.015156
7 | NO | 20.2675 | 31.4409 | 0.0096277 | -0.1278 | -0.0054398 | 0.12236
7 | NO | 9.9188 | 30.807 | 0.0089025 | -0.37171 | -0.005465 | 0.36625
7 | NO | 5.1223 | 29.089 | 0.15724 | 1.3644 | -0.0054767 | -1.3698
7 | NO | 0.77146 | 24.8025 | 0.1441 | -0.54194 | -0.0054873 | 0.53645
8 | YES | 365.861 | 32.9594 | 0.00023941 | -0.0038948 | -0.0045987 | -0.00070392
8 | YES | 299.935 | 32.9557 | 0.00020033 | -0.0023956 | -0.0047592 | -0.0023635
8 | YES | 199.991 | 32.9143 | 0.0004072 | -0.0064735 | -0.0050024 | 0.0014711
8 | NO | 100.927 | 32.7653 | 0.0030382 | -0.0098381 | -0.0052435 | 0.0045946
8 | NO | 50.3263 | 32.1112 | 0.0067474 | -0.058174 | -0.0053667 | 0.052807
8 | NO | 20.7948 | 31.8343 | 0.00061997 | -0.10745 | -0.0054385 | 0.10201
8 | NO | 10.2701 | 31.6912 | 0.0029368 | -0.07218 | -0.0054641 | 0.066716
8 | NO | 5.3449 | 31.3717 | 0.0033182 | -0.27337 | -0.0054761 | 0.26789
8 | NO | 1.0399 | 31.0294 | 0.24394 | 2.7503 | -0.0054866 | -2.7558
9 | NO | 5.2115 | 23.7651 | 0.10977 | 0.47266 | -0.0054764 | -0.47814
9 | NO | 0.99503 | 21.2485 | 0.38476 | 4.9188 | -0.0054867 | -4.9243
NO
NO
NO
NO
NO
NO
NO
YES
NO
NO
NO
NO
NO
NO
YES
YES
YES
NO
NO
NO
NO
NO
YES
NO
NO
YES
NO
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
NO
NO
NO
NO
NO
NO
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
FIT vs. Pressure
FIT vs. Sal1
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_3
Calculation of Linear Least Squares Fit For
(Sal0 - Sal_Bot)
3
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal0 - Sal_Bot), Prop=File Pair #
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.0040433 |  | X | CALC Y
m= | 8.1749e-05 | 0.00032076 |  | Std Dev of Resid |  | 0.0024521 |  | 2 | -0.010171
b= | -0.010334 | 0.0018734 |  | Resid Sum of Squares |  | 9.6208e-05 |  | 8 | -0.00968
# of Values Extracted = |  |  | 61
# of Values Used in Fit = |  |  | 18
File | Use | Pressure | Sal0 | Sal0 | Diff (Sal0 - Sal_Bot) |  |  |  | Fit vs. | File Pair #
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 140.104 | 33.6583 | 0.0045588 | -0.0070915 | -0.010252 | -0.0031607 |  | 2 | -0.00774
1 | NO | 101.646 | 32.9244 | 0.003603 | 0.017223 | -0.010252 | -0.027476 |  | 3 | -0.0085907
1 | NO | 50.414 | 32.0218 | 0.0076446 | -0.012764 | -0.010252 | 0.0025117 |  | 3 | -0.0093193
1 | NO | 20.3142 | 31.4235 | 0.0054459 | 0.018726 | -0.010252 | -0.028979 |  | 3 | -0.017887 |  | Bottles below 140db and std dev in CTD SAL < 0.001psu
1 | NO | 9.9402 | 31.2196 | 0.015861 | -0.056549 | -0.010252 | 0.046297 |  | 4 | -0.0066147 |  | without bottom bottles | -0.0102
1 | NO | 4.9201 | 31.088 | 0.023121 | -0.033329 | -0.010252 | 0.023077 |  | 5 | -0.010132 |  | all bottles | -0.0099
1 | NO | 2.0698 | 30.9181 | 0.020837 | -0.053278 | -0.010252 | 0.043026 |  | 6 | -0.010635 |  | only bottom bottles | -0.0088
2 | YES | 165.742 | 32.312 | 0.00030883 | -0.00774 | -0.010171 | -0.0024305 |  | 6 | -0.0098534
2 | NO | 99.7583 | 32.3022 | 0.00031241 | -0.0099487 | -0.010171 | -0.00022178 |  | 6 | -0.0092888
2 | NO | 49.7348 | 32.0196 | 0.0011668 | -0.068642 | -0.010171 | 0.058471 |  | 6 | -0.010429
2 | NO | 19.5967 | 31.1793 | 0.00091162 | -0.12506 | -0.010171 | 0.11489 |  | 6 | -0.0095749
2 | NO | 10.1288 | 30.3544 | 0.085225 | -0.28028 | -0.010171 | 0.27011 |  | 6 | -0.009758
2 | NO | 5.125 | 29.895 | 0.077612 | 0.082355 | -0.010171 | -0.092526 |  | 6 | -0.008812
2 | NO | 1.0524 | 25.2094 | 0.26061 | 2.0114 | -0.010171 | -2.0216 |  | 7 | -0.0085335
3 | YES | 400.574 | 33.6612 | 0.00028395 | -0.0085907 | -0.010089 | -0.0014981 |  | 7 | -0.011974
3 | YES | 300.472 | 33.6483 | 0.00029008 | -0.0093193 | -0.010089 | -0.00076945 |  | 8 | -0.0094109
3 | YES | 200.315 | 33.5923 | 0.00042303 | -0.017887 | -0.010089 | 0.0077983 |  | 8 | -0.0078812
3 | NO | 100.742 | 33.2637 | 0.0012357 | -0.051113 | -0.010089 | 0.041024 |  | 8 | -0.011402 | -0.0098799
3 | NO | 50.4597 | 32.1227 | 0.0075451 | -0.11838 | -0.010089 | 0.10829
3 | NO | 20.8778 | 31.3457 | 0.026631 | -0.10202 | -0.010089 | 0.091931
3 | NO | 4.9961 | 30.2433 | 0.019079 | -0.25619 | -0.010089 | 0.2461
3 | NO | 0.60851 | 27.67 | 0.3077 | 1.8431 | -0.010089 | -1.8532
4 | YES | 408.204 | 33.043 | 0.00025733 | -0.0066147 | -0.010007 | -0.0033923
4 | NO | 49.8762 | 32.3624 | 0.0043851 | -0.054031 | -0.010007 | 0.044024
4 | NO | 0.7822 | 29.3222 | 0.092595 | -0.20071 | -0.010007 | 0.1907
5 | YES | 614.318 | 33.0676 | 0.00024755 | -0.010132 | -0.0099253 | 0.00020653
5 | NO | 49.9757 | 32.1782 | 0.0013148 | -0.12478 | -0.0099253 | 0.11485
6 | YES | 763.484 | 33.6617 | 0.00028165 | -0.010635 | -0.0098435 | 0.00079188
6 | YES | 699.995 | 33.6616 | 0.00029148 | -0.0098534 | -0.0098435 | 9.8423e-06
6 | YES | 600.34 | 33.66 | 0.00027003 | -0.0092888 | -0.0098435 | -0.00055473
6 | YES | 499.977 | 33.6551 | 0.0002673 | -0.010429 | -0.0098435 | 0.00058588
6 | YES | 400.338 | 33.6468 | 0.00027851 | -0.0095749 | -0.0098435 | -0.00026863
6 | YES | 300.733 | 33.6401 | 0.00030403 | -0.009758 | -0.0098435 | -8.5518e-05
6 | YES | 200.943 | 33.6192 | 0.00030047 | -0.008812 | -0.0098435 | -0.0010316
6 | NO | 100.177 | 33.1735 | 0.014541 | 0.015305 | -0.0098435 | -0.025148
6 | NO | 50.4141 | 32.0349 | 0.00061061 | -0.099445 | -0.0098435 | 0.089602
6 | NO | 20.3752 | 31.4093 | 0.0036862 | -0.094595 | -0.0098435 | 0.084751
6 | NO | 10.2843 | 31.0757 | 0.00072816 | -0.13362 | -0.0098435 | 0.12378
6 | NO | 5.3594 | 31.0269 | 0.00049115 | -0.045202 | -0.0098435 | 0.035359
6 | NO | 0.48351 | 31.0147 | 0.00031885 | -2.3319 | -0.0098435 | 2.322
6 | NO | 0.48351 | 31.0147 | 0.00031885 | -0.011656 | -0.0098435 | 0.0018123
6 | NO | 0.48351 | 31.0147 | 0.00031885 | 0.075943 | -0.0098435 | -0.085787
7 | YES | 517.809 | 33.2652 | 0.0002937 | -0.0085335 | -0.0097618 | -0.0012283
7 | YES | 200.577 | 33.2012 | 0.00037732 | -0.011974 | -0.0097618 | 0.0022125
7 | NO | 100.244 | 32.9854 | 0.00074611 | -0.035214 | -0.0097618 | 0.025452
7 | NO | 50.3362 | 32.1131 | 0.0095104 | -0.02998 | -0.0097618 | 0.020218
7 | NO | 20.2675 | 31.4616 | 0.0056697 | -0.10709 | -0.0097618 | 0.097326
7 | NO | 9.9188 | 30.8264 | 0.020065 | -0.35234 | -0.0097618 | 0.34258
7 | NO | 5.1223 | 29.0417 | 0.17793 | 1.3171 | -0.0097618 | -1.3269
7 | NO | 0.77146 | 24.6655 | 0.022731 | -0.67887 | -0.0097618 | 0.66911
8 | YES | 365.861 | 32.9539 | 0.00028597 | -0.0094109 | -0.00968 | -0.00026916
8 | YES | 299.935 | 32.9502 | 0.00029136 | -0.0078812 | -0.00968 | -0.0017989
8 | YES | 199.991 | 32.9094 | 0.00081142 | -0.011402 | -0.00968 | 0.0017221
8 | NO | 100.927 | 32.7632 | 0.0017323 | -0.011864 | -0.00968 | 0.0021837
8 | NO | 50.3263 | 32.1093 | 0.0018971 | -0.060097 | -0.00968 | 0.050417
8 | NO | 20.7948 | 31.8283 | 0.00039716 | -0.11335 | -0.00968 | 0.10367
8 | NO | 10.2701 | 31.6885 | 0.0019151 | -0.074911 | -0.00968 | 0.065231
8 | NO | 5.3449 | 31.3613 | 0.0035346 | -0.28382 | -0.00968 | 0.27414
8 | NO | 1.0399 | 30.7979 | 0.23754 | 2.5188 | -0.00968 | -2.5285
9 | NO | 5.2115 | 23.6972 | 0.10552 | 0.40481 | -0.0095983 | -0.41441
9 | NO | 0.99503 | 21.3267 | 0.4505 | 4.997 | -0.0095983 | -5.0066
NO
NO
NO
NO
NO
NO
NO
YES
NO
NO
NO
NO
NO
NO
YES
YES
YES
NO
NO
NO
NO
NO
YES
NO
NO
YES
NO
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
NO
NO
NO
NO
NO
NO
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
FIT vs. Pressure
FIT vs. Sal0
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_4
Calculation of Linear Least Squares Fit For
(Sal1 - Sal_Bot)
3
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal1 - Sal_Bot), Prop=File Pair #
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.0044256 |  | X | CALC Y
m= | 9.0981e-05 | 0.00034115 |  | Std Dev of Resid |  | 0.002608 |  | 2 | -0.0048473
b= | -0.0050293 | 0.0019925 |  | Resid Sum of Squares |  | 0.00010883 |  | 8 | -0.0043014
# of Values Extracted = |  |  | 61
# of Values Used in Fit = |  |  | 18
File | Use | Pressure | Sal1 | Sal1 | Diff (Sal1 - Sal_Bot) |  |  |  | Fit vs. | File Pair #
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 140.104 | 33.6651 | 0.0044282 | -0.00029755 | -0.0049383 | -0.0046407 |  | 2 | -0.0018043
1 | NO | 101.646 | 32.9293 | 0.0062695 | 0.022076 | -0.0049383 | -0.027014 |  | 3 | -0.003624
1 | NO | 50.414 | 32.0313 | 0.0079106 | -0.0033417 | -0.0049383 | -0.0015966 |  | 3 | -0.0043373
1 | NO | 20.3142 | 31.4278 | 0.0051295 | 0.023018 | -0.0049383 | -0.027956 |  | 3 | -0.012943
1 | NO | 9.9402 | 31.2281 | 0.012788 | -0.048019 | -0.0049383 | 0.043081 |  | 4 | -0.0010109
1 | NO | 4.9201 | 31.0847 | 0.019831 | -0.03665 | -0.0049383 | 0.031711 |  | 5 | -0.0046425
1 | NO | 2.0698 | 30.9249 | 0.018099 | -0.046549 | -0.0049383 | 0.041611 |  | 6 | -0.0052376
2 | YES | 165.742 | 32.3179 | 0.0002274 | -0.0018043 | -0.0048473 | -0.0030429 |  | 6 | -0.0044289
2 | NO | 99.7583 | 32.3079 | 0.00022213 | -0.0042076 | -0.0048473 | -0.00063969 |  | 6 | -0.0038719
2 | NO | 49.7348 | 32.029 | 0.0038889 | -0.059185 | -0.0048473 | 0.054338 |  | 6 | -0.0050812
2 | NO | 19.5967 | 31.1847 | 0.0017084 | -0.11972 | -0.0048473 | 0.11487 |  | 6 | -0.0042381 | -0.0045238
2 | NO | 10.1288 | 30.4127 | 0.028532 | -0.22195 | -0.0048473 | 0.21711 |  | 6 | -0.0044022
2 | NO | 5.125 | 29.8711 | 0.060408 | 0.058479 | -0.0048473 | -0.063327 |  | 6 | -0.0034828
2 | NO | 1.0524 | 25.1706 | 0.27005 | 1.9726 | -0.0048473 | -1.9775 |  | 7 | -0.0029717
3 | YES | 400.574 | 33.6662 | 0.00021143 | -0.003624 | -0.0047563 | -0.0011324 |  | 7 | -0.006588
3 | YES | 300.472 | 33.6533 | 0.00022966 | -0.0043373 | -0.0047563 | -0.00041901 |  | 8 | -0.0038948 |  | Bottles below 140db and std dev in CTD SAL < 0.001psu
3 | YES | 200.315 | 33.5973 | 0.00021376 | -0.012943 | -0.0047563 | 0.008187 |  | 8 | -0.0023956 |  | without bottom bottles | -0.0053
3 | NO | 100.742 | 33.2696 | 0.0023026 | -0.045166 | -0.0047563 | 0.04041 |  | 8 | -0.0064735 |  | all bottles | -0.0045
3 | NO | 50.4597 | 32.1242 | 0.0067382 | -0.11686 | -0.0047563 | 0.1121 |  |  |  |  | only bottom bottles | -0.0033
3 | NO | 20.8778 | 31.366 | 0.032925 | -0.081732 | -0.0047563 | 0.076975
3 | NO | 4.9961 | 30.2447 | 0.022111 | -0.25477 | -0.0047563 | 0.25001
3 | NO | 0.60851 | 27.6062 | 0.22802 | 1.7793 | -0.0047563 | -1.784
4 | YES | 408.204 | 33.0486 | 0.00022454 | -0.0010109 | -0.0046653 | -0.0036544
4 | NO | 49.8762 | 32.3735 | 0.0048893 | -0.0429 | -0.0046653 | 0.038235
4 | NO | 0.7822 | 29.4263 | 0.066218 | -0.096649 | -0.0046653 | 0.091984
5 | YES | 614.318 | 33.0731 | 0.000223 | -0.0046425 | -0.0045744 | 6.8136e-05
5 | NO | 49.9757 | 32.1869 | 0.0037931 | -0.11613 | -0.0045744 | 0.11156
6 | YES | 763.484 | 33.6671 | 0.00020928 | -0.0052376 | -0.0044834 | 0.00075421
6 | YES | 699.995 | 33.6671 | 0.0002098 | -0.0044289 | -0.0044834 | -5.4513e-05
6 | YES | 600.34 | 33.6654 | 0.0002151 | -0.0038719 | -0.0044834 | -0.00061145
6 | YES | 499.977 | 33.6604 | 0.00044776 | -0.0050812 | -0.0044834 | 0.00059781
6 | YES | 400.338 | 33.6522 | 0.00021008 | -0.0042381 | -0.0044834 | -0.00024524
6 | YES | 300.733 | 33.6455 | 0.00022587 | -0.0044022 | -0.0044834 | -8.1213e-05
6 | YES | 200.943 | 33.6245 | 0.0002199 | -0.0034828 | -0.0044834 | -0.0010006
6 | NO | 100.177 | 33.1768 | 0.011935 | 0.01865 | -0.0044834 | -0.023133
6 | NO | 50.4141 | 32.0403 | 0.00048319 | -0.093956 | -0.0044834 | 0.089473
6 | NO | 20.3752 | 31.4149 | 0.0024565 | -0.088989 | -0.0044834 | 0.084506
6 | NO | 10.2843 | 31.0811 | 0.00095705 | -0.1282 | -0.0044834 | 0.12371
6 | NO | 5.3594 | 31.0318 | 0.0003757 | -0.040302 | -0.0044834 | 0.035819
6 | NO | 0.48351 | 31.0201 | 0.00029634 | -2.3265 | -0.0044834 | 2.322
6 | NO | 0.48351 | 31.0201 | 0.00029634 | -0.0063209 | -0.0044834 | 0.0018376
6 | NO | 0.48351 | 31.0201 | 0.00029634 | 0.081278 | -0.0044834 | -0.085761
7 | YES | 517.809 | 33.2707 | 0.00026625 | -0.0029717 | -0.0043924 | -0.0014207
7 | YES | 200.577 | 33.2066 | 0.00033544 | -0.006588 | -0.0043924 | 0.0021956
7 | NO | 100.244 | 32.9917 | 0.00081813 | -0.028942 | -0.0043924 | 0.02455
7 | NO | 50.3362 | 32.1226 | 0.010156 | -0.020523 | -0.0043924 | 0.016131
7 | NO | 20.2675 | 31.4409 | 0.0096277 | -0.1278 | -0.0043924 | 0.1234
7 | NO | 9.9188 | 30.807 | 0.0089025 | -0.37171 | -0.0043924 | 0.36732
7 | NO | 5.1223 | 29.089 | 0.15724 | 1.3644 | -0.0043924 | -1.3688
7 | NO | 0.77146 | 24.8025 | 0.1441 | -0.54194 | -0.0043924 | 0.53755
8 | YES | 365.861 | 32.9594 | 0.00023941 | -0.0038948 | -0.0043014 | -0.0004066
8 | YES | 299.935 | 32.9557 | 0.00020033 | -0.0023956 | -0.0043014 | -0.0019058
8 | YES | 199.991 | 32.9143 | 0.0004072 | -0.0064735 | -0.0043014 | 0.0021721
8 | NO | 100.927 | 32.7653 | 0.0030382 | -0.0098381 | -0.0043014 | 0.0055367
8 | NO | 50.3263 | 32.1112 | 0.0067474 | -0.058174 | -0.0043014 | 0.053873
8 | NO | 20.7948 | 31.8343 | 0.00061997 | -0.10745 | -0.0043014 | 0.10315
8 | NO | 10.2701 | 31.6912 | 0.0029368 | -0.07218 | -0.0043014 | 0.067878
8 | NO | 5.3449 | 31.3717 | 0.0033182 | -0.27337 | -0.0043014 | 0.26907
8 | NO | 1.0399 | 31.0294 | 0.24394 | 2.7503 | -0.0043014 | -2.7546
9 | NO | 5.2115 | 23.7651 | 0.10977 | 0.47266 | -0.0042104 | -0.47687
9 | NO | 0.99503 | 21.2485 | 0.38476 | 4.9188 | -0.0042104 | -4.923
NO
NO
NO
NO
NO
NO
NO
YES
NO
NO
NO
NO
NO
NO
YES
YES
YES
NO
NO
NO
NO
NO
YES
NO
NO
YES
NO
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
NO
NO
NO
NO
NO
NO
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
FIT vs. Pressure
FIT vs. Sal1
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE