## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 13 March 2025 | Updated TSG channel names and formats. GG |

## PROCESSING NOTES

Cruise: 2022-081

Agency: Ocean Sciences Division (OSD)

Location: WCVI & Juan de Fuca Strait & Strait of Georgia

Project:  Orphan Moorings

Chief Scientist: Perreault L.

Platform: Vector

Date: 8 November 2022 – 13 November 2022

Processed by: Germaine Gatien

Date of Processing: 12 May 2023 –   15 May 2023

Number of HEX files:  16	Number of CTD casts processed: 16

Number of TSG files:   2	Number of TSG files processed:  2

# INSTRUMENT SUMMARY

An SBE911+ CTD #1453 was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#983DR), a SBE 43 DO sensor (#3038), a SeaPoint fluorometer (#3949), an SBE pH sensor (#0691), a Biospherical PAR sensor (#4565), a SPAR sensor (#20518) and an altimeter (#76341).

Seasave version 7.26.7.121 was used for acquisition.

The salinometer used at IOS was a Guildline model 8400B Portasal, serial # 73274.

A thermosalinograph (SeaBird 45 S/N 3411) was in use.

# SUMMARY OF QUALITY AND CONCERNS

The Science Log was in digital format. The header page had complete information about participants and equipment.

Data from the 10m soak were included in the files; those data were removed.

The 2 CTD salinity channels differed by about 0.004psu.

There was no bottle sampling during this cruise. Recalibration of CTD data was based on the results of cruise 2022-021.

There are 2 Thermosalinograph files with a 3.5 minute gap between them. There was no flow meter, intake thermistor or loop sampling. A proxy for intake temperature was derived by subtracting 0.2C° based on comparison with CTD casts. Salinity was recalibrated by adding 0.055psu based on comparisons during cruise 2022-070 since there was high variability in the TSG salinity leading to huge scatter in the comparison with CTD data from this cruise.

# PROCESSING SUMMARY

GENERAL

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

- The digital daily log  was obtained and checked for comments.

- The cruise summary sheet was completed.

- The history of the pressure sensor, conductivity and dissolved oxygen sensors were checked.

- The calibration control files were obtained; there were no errors in parameters.

##### Conversion of Full Files from Raw Data

All files were converted using files 2022-081-0001.xmlcon.

The Tau function was selected but not the hysteresis function as there was no deep sampling.

A few casts were examined and all expected channels are present.

- The alignment of the DO traces look good.

- The minimum DO value in Saanich Inlet was 0.08mL/L.

- The temperature and conductivity traces are close. No spikes were noted.

- The transmissivity, fluorescence PAR and SPAR look good.

- The buffer bottle was left on the pH sensor during cast 1. Other casts look ok.

- The altimetry looks excellent for most casts but at least 1 cast had a spike at the bottom.

##### WILDEDIT

WILDEDIT won’t help with the sort of spikes noted above, but the routine was run in case there are any single-point spikes in other casts.

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the

full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

A selection of casts were examined; the temperature profiles were complex, so the tests were not easy to interpret, but using +2.5s certainly improves the alignment and overall looks like a good choice for the DO sensor. ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

CELLTM was run using the default settings (α = 0.0245, β=9.5) for both the primary and secondary conductivity. A few casts were checked and those parameters appear to have worked well.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

The differences between the channel pairs was examined for a few casts but all are shallow. The same was the case on the 2 other cruise that had used these sensors since they were last serviced. An example of each cruise is reported below.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2022-019-0030 | 200 | -0.0006 | -0.0004 | -0.003 | shallow |
| 2022-021-0009 | 300 | +0.0007 | -0.0004 | -0.004 | steady |
| 2022-081-0001 | 280 | +0.0011 | -0.00035 | -0.0045 | High, steady |

The differences suggest some drift in calibration but the evidence is weak.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

To remove the 10m soak data file clip.csv was prepared based on plots of pressure versus scan number.

CLIP was run on CLN files; plots were examined to fine-tune the number of records removed.

##### Checking Headers

A header check was run on CLN files and there is evidence of spikes. When run on CLIP files most disappear though there are fairly large negative pressures at the end of one cast. The pumps were off, and the CTD was probably out of water as indicated by 0 fluorescence and very high transmissivity.

There are no negative fluorescence values.

A cross-reference list was checked. Some of the station names are slightly different from those in the log; they were not changed as they appear to relate to moorings at the same site.

The track plots look good and were added to the end of this report.

A check was made of the water depths entered in the file headers by adding the maximum depth sampled plus the altimetry reading when the CTD was at the bottom of the cast. We usually expect to get values  within ± 5m. The only event with a check value >5m was #15 and it was just 5.8m. The water depth in the header matches that in the log, but there may have been some variation through the stop. The digital logs at the time of this cruise entered the same positions and water depth for start, bottom and end of the cast, so this can’t be checked.  No adjustments were made to the files, as the differences are small. The overall results show the sounder was working well.

Surface check was run. The average was -0.09db before removal of the 10m soak data.

##### Shift

The same shift values were selected as were used for the previous use during 2022-021. Tests were run after each step and the results were found to be good.

Fluorescence

SHIFT was run on the fluorescence channel with an advance of 24 records. The fluorescence was very noisy making a judgement difficult, but cast #1 looked better and this setting is generally found to be the best choice.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. No further alignment is needed for the DO concentration channel.

Conductivity

This CTD had been used recently on cruise 2022-019 when the best settings found for alignment of conductivity were -0.9 records for the primary sensor and -0.3 records for the secondary.

These profiles have such little variation that it is harder to judge what works best, but the values used for the previous cruise produce much smoother T-S curves.

pH

There is too little variability and too much hysteresis to make a fine-tuned judgement about the alignment of pH with temperature, but the last time the sensor was used in 2022 a shift of +15 records was applied, so SHIFT was run with that setting. The results look ok.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

The P, T, S and DO sensors had been used on 1 other cruise since they were last recalibrated.

- 2022-019 – No calibration sampling; no recalibration applied.

- 2022-021 – Primary salinity within 0.002psu (only 5 samples)   Dissolved oxygen low by 5.5%.

Historic ranges  All temperature and salinity data fell within the local climatology where available.

Repeat Casts – The only repeat casts were #9 and 10. They were too shallow (and far apart in time) to provide any indication of CTD repeatability.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

The primary T-S pair were chosen for editing and eventual archiving because primary salinity was likely closer to bottles than the secondary salinity and the primary sensors were chosen during previous 2 cruises.

All DEL files were copied to *.EDT.

Many of the casts are in areas of active mixing with many unstable features that are likely real.

CTDEDIT was used to remove records that appear to be corrupted by shed wakes near the top and bottom of casts. Salinity was cleaned to remove spikes that appear to be due to small misalignments, but in most cases unstable features were left unedited.

Most files required editing near the top and bottom of casts.

No editing was applied to events #21 and 25.

The edited files were copied to *.EDT.

After editing T-S plots were examined for all casts; there remain many small unstable features but they are in areas when they may well reflect real conditions.

##### Recalibration

Pressure appears to be accurate and will not be recalibrated.

Based on the comparison with bottles during 2022-021 Salinity:T0:C0 appears to be accurate.

Based on the comparison with bottles during 2022-021 dissolved oxygen will be multiplied by 1.055.

File 2022-081-recal1.ccf was prepared to apply a correction as follows:

CTD DO Corrected = CTD DO * 1.055

CALIBRATE was run on the *EDT files.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel. It worked well to reduce noise.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined.

Profile plots were examined to see if there any problems. No problems were noted.

##### Final CTD File Steps (REMOVE and HEADEDIT)

- For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Secondary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T1:C1 and Flag.

- For cast #1 channel pH:SBE was also removed.

- A second SBE DO channel (with umol/kg units) was derived.

- Oxygen saturation was calculated and surface values were very low in Saanich Inlet and in the Gulf Islands, at about 60-73%. In the offshore and the central Strait of Georgia they ran from 80% and 88%.  These casts were fairly well mixed which usually leads to low values, so this offers no evidence about the quality of the DO recalibration.

- HEADER EDIT was used to fix formats and channel names and to add comments to the headers.

- The Standards Check routine was run and no problems were found.

- The Header Check was run; no problems were found.

- Profile and T-S plots were examined and no problems were found.

- The sensor history was updated.

##### Thermosalinograph processing

There were 2 thermosalinograph files.

There were no loop samples, flow meter or intake thermistor. The intake is at about 2m. The only method to check calibration is to compare with the CTD casts.

a.) Checking calibrations

The configuration file parameters were checked and are correct.

b.) Conversion of Files

The file was converted to CNV format.

The format of the file names was corrected by adding a 0.

There is a 3.5-minute gap between the 2 files.

The file was converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels.

A time-series plot was produced. The temperature and salinity channels have a single large spike.

The 2nd file was clearly running after the flow turned off at about 16:30 on November 12th.

Editing will be needed for both files.

The track plot looks fine and was added to the end of this report.

c.)  Checking Time Channel

- The CTD files were thinned to reduce the files to a single point from the downcast within 0.5db of 2db. These were exported to a spreadsheet which was saved as 2022-081-tsg-ctd-comp.xlsx. There were 14 CTD casts with data available that overlapped with TSG data, but the flow was obviously off for 3 of those casts, leaving just 11 for comparisons..

- The TSG ATC files were reduced to the times of CTD casts. Those data were added to 2022-081-tsg-ctd-comp.xlss.

- To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. The median differences were both 0.0000º with a maximum difference of 0.0007º. So the TSG clock worked well.

d.) Comparison of Temperature and Salinity from TSG and CTD data

The differences in temperature and salinity between the TSG and CTD are highly variable, especially for salinity.

|  | Median | Std Dev | Median | Std Dev |
| --- | --- | --- | --- | --- |
|  | Ttsg-Tctd | Ttsg-Tctd | SALtsg-SALctd | SALtsg-SALctd |
| 11 casts | 0.1905 | 0.6261 | -0.4119 | 0.5006 |
| 7 casts with lowest variation in salinity differences | 0.1978 | 0.1399 | -0.2014 | 0.3051 |

The standard deviations in the TSG data over 2 minutes show that variability in salinity was extremely high except for event #2, so the comparison is definitely limited. The salinity for event #2 is lower than the CTD by 0.060psu. The temperature variability is not particularly high.

There are insufficient data to inform salinity recalibration. The one cast with a salinity difference close to those found in 2 previous cruises was cast #2 in the offshore where variability was low. For temperature, the flow rate is important to know and the differences may be smaller than in previous cruises if that was set higher. So using a difference of 0.2Cº to make an estimate for intake temperature looks like the best choice, and that difference was found in the first 4 cruises that used this equipment since servicing.

See 2022-081-tsg-ctd-comp.xls for details.

Calibration History

This TSG has been used on 7 previous cruises.

- For the first 4 a proxy for intake temperature was created by subtracting 0.2C° from the lab temperature. The only cruise with sufficient sampling to estimate salinity calibration was from 2022-017 in June 2022 when salinity was estimated to be low by 0.02psu.

- F0r 2022-070 a proxy for intake temperature was created by subtracting 0.4C° from the lab temperature and 0.055psu was added to salinity

- For 2022-019 there was insufficient information so the 2022-070 recalibration was applied.

- For 2022-021 the differences found were close to those from 2022-070 but based on fewer points of comparison, so the 2022-070 results were applied.

Conclusions

1. The TSG clock worked well.

2. The TSG lab temperature was higher than the CTD temperature by roughly 0.2Cº.  This matches results seen in the first 4 uses of this equipment but is lower than seen during the 2 most recent ones. This may well be due to flow rate variations or ambient conditions of the ship or mismatches in depth of  TSG and CTD sampling. The comparison is reasonably consistent, so a proxy will be created for intake temperature by subtracting 0.2C° from lab temperatures.

3. The TSG Salinity is very noisy as is the comparison to CTD salinity. The one cast with low standard deviation in TSG salinity over 2 minutes shows a difference similar to those found during cruise 2022-070 which was the strongest recent comparison for this equipment. 0.055psu will be added to TSG salinity.

f.) Editing

1. CTDEDIT was used to remove a single-point spike in temperature conductivity and salinity- scan 928.

2. CTDEDIT was used to pad temperature conductivity and salinity values from scan 198 to the end of the file because flow in the loop had obviously stopped.

g.) Remove, Add Channel and Recalibration

REMOVE was used to remove the following channels: Scan Number and Flag channels.

Add Channels was used to add Channel Temperature:Lab with values set equal to Temperature:Primary.

Calibrate was run using file 2022-081-tsg-recal1.ccf to subtract 0.20C° from Temperature:Primary and to add 0.055psu to Salinity.

h.) Preparing Final Files

HEADER EDIT was used to change the DATA DESCRIPTION to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and time-series and all look fine.

The cruise plot was added to the end of this report.

Particulars

1. pH sensor data bad – buffer bottle left on.

25. Deck pressure -0.1db

CRUISE SUMMARY – CTD & TSG

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 1453 | No | Yes |  |  |
| Calibration Information - 1453 |  |  |  |  |  |  |  |
| Sensor |  |  | Pre-Cruise |  | Post Cruise |  |  |
| Name |  | S/N | Date | Location | Date | Location |  |
| Temperature |  | 2023 | 15Jul2022 | Factory |  |  |  |
| Conductivity |  | 3500 | 19Jul2022 | Factory |  |  |  |
| Secondary Temp. |  | 6565 | 15Jul2022 | Factory |  |  |  |
| Secondary Cond. |  | 5043 | 15Jul2022 | Factory |  |  |  |
| SBE 43 DO sensor |  | 3038 | 26Febl2022 | Factory |  |  |  |
| Altimeter |  | 76341 | 2Oct2021 |  |  |  |  |
| Pressure Sensor |  | 1453 | 21Jul2022 | Factory |  |  |  |
| Transmissometer |  | 983DR | 28Apr2021 | IOS |  |  |  |
| PAR |  | 4565 | 24Feb2021 | Factory |  |  |  |
| SPAR |  | 20518 |  |  |  |  |  |
| pH:SBE |  | 0692 | 14Jan2020 | Factory |  |  |  |
| SeaPoint Fluor. |  | 3949 | ? | Factory |  |  |  |

| Calibration Information – TSG 3411 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 3411 | 1Dec2019 | Factory |  |  |
| Conductivity | 3411 | 1Dec2019 | Factory |  |  |