## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2023-007

Agency: Ocean Sciences Division (OSD)

Location: Strait of Georgia

Project:  Strait of Georgia Pelagic Eco Survey

Chief Scientist: Gauthier S.

Platform: Sir John Franklin

Date: 22 February 2023 – 10 March 2023

Processed by: Germaine Gatien

Date of Processing: 5 July 2023 –  14 July 2023

Number of HEX files: 68 	Number of CTD casts processed: 66 (2 tests only- not processed)

Number of ROS files: 33	Number of bottle casts processed: 29 (2 non-DFO sampling + 2 tests)

# INSTRUMENT SUMMARY

An SBE911+ CTD #1453 was mounted in a rosette and attached were 2 Wetlabs CSTAR transmissometers (#1183DR and #1883DG), an SBE 43 DO sensor (#3038) mounted on the secondary pump, a SeaPoint fluorometer (#3641) mounted on the secondary pump, and an altimeter (#75321).

Seasave version 7.26.7.121 was used for acquisition.

The data logging computer was #103.

The salinometer used at IOS was a Guildline model 8400B Portasal, serial # 73274.

An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

The Science Log was in digital format. The header page had complete information about participants and equipment. There were also sampling logs for each bottle cast. Both logs included useful comments about problems encountered.

The deployment method generally included a 10m soak except for cast #44 when undisturbed surface water was required. Waits before firing bottles were at least 60s from 0 to 300db and at least 30s below that.

Most of the salinity calibration sampling from this cruise comes from bottles fired at the bottom of the cast which was usually about 10m above the seabed or between 0 and 50db. Both are areas liable to be compromised by poor flushing of Niskin bottles. In the top 50db high salinity gradients generally mean the Niskin bottles will contain higher salinity than ambient values. At the bottom the opposite is likely to be true due to incomplete flushing and shed wakes from above bouncing between the seabed and the rosette. Sampling at least 10m above the bottom of the cast (rather than 10m above the sea bed) provides better information for salinity calibration.

The primary temperature and salinity were selected for archiving since the primary salinity was closer to bottle samples and had slightly less pressure dependence. CTD salinity is estimated to be high by roughly 0.0015psu.

The SBE DO sensor has a fairly long response time so data accuracy is not as high when it is in motion as it is during stops for bottles. This will be especially true when vertical DO gradients are large. To get an estimate of the accuracy of the SBE DO data during downcasts (after recalibration) a rough comparison was made between downcast SBE DO and upcast titrated samples. Some of the difference will be due to problems with flushing of Niskin bottles and/or analysis errors and small mismatches in depth in the presence of large DO gradients.

Downcast (CTD files) Oxygen:Dissolved:SBE data for this cruise are considered, very roughly, to be:

±0.20 mL/L from 0-100db except in areas of very large DO gradients

±0.10 mL/L from 100db-200db

±0.04 mL/L from 200db-300db

±0.02 mL/L below 300db

# PROCESSING SUMMARY

GENERAL

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

- The digital daily log and rosette log sheets were obtained and checked for comments.

- Salinity, Nutrient, Extracted CHL and Dissolved Oxygen data were obtained in QF spreadsheets.

- The cruise summary sheet was completed.

- The history of the pressure sensor, conductivity and dissolved oxygen sensors were checked.

- The calibration control files were obtained; there were no errors in parameters and no changes through the cruise.

##### BOTTLE FILE PREPARATION

The ROS files were created using files 2023-007-ctd.xmlcon.

The ROS files were converted to IOS format; the hysteresis correction was not selected since there were no very deep casts; the Tau correction was selected.

The files were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for outliers. None were found.

A preliminary header check was run. No problems were noted.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. First, the file was sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs.

A few adjustments had to be made:

- File #1 – bottles fired as test – No sampling, no sample #s. Removed from processing stream.

- File #30 – As mentioned in log bottle was fired for non-DFO sampling; no sample # assigned. Removed from processing stream.

- File #33 – should be named #34. #33 was a net cast. The file was renamed and headers fixed.

- File #89 – No original sampling log sheet was found for this cast, nor in the scanned log sheets. It was found in the digital log. Samples #219 were meant to come from Niskin #2 but actually came from Niskin #3 but sample numbers were not changed. The 2 bottles were at the same depth. It was unclear how to fix this –list with Niskin #2 but with 2 flag? Or change sample #s.  The latter would cause much confusion, so it was put in with Niskin #2.

- File #96 – As mentioned in the log, bottles fired here were for non-DFO purposes, so no sample #s were assigned. Removed from processing stream.

The ADDSAMP file was then sorted on event number & then sample number.

It was used to add sample numbers to the BOT files – output *.SAM.

The SAM files were bin-averaged on bottle # and called SAMAVG.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

Next, the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were saved in file 2023-007-bot-hdr.txt.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2023-007_OXY*.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and saved as 2023-007oxy.csv. That file was converted into individual *.OXY files.

There were 2 samples in the DO file that had flag 3 and comments starting with “ALL:”, so flag 3 was added to the oxygen, chlorophyll and nutrient samples #33 and 79.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2023-007_CHL QF*.xlsx. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared and saved as 2023-007chl.csv. The csv file was then converted to individual CHL files.

SALINITY

Salinity analysis was obtained in file QF2023-007 SAL*.xlsx. The analyses were carried out in a temperature-controlled lab within 8-17 days of collection. The files were simplified and saved as 2023-007sal.csv. That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF_NUTS_2023-007*.xlsx. This includes a precision study. The file was simplified, saved as 2023-007nuts.csv and converted to individual NUT files.

The SAL, OXY, CHL and NUT files were merged with CST files in 4 steps.

The files were then put through CLEAN to reduce the headers to File and Comment sections only.

These files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so the MRGCLN1 files were reordered on Bottle_Number and saved as *. MRGCLN1s.

The MRGCLN1s files were then merged with SAMAVG files using merge channel Bottle_Number.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. None were found but one nutrient sample was found that was not expected. It was removed by the analyst and a comment was added to sample #168 silicate.

The minimum CTD Salinity in the spreadsheet was >25psu so no correction to silicate samples is required.

A header check was run and no problems were found.

##### Compare

Salinity

Most of the salinity calibration sampling from this cruise comes from bottles fired at the bottom of the cast which was close to the seabed or between 0 and 50db. Both are areas liable to be compromised by poor flushing of Niskin bottles. In the top 50db high salinity gradients generally mean the Niskin bottles will contain higher salinity than ambient values. At the bottom the opposite is likely to be true due to incomplete flushing and shed wakes from above bouncing between the seabed and the rosette.

When all data from 300db to the bottom are included the primary salinity is high by an average of 0.0018psu (std dev 0.0015) and the secondary is low by an average of 0.0034psu (std dev 0.0014).

Most of the differences above 200db suggest the bottles contained water with higher salinity values than ambient ones, as expected in high gradient zones. Between 200 and 300db the CTD salinity values are a little on the higher side, but all of those bottles were fired at the bottom of casts where flushing errors are of the opposite sign. However, The errors are quite small as local gradients were fairly low and waits before firing bottles were fairly long. The bottles between 300 and 400m were also fired at the bottom of casts and appear to be very slightly affected by poor flushing, but the errors are very small, likely similar to the errors during upcasts, just opposite sign..

There were 3 samples taken between 300 and 400db during the upcast. The primary salinity was higher than bottles by an average of 0.0010psu and the secondary was low by an average of 0.0037psu. Incomplete flushing of Niskin bottles would lead to the CTD salinity looking lower than it really is. In contrast, there were 4 bottles from between 300 and 400m fired at the bottom and the primary salinity is higher than bottles by an average of 0.003psu and secondary is lower by 0.0024psu. The difference between the two sets is 0.002 for the primary and 0.0013 for the secondary, giving a rough idea of the flushing errors below 300db as ~0.001psu.

A rough estimate is that primary salinity is high by ~0.0015psu and secondary low by about ~0.003psu.

Using the same points there is a little more pressure dependence in the secondary fit than in the primary, though neither is significant given the scatter.

.

For more detail see file 2023-007-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

There were 4 casts with DO sampling and they had quite different ranges of DO values.

Cast #7 was very well-mixed. Cast #47 had very large DO gradients.

When outliers were removed based on residuals plus a few near-surface samples where local gradients reversed, the fit was

CTD DO Corrected = CTD DO * 1.0445 +0.048

Most of the outliers correspond to high DO gradients and all but 1 involve the sample values being lower than the CTD; these both suggest outliers are primarily due to poor flushing. The one case where the sample was higher was from a depth where the DO gradient reversed, so again may be due to incomplete flushing of the Niskin bottle.

While cast #7 would have the most reliable flushing since gradients were very low, the offset is not reliable because there is such a small range of DO. So the data from cast #7 were forced to have the same offset as for the fit above and the resulting fit was:

CTD DO Corrected = CTD DO * 1.0452 -0.048

That suggests the main fit is reasonable.

None of the DO samples require changes to quality flags as they are not severely out of line and likely truly reflect the contents of the Niskin bottles.

For more detail see file 2023-007-dox-comp1.xls.

SBE Fluorescence vs Extracted Chlorophyll

COMPARE was run with Pressure as the reference channel.

The data were manipulated to allow plotting of FL vs CHL and FL/CHL vs CHL.

The usual pattern was found with the SBE fluorescence reading higher than extracted chlorophyll when CHL values were low and vice versa.

The range of CHL was 0.11 to 11.93ug/L with only 3 readings over 1.5ug/L.

At the low end of the CHL range fluorescence was 2.6 times the CHL and at the high end it was 0.4 times CHL.

The median ratio FL/CHL was 1.2 with standard deviation 0.4.

The fluorometer corresponds reasonably well to extracted CHL when CHL is <4ug/L.

For details see document 2023-007-fl-chl-comp1.xls.

##### Conversion of Full Files from Raw Data

All files were converted using files 2023-007-ctd.xmlcon.

The Tau function was selected but not the hysteresis function as there was no deep sampling).

A few casts were examined and all expected channels are present.

- The temperature and conductivity traces are reasonably close during downcasts. There are some spikes in conductivity though most casts appear to be spike free.

- The upcast channel pairs are farther apart than downcasts, as usual, but the primary conductivity looks particularly noisy.

- The dissolved oxygen, transmissivity and fluorescence traces look reasonable.

- The altimetry looks excellent for most casts but at least 1 cast had spikes at the bottom.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the

full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

This step removed all spikes in the secondary conductivity but there remain some spikes in at least 1 cast in the primary conductivity; that involved many multiple bad values, not single-point spikes.

##### ALIGN DO

A selection of casts were examined; the temperature profiles were complex but using +2.5s certainly improves the alignment and overall looks like a good choice for the DO sensor.

ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

CELLTM was run using the default settings (α = 0.0245, β=9.5) for both the primary and secondary conductivity. A few casts were checked and those parameters appear to have worked well.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

The differences between the channel pairs was examined for a few casts but the differences are noisy since sampling was fairly shallow. The sampling was even shallower for the only other cruises that have used these sensors since they were last serviced. An example of each cruise is reported below, with those highlighted in grey being from previous cruises.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2022-019-0030 | 200 | -0.0006 | -0.0004 | -0.003 | shallow |
| 2023-007-0009 | 300 | +0.0007 | -0.0004 | -0.004 | steady |
| 2022-081-0001 | 280 | +0.0011 | -0.00035 | -0.0045 | High, steady |
| 2023-007-0019 | 300 | +0.0017 | -0.00042 | -0.0058 | High, steady |
| 2023-007-0047 | 300 | +0.0018 | -0.00041 | -0.0056 | High, steady |
|  | 650 | +0.0008 | -0.00044 | -0.0051 | “ |
| 2023-007-0061 | 300 | +0.0009 | -0.00040 | -0.0057 | High, steady |
|  | 650 | +0.0012 | -0.00045 | -0.0051 | “ |
| 2023-007-0075 | 300 | +0.0015 | -0.00040 | -0.0056 | High, steady |
|  | 590 | +0.0009 | -0.00045 | -0.0052 | “ |

The temperature differences are highly variable but all larger than usual.

The conductivity differences are steady and consistent with previous cruises.

Salinity differences are a little larger than usual, no doubt due to the larger temperature differences.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

A header check was run.

There are no negative fluorescence values; the minimum was 0.095ug/L.

The minimum pressure was +0.36db. There were no data acquired out of water during the cruise, but during the previous use a deck measurement of -0.1db was taken. During a test cast in the lab values were between -0.3db and -0.5db with most values ~-0.4db.

These readings are within the accuracy expected for the pressure sensor. No recalibration will be applied.

A cross-reference list was checked and no problems were found.

The track plots look good and were added to the end of this report.

A check was made of the water depths entered in the file headers by adding the maximum depth sampled plus the altimetry reading when the CTD was at the bottom of the cast. In the offshore we usually expect to get check values within ± 5m, but that is likely too severe for this region.

There were 16 casts outside that limit. They were divided between positive and negative values, so this is not likely to be due to a poorly calibrated sounder. Plots of altimetry were checked to ensure the header entries were reasonable and the algorithm worked well in all cases. Sounder readings were known to be variable in a few cases due to shoaling during the cast or due to steep sides in narrow channels. Since the water depth at the time the CTD was close to the bottom is the most useful value, estimates were made of water depth for the cases with a check value >8m by adding the altimetry header reading to maximum depth sampled.

The water depth entry in the headers were adjusted for 9 CLN files and for 4 affected bottle files.

##### Shift

Fluorescence   SHIFT was run on the fluorescence channel with setting +24 records. The fluorescence was very noisy making judgment difficult, but cast #1 looked better; this setting is generally found to be the best choice.

Dissolved Oxygen  The Dissolved Oxygen voltage channel was aligned earlier. No further alignment is needed for the DO concentration channel.

Conductivity  This CTD had been used recently on cruise 2022-081 when the best settings found for alignment of conductivity were -0.9 records for the primary sensor and -0.3 records for the secondary. For this cruise larger shifts looked best, at -1.1 and -0.7 records.

SHIFT was run twice to “advance” the primary conductivity by -1.1 records and the secondary by -0.7 records. Salinity was recalculated.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

The P, T, S and DO sensors had been used on 3 other cruises since they were last recalibrated.

- 2022-019 – No calibration sampling; no recalibration applied.

- 2022-021 – Primary salinity within 0.002psu (only 5 samples)   Dissolved oxygen low by 5.5%.

- 2022-081 – No recalibration sampling. Recal based on 2022-021. Sal difference ~0.004

Historic ranges  For many casts there was no local climatology available.

- Where it was available, all salinity data were within the historic ranges except for slightly low salinity at the bottom of the cast at Station 2.

- There were many excursions from the local temperature climatology. This is a 3-standard deviation climatology which is not well suited to areas close to shore, and the date of this cruise straddled the end of the 3-month range of one climatology and the beginning of the next one. There are few data in the climatology from late winter in this area. In the Gulf Islands area casts were very well mixed with some sections having temperature below the climatology minimum. In the southern Strait of Georgia temperatures were within the climatology but there were some low values in the northern Strait of Georgia.

Given the limitations in the climatology the low temperatures are not considered evidence of calibration drift.

##### DETAILED EDITING

The primary T-S pair were chosen for editing and eventual archiving because primary salinity was likely closer to bottles than the secondary salinity, had slightly less pressure dependence and the primary sensors were chosen during previous uses.

All DEL files were copied to *.EDT.

Many of the casts are in areas of active mixing with many unstable features that are likely real.

CTDEDIT was used to remove records that appear to be corrupted by shed wakes near the top and bottom of casts. Salinity was cleaned to remove spikes that appear to be due to small misalignments of T and C, but in many cases unstable features were left unedited.

Most files required editing near the top and/or bottom of casts.

No editing was applied to events #52, 61

The edited files were copied to *.EDT.

After editing T-S plots were examined for all casts; there remain many small unstable features but they are in areas when they may well reflect real conditions. No further editing was applied.

##### Recalibration

All salinity in bottle files are >25psu, so silicate does not need recalibration.

A rough estimate is that primary salinity is high by ~0.0015psu and secondary low by about ~0.003psu.

There is a little more pressure dependence in the secondary fit than in the primary, though neither is significant given the scatter. No recalibration will be applied due to the limited calibration data and the primary salinity being within expected accuracy.

Pressure will not be recalibrated.

File 2023-007-recal1.ccf was prepared to apply a correction as follows:

CTD DO Corrected = CTD DO * 1.0445 +0.048

CALIBRATE was run on the MRGCLN2 and SAM files – output MRGCOR1 and SAMCOR1.

COMPARE was then run with input recalibrated CTD DO and bottle data. When the same data were excluded as in the original fit the CTD DO was found to be lower than bottles by 0.001mL/L (std dev 0.012mL/L).  So the calibration applied worked well.

CALIBRATE was run on the *EDT files.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel. It worked well to reduce noise.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from roughly the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and delayed response and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles.

The CTD DO was higher than the titrated samples by an average of 0.083mL/L when all bottles were included, but when the same values were rejected as in  the original analysis, it was high by 0.033mL/L with a standard deviation of 0.036mL/L. The recalibration was obviously effective. We expect the downcast CTD DO to be a little higher than bottles due to a combination of slow sensor response and incomplete flushing of bottles.

Downcast (CTD files) Oxygen:Dissolved:SBE data for this cruise are considered, very roughly, to be:

±0.20 mL/L from 0-100db except in areas of very large DO gradients

±0.10 mL/L from 100db-200db

±0.04 mL/L from 300db-200db

±0.02 mL/L below 300db

For more detail see file 2023-007-dox-comp3.xls.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average of values in the bin will be used.    Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined.

Profile plots were examined to see if there any problems. No problems were noted.

##### Final CTD File Steps (REMOVE and HEADEDIT)

- For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Secondary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T1:C1 and Flag.

- A second SBE DO channel (with umol/kg units) was derived.

- Oxygen saturation was calculated and surface values were fairly low, between 88% and 98%. .  These casts were well generally well mixed which usually leads to low values, so this offers no evidence about the quality of the DO recalibration.

- HEADER EDIT was used to fix formats and channel names and to add comments to the headers.

- The Standards Check routine was run and no problems were found.

- The Header Check was run; no problems were found.

- Profile and T-S plots were examined and no problems were found.

- The sensor history was updated.

##### Final Bottle Files

- The MRGCOR1 files were put through SORT to order on increasing pressure.

- For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Secondary,  Oxygen:Voltage:SBE, Transmissivity, Altimeter, Status:Pump, Descent_Rate and Flag.

- A second SBE DO channel with mass units was added for both the CTD DO and Titrated DO.

- REORDER was run to get the pairs of DO channels together.

- HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

- Standards check was run and no problems were found..

- Bottle data were exported to a spreadsheet and no problems were found.

- A header check were run. No problems were found.

- Plots of each file were examined and no problems were found.

- A cross-reference listing was produced for the CHE files.

Particulars (notes from logs of concern in processing plus few remarks on processing)

1. All bottles fired at 50m - test

30. Surface bottle tripped for DNA – no sample # assigned.

44. Surface soak only – no 10m soak.

59. Soft bottom, noisy altimeter, sounder seeing sidewalls.

64. Forgot to start acquisition after 10m soak, returned to surface and waited ~2min; overwrote file.

77. Sounder indicates shallowing through downcast.

83. Comparison cast.

86. BE time late by 3min

89. 7m above bottom. Niskin #2 didn’t fire – all Niskin #2 samples taken from Niskin #3 but sample numbers not changed.

96. Bottles fired for non-DFO purposes – no sample numbers assigned.

CRUISE SUMMARY – CTD

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 1453 | Yes | Yes |  |  |
| Calibration Information - 1453 |  |  |  |  |  |  |  |
| Sensor |  |  | Pre-Cruise |  | Post Cruise |  |  |
| Name |  | S/N | Date | Location | Date | Location |  |
| Temperature |  | 2023 | 15Jul2022 | Factory |  |  |  |
| Conductivity |  | 3500 | 19Jul2022 | Factory |  |  |  |
| Secondary Temp. |  | 6565 | 15Jul2022 | Factory |  |  |  |
| Secondary Cond. |  | 5043 | 15Jul2022 | Factory |  |  |  |
| SBE 43 DO sensor |  | 3038 | 26Febl2022 | Factory |  |  |  |
| Altimeter |  | 76341 | 10Feb2021 |  |  |  |  |
| Pressure Sensor |  | 1453 | 21Jul2022 | Factory |  |  |  |
| Transmissometer |  | 1185DR | 21Feb2023 | IOS |  |  |  |
| Transmissometer |  | 1883DG | 28Apr2021 | Factory |  |  |  |
| SeaPoint Fluor. |  | 3641 | ? | Factory |  |  |  |