## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2023-009

Agency: Ocean Sciences Division (OSD)

Location: WCVI

Project:  WCVI Juvenile Salmon Survey

Chief Scientist: Tabata A.

Platform: Sir John Franklin

Date: 6 October 2023 – 13 October 2023

Processed by: Germaine Gatien

Date of Processing: 16 July 2024 –  19 July 2024

Number of HEX files:  20		Number of CTD casts processed:  20

Number of bottle files: 18 		Number of bottle files processed:  17 (1 file had no samples)

# INSTRUMENT SUMMARY

An SBE911+ CTD #1453 was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#983DR), a SBE 43 DO sensor (#3038) on the primary pump, a SeaPoint fluorometer (#3849) on the secondary pump and an altimeter (#76341).

Seasave version 7.26.7.121 was used for acquisition.

The data logging computer was Lenovo #102. The CTD Deck Unit was SBE11 #1043.

The salinometer used at IOS was a Guildline model 8400B Portasal, serial # 73274.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log and rosette logs were available. There were useful comments including some deck pressure readings and notes about the deployment strategy used which varied through the cruise. For  casts #19-43 acquisition started during a surface soak. There was a 10m soak for events #1-13 and 47-53.

The bl file for cast #35 was empty; appropriate data were extracted from the full file from the level at which sampling was done. This enabled creation of a rosette file.

For most casts there were 60s waits before firing shallow bottles, but for cast #51 the waits were shorter. Even in the offshore area a 60s wait is recommended for sampling down to 400m, especially if there is dissolved oxygen sampling.

The sampling plan was good given the nature of the cruise. Unfortunately, this CTD has only been used on cruises with little sampling and some contradictory results in comparisons. There was a post-cruise factory calibration in early 2024 which enables a rough estimate of calibration drift. The comparison of 2023-009 salinity samples with CTD salinity are reasonably close to that estimate. The comparison of dissolved oxygen samples with CTD dissolved oxygen values was also in reasonable agreement with the post-cruise results at depth, but near-surface values were not. There are too few samples to make an estimate of SBE dissolved oxygen accuracy except to say that surface values may be low by up to 2%.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

- The digital daily log and rosette log sheets were obtained and checked for comments.

- Salinity, Extracted CHL, Dissolved Oxygen and Nutrient data were obtained in QF spreadsheets.

- The cruise summary sheet was completed.

- The history of the pressure sensor, conductivity and dissolved oxygen sensors were checked.

- The calibration control files were obtained; there were no errors in parameters and no changes during the cruise. One file was saved as 2023-009-ctd.xmlcon.

##### BOTTLE FILE PREPARATION

The ROS files were created using files 2023-009-ctd.xmlcon.

The ROS files were converted to IOS format; the hysteresis correction was not selected since there was only 1 deep cast and it only went to 1300db; the Tau correction was selected.

The files were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files and no significant outliers were found.

A preliminary header check was run. No obvious problems were found.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. First, the file was sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs.

The ADDSAMP file was then sorted on event number & then sample number.

It was used to add sample numbers to the BOT files – output *.SAM.

The SAM files were bin-averaged on bottle # and called SAMAVG.

A few adjustments were needed:

- File 2023-009-0003.samavg was edited to remove the data from Niskin #2 which did not close, and to move sample #2 data to Niskin #1.

- The MRGCLN1S file also needed to be adjusted to reflect that Sample #s came from Niskin #1.

- The file for event #21 was misnamed as #20.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

Next, the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were saved in file 2023-009-bot-hdr.txt.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2023-009_OXY*.xlsx which includes flags and comments but there were too few samples for a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and saved as 2023-009oxy.csv. That file was converted into individual *.OXY files.

No samples had a flag that applies to ALL sample types.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2023-009_NUTS*.xlsx. This includes a precision study. The file was simplified, saved as 2023-009nuts.csv. The file was converted to individual NUT files.

SALINITY

Salinity analysis was obtained in file QF2023-009 SAL*.xlsx. The analyses were carried out in a temperature-controlled lab within 19-26 days of collection. The files were simplified and saved as 2023-009sal.csv. That file was then converted to individual SAL files.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2023-009_CHL QF*.xlsx. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared and saved as 2023-009chl.csv. The csv file was then converted to individual CHL files.

The SAL, CHL, OXY and NUTS CSV files were converted into individual files for each cast.

Then they were merged with CST files in 4 steps.

The files were then put through CLEAN to reduce the headers to File and Comment sections only.

These files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so the MRGCLN1 files were reordered on Bottle_Number and saved as *. MRGCLN1s.

The MRGCLN1s files were then merged with SAMAVG files using merge channel Bottle_Number.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets.

Cast #1 had no sampling so was dropped from the list to be processed.

Cast #3 had 2 bottles fired but only Niskin 1 was sampled. Niskin 2 did not close so that line was removed from the SAMAVG file and the merge rerun..

There are no salinity values <25psu when bottles were fired, so silicate recalibration is not needed.

A header check was run and no problems were found.

##### Compare

Salinity

There were only 3 casts with salinity sampling but one was sampled to 1306db. Using only the 9 samples below 200db the primary salinity was found to be low by an average of 0.0020psu and the secondary low by 0.0079psu. The standard deviation was 0.0007psu for both. There is more pressure dependence in the secondary fit than in the primary.

All significant outliers are likely due to noise in the CTD data. No quality flags are recommended.

For more detail see file 2023-009-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

The data come from only 2 casts, one shallow close to shore and one deep offshore. The fit found when excluding only obvious outliers was:

CTD DO Corrected = CTD DO * 1.077 +0.007 		(Fit 1)

The near-surface data were very noisy and when most bottles above 100db were excluded the fit was:

CTD DO Corrected = CTD DO * 1.060 +0.028	 	(Fit 2)

The sample from 25db of cast #33 did look ok in the general fit while all other shallow samples had much larger differences.

That fit is much closer to previous results from this sensor and more in line with the post-cruise calibration than was found during 2023-061.

Cast #33 had a steep vertical DO gradient in the top 100db and cast #51 had a mix of steep gradients and reversals. These conditions are definitely challenging in making a comparison. The near-surface data are in well-mixed layers so flushing of bottles is likely good there while but poorer in the high gradients below. the CTD response is good in the well-mixed layers but likely a little “behind” in the high gradients. All samples from the high-gradient region were outliers in the comparison.

When separated into near-shore and off-shore groups some of the variability becomes easier to understand. Near-shore the bottles likely do not flush as well as offshore which tends to reduce the differences between CTD and bottles. Offshore the gradients are small at depth but large above 100db. The bottles likely flush better, but the CTD may not do so well in responding, so the CTD may look lower compared to bottles, though that effect should not be large. There are many small-scale reversals in the offshore cast that confound confidence in the fits, especially in the top 100db.

The deep fit (Fit 2) is reasonably close to results of previous cruises (none had great DO sampling) and to the post-cruise estimate that SBE DO is low by about 6%. Given the few data available it seems best to use that figure for the correction.

No outliers appear to warrant further flagging of samples since the CTD data are noisy.

Plots of SBE DO and Titrated DO versus salinity turned up no outliers.

For more detail see file 2023-009-oxy-comp1.xls.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable. Chlorophyll values ranged from 0.42 to 2.55ug/L. Overall the fluorescence reads higher than CHL, but that is because there are many low CHL values. When a fit is forced through the origin, Fluorescence is 77% of extracted chlorophyll which is close to the results of cruise 2023.008.

When the ratio FL/CHL is plotted against CHL, the usual pattern is seen with  FL higher than CHL when CHL is <1ug/L and lower at higher CHL value. The shape is typcial of the performance for the fluroometer.

No updates to flag samples are recommended.

For full details for the COMPARE run see file 2023-009-fl-chl-comp1.xls.

##### Conversion of Full Files from Raw Data

All files were converted using files 2023-009-ctd.xmlcon.

The Tau function was selected but not the hysteresis function as there was little sampling below 200db with only one deep cast to just 1300db.

A few casts were examined and all expected channels are present.

- There seems to be a vertical offset between temperature traces during the downcast that could be misalignment, or it could be something affecting response in one of the pair. Small misalignment is expected at this stage, and it is very slight when stopped for bottle sampling.

- The transmissivity and fluorescence look good.

- The altimetry looks excellent .

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the

full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

A selection of casts were examined; the temperature profiles were complex, so the tests were not easy to interpret, but using +2.5s certainly improves the alignment and overall looks like a good choice for the DO sensor. ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

CELLTM was run using the default settings (α = 0.0245, β=9.5) for both the primary and secondary conductivity. A few casts were checked and those parameters appear to have worked well.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

At this stage casts were examined to see how well the previous steps worked.

A few spikes noted in conductivity were removed by WILDEDIT.

The alignment of dissolved oxygen looks good.

The salinity channels differ by about 0.0055psu during downcasts and upcast stops, but underway upcast data are extremely noisy, especially the primary salinity. So it is hard to judge how well CELLTM worked.

The differences between the channel pairs was examined for the only cast that sampled deeper than 250m. Samples from 4 other cruises that used these sensors previously and since they were last serviced are reported as well as the current cruise. Most sampling was shallow.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2022-019-0030 | 200 | -0.0006 | -0.0004 | -0.003 | shallow |
| 2023-021-0009 | 300 | +0.0007 | -0.0004 | -0.004 | steady |
| 2022-081-0001 | 280 | +0.0011 | -0.00035 | -0.0045 | High, steady |
| 2023-007-0047 | 300 | +0.0018 | -0.00041 | -0.0056 | High, steady |
|  | 650 | +0.0008 | -0.00044 | -0.0051 | “ |
| 2023-007-0075 | 300 | +0.0015 | -0.00040 | -0.0056 | High, steady |
|  | 590 | +0.0009 | -0.00045 | -0.0052 | “ |
| 2023-008-0022 | 300 | -0.0004 | -0.00051 | 0.0055 | High, V Noisy |
|  | 720 | -0.0004 | -0.00057 | -0.0055 | “ |
| 2023-061-0018 | 300 | -0.0008 | -0.00038 | -0.0028 | High Noisy |
|  | 500 | -0.0002 | -0.00039 | -0.0029 | “ |
|  | 1000 | +0.0002 | -0.00029 | -0.0033 | “ |
| 2023-061-0054 | 300 | -0.0004 | -0.00025 | -0.0027 |  |
| 2023-061-0090 | 300 | +0.0001 | -0.00019 | -0.0022 | High, X Noisy |
|  | 500 | Too noisy | Too noisy | Too noisy |  |
|  | 1000 | +0.0015 | -0.00008 | -0.0024 | “ |
|  | 1500 | +0.0012 | -0.00008 | -0.0022 | “ |
| 2023-009-0051 | 300 | -0.0001 | -0.00046 | -0.0051 | High, X Noisy |
|  | 500 | +0.0013XN | -0.00026XN | -0.0050XN | “ |
|  | 700 | +0.0007 | -0.00039 | -0.0053 | “ |
|  | 1000 | +0.0008 | -0.00036 | -0.0050 | “ |

The differences in salinity are similar to those from cruises 2023-007 and 2023-008, but differ from cruise 2023-061, a cruise that stood out as odd in many ways. Conductivity differences are slightly smaller than in 2023-008 but similar to 2023-007. Temperature differences differ in sign from 2023-008 but are closer to 2023-007, though the variations are small and likely not significant. Given the noise in these plots, the only clear conclusion is that 2023-009 salinity differences differ from 2023-061 and look more like cruises before 2023-061.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

CLIP was not needed.

##### Checking Headers

A header check was run on CLN files and no problems were found.

There are no negative fluorescence values.

The lowest pressure recorded was 0.096db at the end of cast #10. The pumps were on and there are pad values in conductivity, the transmissivity is 0. This is likely right at the surface. There were also deck pressure readings ranging from -0.17db to -0.25db. Pressure appears to be very accurate at the surface.

A cross-reference list was checked. No problems were found.

A check was made of the water depths entered in the file headers by adding the maximum depth sampled plus the altimetry reading when the CTD was at the bottom of the cast. There were no differences >10m.  We investigate differences >5m to ensure header entries for water depth agree with the log and that the altimetry minimum was clear. The following casts were checked:

- Casts #1 and #24 had no entries in the file headers. The log entries look appropriate and were entered in the file headers. This change affects the bottle file for cast #24 as well.

- Cast #20 had a different entry from the log but the file entry looks better, so no change was made.

- Casts #25 and #51 had log entries the same or close to the log entries; altimetry looked good. No changes were made.

Surface check was run. The average was 1.6db.

##### Shift

Fluorescence

SHIFT was run on the fluorescence channel with an advance of 24 records. The fluorescence was very noisy making a judgement difficult, the data are definitely better aligned after this step.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. It is very hard to judge the alignment due to extremely noisy upcast data, but it appears that no further adjustment is needed for the DO concentration channel.

Conductivity

The settings used on other recent cruises were:

2023-081: -0.9 and -0.3  records for the primary and secondary conductivity. There was a pH

2023-007: -1.1 and -0.7 records for the primary and secondary conductivity. No pH

2023-008: -0.9 and -0.3 but  the differences between the two choices were small. There was a pH

2023-061: -1.0 and -0.6 records for the primary and secondary conductivity. No pH

For this cruise values of -1.0 and -0.6 also looked best for the primary and secondary conductivity, though there was little difference using slightly different values for each. There was no pH sensor.

SHIFT was run apply a shift of -1.0 for the primary conductivity and -0.6 for the secondary. Salinity was recalculated for each channel.

DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

- 2022-019 – No calibration sampling; no recalibration applied.

- 2022-021 – Primary salinity within 0.002psu (only 5 samples)   Dissolved oxygen low by 5.5%.

- 2022-081 – No recalibration sampling. Recal based on 2022-021. Sal difference ~0.004

- 2023-007 – Limited sampling mostly fired at bottom of casts which tends to lead to salinity samples lower than ambient values, so CTD salinity looks relatively higher than it really is. Primary salinity high by ~0.0015 and secondary low by ~0.003psu. DO recalibrated using slope 1.0445 and offset +0.048. Surface pressure variable but all within ±0.5db.

- 2023-008 – Limited sampling with Niskin flushing likely poor. Estimate DO low by 6% and primary and secondary salinity low by 0.003psu and 0.0075psu, respectively based on previous cruises and factory post-cruise calibrations in early 2024.

- 2023-061 – Salinity sampling only from 1 cast. No DO sampling. Recal by adding 0.002psu to Sal0, DO multiplied by 1.06.

Historic ranges  Local climatology was not available for all casts. A few casts had brief and small sections with temperatures slightly high – one at 400db and another at 70db. The only excursion in salinity from the climatology was low salinity near the bottom of a cast in Queen Charlotte Strait, not an area best suited to this type of climatology. None of these cases suggest a problem with calibration.

Repeat Casts – There were no repeat casts.

Post-Cruise Calibration – Very rough estimates of calibration drift were made based on the post-cruise calibration in early 2024. The dissolved oxygen sensor was reading low by about 6%, primary salinity was low by roughly 0.003psu and secondary salinity low by roughly 0.008psu.

##### DETAILED EDITING

The DEL files were sent to the QC-CTD site for predictions.

All DEL files were copied to *.EDT.

As noted in the previous use of this CTD there is a small-scale jitter in the salinity. However, the variability is much smaller and looks similar to variations in conductivity.

The choice of which sensor pair to edit is not obvious. The primary salinity has a larger jitter, but the comparison of secondary sensors to bottles had some pressure dependence while the primary had very little. The primary salinity is closer to bottles and has been chosen during previous uses since the secondary had more unstable features.

The primary temperature and salinity were selected for editing.

CTDEDIT was used to remove records that appear to be corrupted by shed wakes near the top and bottom of casts. Salinity was cleaned to remove spikes that appear to be due to small misalignments. In the casts in narrow channels there are many unstable features that could be real, so editing was kept very light.

All files required editing, mostly near the top and bottom of casts.

The edited files were copied to *.EDT.

After editing T-S plots were examined for all casts; there remain many small unstable features but they are in areas when they may well reflect real conditions.

##### Recalibration

Pressure appears to be accurate and will not be recalibrated.

Salinity:T0:C0 was found to be low by about 0.0020psu. The evidence for salinity calibration is weak from previous cruises that used the same sensors since they were last serviced. However, it is close to the results of cruise 2023-007 and is not far off from the estimated error when the sensor was next serviced in early 2024. Although the secondary salinity is not being chosen for the archive and has not been through the graphical editor, it was recalibrated as well based on the COMPARE result.

Sampling for dissolved oxygen was also limited but the fit of deeper data is reasonably close to the estimate based on the post cruise calibration when it appears to have been reading low by roughly 6% in early 2024. In the absence of a stronger calibration sampling history, that value looks like the best choice for recalibration of these data.

File 2023-009-recal1.ccf was prepared to add 0.002psu to channel Salinity:T0:C0 0.0079 to channel Salinity:T1:C1 and to apply a correction as follows:

CTD DO Corrected = CTD DO * 1.060 +0.028

There was no salinity<25psu in bottle files so the silicate correction was not needed.

CALIBRATE was first run on the MRGCLN2 files and SAM files.

COMPARE was rerun on salinity and dissolved oxygen to check that the corrections were applied properly and they were.

CALIBRATE was run on the *EDT files.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel. It worked well to reduce noise.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from roughly the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and delayed response and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles.

The CTD DO was higher than the titrated samples by an average of 0.07mL/L and the standard deviation was 0.026mL/L. We expect the CTD data to be higher than bottles due to incomplete flushing of bottles and slow response of the SBE DO.

There are too few samples to make an estimate of SBE dissolved oxygen accuracy except to say that  shallow SBE DO values may be too low by up to 2%.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined.

Profile plots were examined to see if there any problems. No problems were noted.

##### Final CTD File Steps (REMOVE and HEADEDIT)

- For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Secondary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T1:C1, Flag and Prediction_Flag.

- A second SBE DO channel (with umol/kg units) was derived.

- REORDER was run to get the 2 DO channels together.

- Oxygen saturation was calculated and surface values varied from 57% to 125% with only the Saanich Inlet site being >100%. Saturation was mostly low in inlets. Offshore values were about 98-99%.

- HEADER EDIT was used to fix formats and channel names and to add comments to the headers.

- The Standards Check routine was run and no problems were found.

- The Header Check was run; no problems were found.

- Profile and T-S plots were examined and no problems were found.

- The sensor history was updated.

- The EDT files were zipped and sent to the CTD-QC program (Lee Croft).

##### Final Bottle Files

- The MRGCOR1 files were put through SORT to order on increasing pressure.

- For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Secondary,  Oxygen:Voltage:SBE, Transmissivity, Altimeter, Status:Pump, Descent_Rate and Flag.:

- A second SBE DO channel with mass units was added for both the CTD DO and Bottle DO

- REORDER was run to get the pairs of DO channels together.

- The entry for bottle #2 in event 51 was removed as there was no sampling.

- EDIT HEADERS was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

- Standards check was run and 1 problem was found and fixed.

- Bottle data were exported to a spreadsheet and a few problems were found and corrected in the bottle files.

- A header check were run. No problems were found.

- Plots of each file were examined and no problems were found.

- A cross-reference listing was produced for the CHE files.

Particulars

1. Niskin 2 did not close. (It was fired.)

3. Deck pressure -0.24db

17- 44. Archiving started at surface soak.

21. Deck pressure -0.25db

25. Deck pressure -0.25db but noisy

33. Niskin 2 did not fire

35. BL file empty – created file based on time stopped near end of stop at appropriate level of upcast.

43. Upper 80m fluorescence quite noisy, but did have big storm

47. Deck pressure -0.17db

51. Anticipating that Niskin 2 would not fire, fired 3 at 1000m too. No sample # assigned to Niskin 2.

CRUISE SUMMARY – CTD

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 1453 | Yes | Yes |  |  |
| Calibration Information - 1453 |  |  |  |  |  |  |  |
| Sensor |  |  | Pre-Cruise |  | Post Cruise |  |  |
| Name |  | S/N | Date | Location | Date | Location |  |
| Temperature |  | 2023 | 15Jul2022 | Factory | 8Feb2024 |  |  |
| Conductivity |  | 3500 | 19Jul2022 | Factory | 23Feb2024 |  |  |
| Secondary Temp. |  | 6565 | 15Jul2022 | Factory | 17Feb2024 |  |  |
| Secondary Cond. |  | 5043 | 15Jul2022 | Factory | 23Feb2024 |  |  |
| SBE 43 DO sensor |  | 3038 | 26Febl2022 | Factory | 11March 2024 | Factory |  |
| Altimeter |  | 76341 | 2Oct2021 |  |  |  |  |
| Pressure Sensor |  | 1453 | 21Jul2022 | Factory |  |  |  |
| Transmissometer |  | 983DR | 28Apr2021 | IOS |  |  |  |
| SeaPoint Fluor. |  | 3949 | ? | Factory |  |  |  |

Test cast (Event #1) in Saanich Inlet excluded to enable clearer maps.