## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2023-012

Agency: OSD

Location: Strait of Georgia, Juan de Fuca Strait, WCVI

Project: Acoustic Moorings

Chief Scientist: O’Neill C.

Platform: Vector

Date: 14 October 2023 – 20 October 2023

Processed by: Jeannette Bedard

Date of Processing:  20 November 2023 – 15 December 2023

Number of HEX files: 	76		Number of CTD files processed: 76

# INSTRUMENT SUMMARY

CTD #1222 with the following instruments:

- Conductivity sensor 4434 paired with temperature sensor 5013

- Conductivity sensor 4448 paired with temperature sensor 5725

- Piggybacked: Seapoint Fluorometer (#2228), SBE 43 DO sensor (#4367) and Valeport VA500 altimeter (#76341).

Seasave version 7.26.7.121 was used for acquisition.

# SUMMARY OF QUALITY AND CONCERNS

The digital Daily Science Log Book was in good order with useful comments about the CTD.

There was no calibration sampling for the salinity or dissolved oxygen sensors. The 2 salinity channels were close in value and recently recalibrated, so drift was likely small. The SBE DO sensor was brand new.

For a few casts the fluorometer had surface values of zero, these were removed.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

- Both conductivity sensors and the pressure sensor were calibrated in March 2022. This was the third cruise these sensors were used on since then.

- The oxygen sensor was new and this was the first cruise it was used.

- The digital Log Book was obtained.

- The cruise summary sheet was completed.

The configuration file used at sea had the latest calibration parameters. No changes were made during the cruise. The file was saved as 2023-012-0001.xmlcon.

##### Conversion of Full Files from Raw Data

All files were converted using file 2023-012-0001.xmlcon.

The Tau function was selected but not the hysteresis function since there was no deep sampling.

A few casts were examined. The T and C pairs were close during downcasts with upcasts a little further apart and noisy, but not as much as usual due to the absence of a rosette. Fluorescence looks normal. Dissolved oxygen look normal. Altimetry data looks normal.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, depth, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration. The DO alignment looks reasonably good. There were no deep casts.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

- The cross-reference check and header check were run. A few minor capitalization issues were found in the station names and fixed.

- Surface check was run and found an average  of 1.3db which is reasonable.

- Cruise tracks were plotted and no problems were found. Plots with station names (every third one listed) and event numbers (every second one listed) were added to the end of this report.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Plots show that the vertical offset between downcast and upcast for fluorescence is reasonably close to that for temperature after this step.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked good, and it did. No further alignment is needed for the DO concentration channel.

Conductivity

Tests were run on a few casts to assess what settings are best to align conductivity with temperature (as judged by the effect on salinity as seen in T-S space). The best settings were +0.3 and +0.1 records for the primary and secondary, respectively.

SHIFT was run twice on all SBE911 casts using 0.3 records for the primary and 0.1 records for the secondary conductivity. Salinity was recalculated.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Experience with these sensors since last factory service –  None

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. The few excursions from the climatology were in the Gulf Islands.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

The primary channels were chosen for recalibration.

Active mixing in this region leads to unstable features that may be real, thus only light editing was performed.

Most casts required light editing at the top and bottom where the instrument descent rates were very low.

Casts that did not require editing were: 7, 10, 15, 16, 19, 20, 21, 23, 24, 29, 35, 38, 39, 40, 55, 67, 76 , 77, 83, 88, 90, 92

Notes about editing applied were added to the files.

The edited files were copied to *.EDT.

After editing T-S plots were examined for all casts. No further editing was needed.

##### Corrections to Pressure, Salinity and Dissolved Oxygen Concentration

There is no indication of a problem with pressure calibration and there is no information available upon which to base recalibration of conductivity. No recalibration was applied.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

- Zero surface fluorometer values were removed for casts 15 (to 5.8 db), 16 (to 2 db) and 44 (to 6.6 db).

- Large fluorescence spikes were removed from cast 21 and 86.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined. No problems were found.

Profile plots were examined. No problems were found

##### Final CTD File Steps (REMOVE and HEADEDIT)

For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Secondary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T1:C1 and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add comments about processing.

The Standards Check routine was run and no problems were found.

The Header Check was run; no problems were found.

Profile and T-S plots were examined. No problems were found.

The sensor history was updated.

CRUISE SUMMARY – CTD

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 1222 | No | Yes |  |  |
| Calibration Information - 1222 |  |  |  |  |  |  |  |
| Sensor |  |  | Pre-Cruise |  | Post Cruise |  |  |
| Name |  | S/N | Date | Location | Date | Location |  |
| Temperature |  | 5013 | 4 Mar 2022 | Factory |  |  |  |
| Conductivity |  | 4434 | 4 Mar 2022 | Factory |  |  |  |
| Secondary Temp. |  | 5725 | 4 Mar 2022 | Factory |  |  |  |
| Secondary Cond. |  | 4448 | 4 Mar 2022 | Factory |  |  |  |
| SBE 43 DO sensor |  | 4367 | 27 Feb 2023 | Factory |  |  |  |
| Altimeter |  | 76341 | unknown |  |  |  |  |
| Pressure Sensor |  | 1222 | 17 Mar 2022 | Factory |  |  |  |