## Main
# Channels= | 3 |  |  |  | COMPARE Version 3.6.1 | Run at 2024/04/17 15:03:46 | Compare Excel Add-in Version 2.2
| Primary | Secondary |  | Comparison Name
Ref | Pressure | Pressure |  | Pressure - Pressure
Comp 1 | Sal_Bot | Sal0 |  | Sal0 - Sal_Bot
Comp 2 | Sal_Bot | Sal1 |  | Sal1 - Sal_Bot
Average Size: | 1
|  |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary
Primary | Secondary | Secondary | Pressure | Pressure | Pressure | (S-P) | Sal_Bot | Sal0 | Sal0 | (S-P) | Sal_Bot | Sal1 | Sal1 | (S-P)
File Name | File Name | Count | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff
2023-026-0009.mrg | 2023-026-0009.sam | 241 | 11.666 | 11.666 | 0.048819 | -2.0981e-05 | 31.6763 | 31.7592 | 0.0021347 | 0.082893 | 31.6763 | 31.7574 | 0.0018748 | 0.081078
2023-026-0027.mrg | 2023-026-0027.sam | 337 | 4.939 | 5.0843 | 0.10206 | 0.14533 | 32.2375 | 32.2374 | 0.019249 | -9.1553e-05 | 32.2375 | 32.2382 | 0.023162 | 0.00074768
2023-026-0036.mrg | 2023-026-0036.sam | 241 | 999.676 | 999.676 | 0.34257 | 0.00012207 | 34.367 | 34.36 | 0.00027517 | -0.006958 | 34.367 | 34.3591 | 0.00026898 | -0.0079346
2023-026-0036.mrg | 2023-026-0036.sam | 241 | 750.422 | 750.422 | 0.20489 | 0.00042725 | 34.2439 | 34.2384 | 0.00023329 | -0.0054893 | 34.2439 | 34.2373 | 0.00024923 | -0.0066071
2023-026-0036.mrg | 2023-026-0036.sam | 241 | 499.412 | 499.412 | 0.15841 | 0.00021362 | 34.0912 | 34.0854 | 0.00022847 | -0.0057716 | 34.0912 | 34.0844 | 0.00025273 | -0.0068359
2023-026-0047.mrg | 2023-026-0047.sam | 361 | 4.881 | 4.9436 | 0.055251 | 0.062645 | 31.6286 | 31.6203 | 0.0029989 | -0.0082645 | 31.6286 | 31.6195 | 0.0025438 | -0.0090866
2023-026-0058.mrg | 2023-026-0058.sam | 315 | 5.213 | 5.2706 | 0.08801 | 0.057556 | 32.2236 | 32.2118 | 0.0012103 | -0.011814 | 32.2236 | 32.2115 | 0.0013583 | -0.012135
2023-026-0066.mrg | 2023-026-0066.sam | 241 | 1250.46 | 1250.46 | 0.27149 | 0.00012207 | 34.4511 | 34.4461 | 0.00023507 | -0.0050087 | 34.4511 | 34.4449 | 0.00025478 | -0.006218
2023-026-0066.mrg | 2023-026-0066.sam | 241 | 1000.98 | 1000.98 | 0.22043 | 0.00024414 | 34.3643 | 34.3593 | 0.00037026 | -0.0050278 | 34.3643 | 34.3581 | 0.00041323 | -0.0062218
2023-026-0066.mrg | 2023-026-0066.sam | 241 | 749.707 | 749.707 | 0.088588 | 0.00024414 | 34.2457 | 34.2421 | 0.00023353 | -0.0035667 | 34.2457 | 34.2408 | 0.00023937 | -0.0049248
2023-026-0066.mrg | 2023-026-0066.sam | 241 | 500.565 | 500.565 | 0.21684 | -6.1035e-05 | 34.1113 | 34.106 | 0.0003891 | -0.0052834 | 34.1113 | 34.1048 | 0.00039938 | -0.0064735
2023-026-0066.mrg | 2023-026-0066.sam | 241 | 5.457 | 5.4573 | 0.15611 | 0.00032806 | 32.2015 | 32.1947 | 0.00083066 | -0.0068207 | 32.2015 | 32.1937 | 0.0006505 | -0.007782
2023-026-0071.mrg | 2023-026-0071.sam | 241 | 1499.79 | 1499.79 | 0.34311 | 0.00048828 | 34.499 | 34.4936 | 0.0002311 | -0.0053673 | 34.499 | 34.4927 | 0.00025133 | -0.0062904
2023-026-0071.mrg | 2023-026-0071.sam | 241 | 1250.06 | 1250.06 | 0.14201 | 0.00012207 | 34.4503 | 34.4441 | 0.00024556 | -0.0062294 | 34.4503 | 34.443 | 0.00026968 | -0.0073051
2023-026-0071.mrg | 2023-026-0071.sam | 241 | 1000.44 | 1000.44 | 0.2328 | -0.00030518 | 34.3752 | 34.3692 | 0.00042606 | -0.0060043 | 34.3752 | 34.3681 | 0.00041156 | -0.0070496
2023-026-0071.mrg | 2023-026-0071.sam | 241 | 750.726 | 750.726 | 0.15746 | 0.00036621 | 34.2699 | 34.2647 | 0.00020386 | -0.0052414 | 34.2699 | 34.2635 | 0.00023466 | -0.0063705
2023-026-0071.mrg | 2023-026-0071.sam | 241 | 501.063 | 501.063 | 0.12516 | 0 | 34.1056 | 34.1006 | 0.00021715 | -0.0049591 | 34.1056 | 34.0996 | 0.00022741 | -0.006012
2023-026-0071.mrg | 2023-026-0071.sam | 241 | 400.164 | 400.164 | 0.17952 | -0.00024414 | 34.0535 | 34.0484 | 0.00022571 | -0.0051041 | 34.0535 | 34.0474 | 0.00024505 | -0.0061302
2023-026-0071.mrg | 2023-026-0071.sam | 241 | 300.306 | 300.306 | 0.23926 | 0.00039673 | 34.0026 | 33.996 | 0.00033204 | -0.006649 | 34.0026 | 33.9949 | 0.0003874 | -0.0077209
2023-026-0071.mrg | 2023-026-0071.sam | 241 | 250.891 | 250.891 | 0.1162 | -0.00038147 | 33.9419 | 33.9431 | 0.00055836 | 0.0012054 | 33.9419 | 33.9418 | 0.00049475 | -0.00010681
2023-026-0071.mrg | 2023-026-0071.sam | 241 | 200.259 | 200.259 | 0.2497 | 0.00039673 | 33.8886 | 33.8811 | 0.00025031 | -0.007473 | 33.8886 | 33.8802 | 0.00037385 | -0.0084419
2023-026-0071.mrg | 2023-026-0071.sam | 241 | 175.129 | 175.129 | 0.3881 | 0.00047302 | 33.8802 | 33.8745 | 0.00090458 | -0.0057259 | 33.8802 | 33.8736 | 0.0010013 | -0.0065613
2023-026-0071.mrg | 2023-026-0071.sam | 241 | 150.334 | 150.334 | 0.16908 | 0 | 33.8127 | 33.8062 | 0.00038394 | -0.0064964 | 33.8127 | 33.805 | 0.00032338 | -0.007679
2023-026-0071.mrg | 2023-026-0071.sam | 241 | 125.834 | 125.834 | 0.13578 | -0.00022888 | 33.6981 | 33.6902 | 0.00050522 | -0.0079117 | 33.6981 | 33.6893 | 0.0004873 | -0.0087547
2023-026-0071.mrg | 2023-026-0071.sam | 241 | 100.062 | 100.062 | 0.11417 | -0.00038147 | 33.4025 | 33.3938 | 0.0007654 | -0.0087357 | 33.4025 | 33.3931 | 0.00054894 | -0.009388
2023-026-0071.mrg | 2023-026-0071.sam | 241 | 75.477 | 75.4766 | 0.099854 | -0.00036621 | 32.8748 | 32.8583 | 0.00074282 | -0.016472 | 32.8748 | 32.8573 | 0.0010916 | -0.017464
2023-026-0071.mrg | 2023-026-0071.sam | 241 | 50.273 | 50.2732 | 0.13001 | 0.00018311 | 32.3961 | 32.3862 | 0.0002663 | -0.0099335 | 32.3961 | 32.3853 | 0.00029743 | -0.010761
2023-026-0071.mrg | 2023-026-0071.sam | 241 | 30.225 | 30.2248 | 0.1572 | -0.00024033 | 32.3346 | 32.3212 | 0.0022966 | -0.013374 | 32.3346 | 32.3205 | 0.00094629 | -0.014145
2023-026-0071.mrg | 2023-026-0071.sam | 241 | 20.044 | 20.0443 | 0.051117 | 0.00027657 | 32.3057 | 32.2995 | 0.0061859 | -0.0061646 | 32.3057 | 32.2944 | 0.0073216 | -0.011272
2023-026-0071.mrg | 2023-026-0071.sam | 241 | 10.884 | 10.8838 | 0.07337 | -0.00016212 | 32.2923 | 32.2867 | 0.0014092 | -0.0055924 | 32.2923 | 32.2855 | 0.0023495 | -0.0068054
2023-026-0071.mrg | 2023-026-0071.sam | 312 | 4.684 | 4.686 | 0.058243 | 0.0020065 | 32.2701 | 32.2671 | 0.0067263 | -0.0029755 | 32.2701 | 32.2666 | 0.0046113 | -0.0034561
2023-026-0078.mrg | 2023-026-0078.sam | 241 | 750.513 | 750.513 | 0.19713 | 0.00036621 | 34.2321 | 34.228 | 0.00023241 | -0.0040588 | 34.2321 | 34.2271 | 0.00028306 | -0.0050087
2023-026-0078.mrg | 2023-026-0078.sam | 241 | 500.707 | 500.707 | 0.17339 | -6.1035e-05 | 34.0595 | 34.0536 | 0.00059172 | -0.0059013 | 34.0595 | 34.0525 | 0.00055271 | -0.0070343
2023-026-0081.mrg | 2023-026-0081.sam | 241 | 5.291 | 5.291 | 0.10938 | 2.0981e-05 | 32.5602 | 32.562 | 0.0143 | 0.0018082 | 32.5602 | 32.56 | 0.014032 | -0.00023651
2023-026-0098.mrg | 2023-026-0098.sam | 241 | 750.621 | 750.621 | 0.13727 | 0.00042725 | 34.2192 | 34.2157 | 0.00026536 | -0.0034561 | 34.2192 | 34.2145 | 0.00027849 | -0.0046539
2023-026-0098.mrg | 2023-026-0098.sam | 941 | 5.462 | 5.4539 | 0.14141 | -0.0080543 | 32.116 | 32.1111 | 0.00047203 | -0.0048866 | 32.116 | 32.11 | 0.00044047 | -0.0059815
2023-026-0117.mrg | 2023-026-0117.sam | 241 | 1000.03 | 1000.03 | 0.090496 | 0.00042725 | 34.3541 | 34.352 | 0.00028992 | -0.0021172 | 34.3541 | 34.351 | 0.0003092 | -0.0031014
2023-026-0117.mrg | 2023-026-0117.sam | 241 | 750.314 | 750.314 | 0.20408 | 0.00030518 | 34.2151 | 34.2106 | 0.00028172 | -0.004528 | 34.2151 | 34.2097 | 0.00025944 | -0.0053596
2023-026-0117.mrg | 2023-026-0117.sam | 241 | 499.693 | 499.693 | 0.099414 | 0.00018311 | 34.0426 | 34.04 | 0.00023868 | -0.0025749 | 34.0426 | 34.0391 | 0.00022959 | -0.0034599
2023-026-0117.mrg | 2023-026-0117.sam | 241 | 400.239 | 400.239 | 0.20291 | 0.00045776 | 34.0075 | 34.002 | 0.0002181 | -0.0055199 | 34.0075 | 34.001 | 0.00027684 | -0.0064583
2023-026-0117.mrg | 2023-026-0117.sam | 241 | 300.205 | 300.205 | 0.15319 | -0.00021362 | 33.9814 | 33.9769 | 0.00022431 | -0.0044708 | 33.9814 | 33.976 | 0.00023457 | -0.0054131
2023-026-0117.mrg | 2023-026-0117.sam | 241 | 249.8 | 249.8 | 0.22615 | -0.00048828 | 33.9601 | 33.9579 | 0.0011243 | -0.002243 | 33.9601 | 33.9573 | 0.00094731 | -0.0027504
2023-026-0117.mrg | 2023-026-0117.sam | 241 | 200.026 | 200.026 | 0.16878 | -0.00024414 | 33.9217 | 33.9121 | 0.00037352 | -0.0096321 | 33.9217 | 33.9111 | 0.00038194 | -0.01059
2023-026-0117.mrg | 2023-026-0117.sam | 241 | 175.796 | 175.796 | 0.37047 | -0.00050354 | 33.8967 | 33.8896 | 0.00056285 | -0.0071068 | 33.8967 | 33.8888 | 0.00040257 | -0.0079498
2023-026-0117.mrg | 2023-026-0117.sam | 241 | 150.197 | 150.197 | 0.16526 | 4.5776e-05 | 33.849 | 33.8446 | 0.00038875 | -0.0043983 | 33.849 | 33.8437 | 0.00043273 | -0.0053215
2023-026-0117.mrg | 2023-026-0117.sam | 241 | 124.924 | 124.924 | 0.27079 | 0.00038147 | 33.7275 | 33.7168 | 0.0020446 | -0.010666 | 33.7275 | 33.7182 | 0.002569 | -0.0092621
2023-026-0117.mrg | 2023-026-0117.sam | 241 | 99.953 | 99.9529 | 0.16233 | -9.1553e-05 | 33.3861 | 33.3798 | 0.0012324 | -0.0062828 | 33.3861 | 33.3789 | 0.001222 | -0.0071793
2023-026-0117.mrg | 2023-026-0117.sam | 241 | 74.49 | 74.4902 | 0.11763 | 0.00022125 | 32.8102 | 32.8139 | 0.00079395 | 0.003746 | 32.8102 | 32.8135 | 0.00098871 | 0.003315
2023-026-0117.mrg | 2023-026-0117.sam | 241 | 50.118 | 50.1177 | 0.17958 | -0.0002861 | 32.3833 | 32.362 | 0.00059948 | -0.021328 | 32.3833 | 32.361 | 0.00065488 | -0.022293
2023-026-0117.mrg | 2023-026-0117.sam | 241 | 29.995 | 29.9946 | 0.18534 | -0.00043297 | 32.338 | 32.3249 | 0.00034622 | -0.01313 | 32.338 | 32.324 | 0.00031868 | -0.014023
2023-026-0117.mrg | 2023-026-0117.sam | 241 | 20.311 | 20.3109 | 0.13572 | -6.6757e-05 | 32.2887 | 32.2747 | 0.0010996 | -0.014011 | 32.2887 | 32.2734 | 0.00075513 | -0.015293
2023-026-0117.mrg | 2023-026-0117.sam | 241 | 10.147 | 10.1474 | 0.18829 | 0.00038147 | 32.153 | 32.1667 | 0.012004 | 0.013733 | 32.153 | 32.1624 | 0.010007 | 0.0093803
2023-026-0117.mrg | 2023-026-0117.sam | 297 | 5.176 | 5.1436 | 0.078909 | -0.032367 | 32.1228 | 32.107 | 0.00098445 | -0.015846 | 32.1228 | 32.1061 | 0.0008251 | -0.016655
2023-026-0120.mrg | 2023-026-0120.sam | 241 | 1000.47 | 1000.47 | 0.2329 | 0 | 34.3376 | 34.3309 | 0.00028956 | -0.0067062 | 34.3376 | 34.3299 | 0.00033513 | -0.0077477
2023-026-0123.mrg | 2023-026-0123.sam | 241 | 149.923 | 149.923 | 0.22258 | -1.5259e-05 | 33.8345 | 33.8288 | 0.00024579 | -0.0056763 | 33.8345 | 33.8284 | 0.0002484 | -0.0061188
2023-026-0135.mrg | 2023-026-0135.sam | 241 | 1000.4 | 1000.4 | 0.23395 | 0.00018311 | 34.3567 | 34.3488 | 0.00025448 | -0.0078545 | 34.3567 | 34.348 | 0.00024017 | -0.0087395
2023-026-0135.mrg | 2023-026-0135.sam | 241 | 750.108 | 750.108 | 0.34326 | 0.00024414 | 34.2235 | 34.2175 | 0.00025987 | -0.0059815 | 34.2235 | 34.2168 | 0.00026273 | -0.0067482
2023-026-0135.mrg | 2023-026-0135.sam | 241 | 500.364 | 500.364 | 0.2551 | -0.00045776 | 34.0613 | 34.049 | 0.00041188 | -0.012318 | 34.0613 | 34.048 | 0.00038845 | -0.013313
2023-026-0135.mrg | 2023-026-0135.sam | 241 | 400.565 | 400.565 | 0.33373 | 0.00024414 | 34.0259 | 34.0213 | 0.00044524 | -0.0046005 | 34.0259 | 34.0202 | 0.0002972 | -0.0056877
2023-026-0135.mrg | 2023-026-0135.sam | 241 | 300.179 | 300.179 | 0.41185 | -0.00030518 | 33.9317 | 33.9269 | 0.00027663 | -0.0048027 | 33.9317 | 33.9261 | 0.0002769 | -0.0056191
2023-026-0135.mrg | 2023-026-0135.sam | 241 | 250.425 | 250.425 | 0.20542 | 0.00041199 | 33.9212 | 33.9158 | 0.00045813 | -0.0053596 | 33.9212 | 33.915 | 0.00041516 | -0.0062256
2023-026-0135.mrg | 2023-026-0135.sam | 241 | 200.581 | 200.581 | 0.20479 | -7.6294e-05 | 33.8846 | 33.8882 | 0.00067811 | 0.0035934 | 33.8846 | 33.8873 | 0.00075273 | 0.0027313
2023-026-0135.mrg | 2023-026-0135.sam | 241 | 175.276 | 175.276 | 0.33707 | 0.00018311 | 33.8484 | 33.8466 | 0.00093793 | -0.0018272 | 33.8484 | 33.8455 | 0.00063247 | -0.0028763
2023-026-0135.mrg | 2023-026-0135.sam | 241 | 150.856 | 150.856 | 0.38642 | 0.00015259 | 33.8101 | 33.8104 | 0.0017596 | 0.00029755 | 33.8101 | 33.8099 | 0.0016631 | -0.0001564
2023-026-0135.mrg | 2023-026-0135.sam | 241 | 125.379 | 125.379 | 0.47784 | 7.6294e-06 | 33.6058 | 33.6041 | 0.001459 | -0.0017395 | 33.6058 | 33.6033 | 0.0012771 | -0.002533
2023-026-0135.mrg | 2023-026-0135.sam | 241 | 100.221 | 100.221 | 0.36329 | 0.00035858 | 33.1151 | 33.0463 | 0.0031701 | -0.068809 | 33.1151 | 33.0443 | 0.004068 | -0.070789
2023-026-0135.mrg | 2023-026-0135.sam | 241 | 74.845 | 74.8455 | 0.17637 | 0.00049591 | 32.404 | 32.4002 | 0.00063469 | -0.003849 | 32.404 | 32.3995 | 0.00055841 | -0.0044632
2023-026-0135.mrg | 2023-026-0135.sam | 241 | 50.439 | 50.4387 | 0.34308 | -0.00033951 | 32.3151 | 32.3113 | 0.0016072 | -0.0037613 | 32.3151 | 32.3107 | 0.00098834 | -0.0044022
2023-026-0135.mrg | 2023-026-0135.sam | 241 | 30 | 29.9999 | 0.25542 | -0.00010681 | 32.2731 | 32.2547 | 0.008314 | -0.018372 | 32.2731 | 32.2571 | 0.0078733 | -0.015972
2023-026-0135.mrg | 2023-026-0135.sam | 241 | 20.697 | 20.697 | 0.44965 | -3.0518e-05 | 32.0814 | 32.0689 | 0.0095363 | -0.012543 | 32.0814 | 32.0697 | 0.0076251 | -0.011661
2023-026-0135.mrg | 2023-026-0135.sam | 241 | 10.199 | 10.1986 | 0.156 | -0.00043201 | 31.8695 | 31.8572 | 0.0014768 | -0.012337 | 31.8695 | 31.8559 | 0.0015664 | -0.013613
2023-026-0135.mrg | 2023-026-0135.sam | 408 | 5.136 | 5.0495 | 0.25026 | -0.086466 | 31.8679 | 31.8626 | 0.00026149 | -0.0052566 | 31.8679 | 31.8616 | 0.00030624 | -0.0062752
2023-026-0142.mrg | 2023-026-0142.sam | 241 | 1000.63 | 1000.63 | 0.28908 | -0.00018311 | 34.3536 | 34.3479 | 0.00026039 | -0.0056648 | 34.3536 | 34.3471 | 0.0002776 | -0.0065346
2023-026-0142.mrg | 2023-026-0142.sam | 241 | 749.547 | 749.547 | 0.31975 | -0.00036621 | 34.2278 | 34.2219 | 0.00025055 | -0.0058899 | 34.2278 | 34.2211 | 0.00029332 | -0.0067368
2023-026-0142.mrg | 2023-026-0142.sam | 241 | 500.284 | 500.284 | 0.29029 | -6.1035e-05 | 34.0756 | 34.0709 | 0.00026253 | -0.0047035 | 34.0756 | 34.0701 | 0.00027699 | -0.005455
2023-026-0142.mrg | 2023-026-0142.sam | 436 | 5.358 | 5.2877 | 0.2532 | -0.070321 | 32.1341 | 32.1233 | 0.0033439 | -0.010834 | 32.1341 | 32.122 | 0.0029561 | -0.012081
2023-026-0165.mrg | 2023-026-0165.sam | 354 | 4.81 | 4.8691 | 0.39606 | 0.059071 | 31.4464 | 31.4388 | 0.0014692 | -0.0076485 | 31.4464 | 31.4385 | 0.0012721 | -0.0079441
2023-026-0183.mrg | 2023-026-0183.sam | 241 | 351.694 | 351.694 | 0.028863 | 9.1553e-05 | 30.7361 | 30.7464 | 0.00042067 | 0.01034 | 30.7361 | 30.7461 | 0.0004233 | 0.0099716
2023-026-0198.mrg | 2023-026-0198.sam | 241 | 350.153 | 350.153 | 0.03203 | 0.00012207 | 31.2996 | 31.2936 | 0.00019437 | -0.0060196 | 31.2996 | 31.2933 | 0.00023073 | -0.0062943
2023-026-0198.mrg | 2023-026-0198.sam | 241 | 1.866 | 1.8659 | 0.052901 | -0.00012863 | 28.0179 | 28.1991 | 0.09372 | 0.18123 | 28.0179 | 28.1858 | 0.09072 | 0.16787
2023-026-0201.mrg | 2023-026-0201.sam | 241 | 315.259 | 315.259 | 0.03356 | -0.00027466 | 31.1898 | 31.1856 | 0.00027272 | -0.0042057 | 31.1898 | 31.1858 | 0.00024471 | -0.0039844
2023-026-0201.mrg | 2023-026-0201.sam | 241 | 10.341 | 10.341 | 0.028312 | -1.2398e-05 | 28.1954 | 28.0903 | 0.0025939 | -0.10505 | 28.1954 | 28.0885 | 0.0028468 | -0.10692
2023-026-0203.mrg | 2023-026-0203.sam | 241 | 166.969 | 166.969 | 0.026772 | -0.00028992 | 31.0756 | 31.0698 | 0.00025252 | -0.0058289 | 31.0756 | 31.0696 | 0.0004757 | -0.0059528
2023-026-0203.mrg | 2023-026-0203.sam | 241 | 4.978 | 4.9782 | 0.025992 | 0.00022793 | 27.2683 | 27.1433 | 0.0047869 | -0.12496 | 27.2683 | 27.1455 | 0.0063467 | -0.12278
Create Fit Sheet
Plot
Plot
Plot

## File List
Number of Files= |  | 20
Pair # | Primary File | Secondary File | First Row in Main | Last Row in Main | Comparison
1 | 2023-026-0009.mrg | 2023-026-0009.sam | 12 | 12 | 2023-026-0009.mrg & 2023-026-0009.sam
2 | 2023-026-0027.mrg | 2023-026-0027.sam | 14 | 14 | 2023-026-0027.mrg & 2023-026-0027.sam
3 | 2023-026-0036.mrg | 2023-026-0036.sam | 16 | 18 | 2023-026-0036.mrg & 2023-026-0036.sam
4 | 2023-026-0047.mrg | 2023-026-0047.sam | 20 | 20 | 2023-026-0047.mrg & 2023-026-0047.sam
5 | 2023-026-0058.mrg | 2023-026-0058.sam | 22 | 22 | 2023-026-0058.mrg & 2023-026-0058.sam
6 | 2023-026-0066.mrg | 2023-026-0066.sam | 24 | 28 | 2023-026-0066.mrg & 2023-026-0066.sam
7 | 2023-026-0071.mrg | 2023-026-0071.sam | 30 | 48 | 2023-026-0071.mrg & 2023-026-0071.sam
8 | 2023-026-0078.mrg | 2023-026-0078.sam | 50 | 51 | 2023-026-0078.mrg & 2023-026-0078.sam
9 | 2023-026-0081.mrg | 2023-026-0081.sam | 53 | 53 | 2023-026-0081.mrg & 2023-026-0081.sam
10 | 2023-026-0098.mrg | 2023-026-0098.sam | 55 | 56 | 2023-026-0098.mrg & 2023-026-0098.sam
11 | 2023-026-0117.mrg | 2023-026-0117.sam | 58 | 74 | 2023-026-0117.mrg & 2023-026-0117.sam
12 | 2023-026-0120.mrg | 2023-026-0120.sam | 76 | 76 | 2023-026-0120.mrg & 2023-026-0120.sam
13 | 2023-026-0123.mrg | 2023-026-0123.sam | 78 | 78 | 2023-026-0123.mrg & 2023-026-0123.sam
14 | 2023-026-0135.mrg | 2023-026-0135.sam | 80 | 96 | 2023-026-0135.mrg & 2023-026-0135.sam
15 | 2023-026-0142.mrg | 2023-026-0142.sam | 98 | 101 | 2023-026-0142.mrg & 2023-026-0142.sam
16 | 2023-026-0165.mrg | 2023-026-0165.sam | 103 | 103 | 2023-026-0165.mrg & 2023-026-0165.sam
17 | 2023-026-0183.mrg | 2023-026-0183.sam | 105 | 105 | 2023-026-0183.mrg & 2023-026-0183.sam
18 | 2023-026-0198.mrg | 2023-026-0198.sam | 107 | 108 | 2023-026-0198.mrg & 2023-026-0198.sam
19 | 2023-026-0201.mrg | 2023-026-0201.sam | 110 | 111 | 2023-026-0201.mrg & 2023-026-0201.sam
20 | 2023-026-0203.mrg | 2023-026-0203.sam | 113 | 114 | 2023-026-0203.mrg & 2023-026-0203.sam

## Fit Index
Index to Fit Worksheets
Fit# | Description | Versus | Number of File Pairs
1 | (Sal0 - Sal_Bot) | Pressure | 20
2 | (Sal1 - Sal_Bot) | Pressure | 20
3 | (Sal0 - Sal_Bot) | File Pair # | 20
4 | (Sal1 - Sal_Bot) | File Pair # | 20
5 | (Sal0 - Sal_Bot) | File Pair # | 20
6 | (Sal1 - Sal_Bot) | File Pair # | 20

## Fit_1
Calculation of Linear Least Squares Fit For
(Sal0 - Sal_Bot)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal0 - Sal_Bot), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.030604 |  | X | CALC Y
m= | -6.6465e-07 | 6.8296e-07 |  | Std Dev of Resid |  | 0.0012124 |  | 300.179 | -0.0049051
b= | -0.0047056 | 0.0005312 |  | Resid Sum of Squares |  | 4.4096e-05 |  | 1499.79 | -0.0057025
# of Values Extracted = |  |  | 84
# of Values Used in Fit = |  |  | 32
File | Use | Pressure | Sal0 | Sal0 | Diff (Sal0 - Sal_Bot) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 11.666 | 31.7592 | 0.0021347 | 0.082893 | -0.0047134 | -0.087607 |  | 999.676 | -0.006958
2 | NO | 5.0843 | 32.2374 | 0.019249 | -9.1553e-05 | -0.004709 | -0.0046174 |  | 750.422 | -0.0054893
3 | YES | 999.676 | 34.36 | 0.00027517 | -0.006958 | -0.0053701 | 0.001588 |  | 499.412 | -0.0057716
3 | YES | 750.422 | 34.2384 | 0.00023329 | -0.0054893 | -0.0052044 | 0.00028497 |  | 1250.46 | -0.0050087
3 | YES | 499.412 | 34.0854 | 0.00022847 | -0.0057716 | -0.0050375 | 0.00073409 |  | 1000.98 | -0.0050278
4 | NO | 4.9436 | 31.6203 | 0.0029989 | -0.0082645 | -0.0047089 | 0.0035556 |  | 749.707 | -0.0035667
5 | NO | 5.2706 | 32.2118 | 0.0012103 | -0.011814 | -0.0047091 | 0.007105 |  | 500.565 | -0.0052834
6 | YES | 1250.46 | 34.4461 | 0.00023507 | -0.0050087 | -0.0055367 | -0.00052803 |  | 1499.79 | -0.0053673
6 | YES | 1000.98 | 34.3593 | 0.00037026 | -0.0050278 | -0.0053709 | -0.00034315 |  | 1250.06 | -0.0062294
6 | YES | 749.707 | 34.2421 | 0.00023353 | -0.0035667 | -0.0052039 | -0.0016372 |  | 1000.44 | -0.0060043
6 | YES | 500.565 | 34.106 | 0.0003891 | -0.0052834 | -0.0050383 | 0.00024505 |  | 750.726 | -0.0052414
6 | NO | 5.4573 | 32.1947 | 0.00083066 | -0.0068207 | -0.0047092 | 0.0021114 |  | 501.063 | -0.0049591
7 | YES | 1499.79 | 34.4936 | 0.0002311 | -0.0053673 | -0.0057025 | -0.00033517 |  | 400.164 | -0.0051041
7 | YES | 1250.06 | 34.4441 | 0.00024556 | -0.0062294 | -0.0055365 | 0.00069293 |  | 300.306 | -0.006649
7 | YES | 1000.44 | 34.3692 | 0.00042606 | -0.0060043 | -0.0053706 | 0.00063377 |  | 750.513 | -0.0040588
7 | YES | 750.726 | 34.2647 | 0.00020386 | -0.0052414 | -0.0052046 | 3.6806e-05 |  | 500.707 | -0.0059013
7 | YES | 501.063 | 34.1006 | 0.00021715 | -0.0049591 | -0.0050386 | -7.9535e-05 |  | 750.621 | -0.0034561
7 | YES | 400.164 | 34.0484 | 0.00022571 | -0.0051041 | -0.0049716 | 0.00013248 |  | 1000.03 | -0.0021172
7 | YES | 300.306 | 33.996 | 0.00033204 | -0.006649 | -0.0049052 | 0.0017438 |  | 750.314 | -0.004528
7 | NO | 250.891 | 33.9431 | 0.00055836 | 0.0012054 | -0.0048724 | -0.0060778 |  | 499.693 | -0.0025749
7 | NO | 200.259 | 33.8811 | 0.00025031 | -0.007473 | -0.0048387 | 0.0026343 |  | 400.239 | -0.0055199
7 | NO | 175.129 | 33.8745 | 0.00090458 | -0.0057259 | -0.004822 | 0.00090385 |  | 300.205 | -0.0044708
7 | NO | 150.334 | 33.8062 | 0.00038394 | -0.0064964 | -0.0048055 | 0.0016909 |  | 1000.47 | -0.0067062
7 | NO | 125.834 | 33.6902 | 0.00050522 | -0.0079117 | -0.0047892 | 0.0031224 |  | 1000.4 | -0.0078545
7 | NO | 100.062 | 33.3938 | 0.0007654 | -0.0087357 | -0.0047721 | 0.0039635 |  | 750.108 | -0.0059815
7 | NO | 75.4766 | 32.8583 | 0.00074282 | -0.016472 | -0.0047558 | 0.011716 |  | 400.565 | -0.0046005
7 | NO | 50.2732 | 32.3862 | 0.0002663 | -0.0099335 | -0.004739 | 0.0051944 |  | 300.179 | -0.0048027
7 | NO | 30.2248 | 32.3212 | 0.0022966 | -0.013374 | -0.0047257 | 0.0086486 |  | 1000.63 | -0.0056648
7 | NO | 20.0443 | 32.2995 | 0.0061859 | -0.0061646 | -0.0047189 | 0.0014456 |  | 749.547 | -0.0058899
7 | NO | 10.8838 | 32.2867 | 0.0014092 | -0.0055924 | -0.0047128 | 0.0008795 |  | 500.284 | -0.0047035
7 | NO | 4.686 | 32.2671 | 0.0067263 | -0.0029755 | -0.0047087 | -0.0017333 |  | 350.153 | -0.0060196
8 | YES | 750.513 | 34.228 | 0.00023241 | -0.0040588 | -0.0052044 | -0.0011456 |  | 315.259 | -0.0042057 | -0.0051786 | 0.0012113
8 | YES | 500.707 | 34.0536 | 0.00059172 | -0.0059013 | -0.0050384 | 0.00086293
9 | NO | 5.291 | 32.562 | 0.0143 | 0.0018082 | -0.0047091 | -0.0065173
10 | YES | 750.621 | 34.2157 | 0.00026536 | -0.0034561 | -0.0052045 | -0.0017484
10 | NO | 5.4539 | 32.1111 | 0.00047203 | -0.0048866 | -0.0047092 | 0.00017739
11 | YES | 1000.03 | 34.352 | 0.00028992 | -0.0021172 | -0.0053703 | -0.0032531
11 | YES | 750.314 | 34.2106 | 0.00028172 | -0.004528 | -0.0052043 | -0.00067626
11 | YES | 499.693 | 34.04 | 0.00023868 | -0.0025749 | -0.0050377 | -0.0024628 |  |  |  | x
11 | YES | 400.239 | 34.002 | 0.0002181 | -0.0055199 | -0.0049716 | 0.00054824
11 | YES | 300.205 | 33.9769 | 0.00022431 | -0.0044708 | -0.0049051 | -0.00043431
11 | NO | 249.8 | 33.9579 | 0.0011243 | -0.002243 | -0.0048716 | -0.0026286
11 | NO | 200.026 | 33.9121 | 0.00037352 | -0.0096321 | -0.0048386 | 0.0047936
11 | NO | 175.796 | 33.8896 | 0.00056285 | -0.0071068 | -0.0048225 | 0.0022843
11 | NO | 150.197 | 33.8446 | 0.00038875 | -0.0043983 | -0.0048054 | -0.00040709
11 | NO | 124.924 | 33.7168 | 0.0020446 | -0.010666 | -0.0047886 | 0.0058773
11 | NO | 99.9529 | 33.3798 | 0.0012324 | -0.0062828 | -0.004772 | 0.0015108
11 | NO | 74.4902 | 32.8139 | 0.00079395 | 0.003746 | -0.0047551 | -0.0085012
11 | NO | 50.1177 | 32.362 | 0.00059948 | -0.021328 | -0.0047389 | 0.016589
11 | NO | 29.9946 | 32.3249 | 0.00034622 | -0.01313 | -0.0047255 | 0.0084047
11 | NO | 20.3109 | 32.2747 | 0.0010996 | -0.014011 | -0.0047191 | 0.0092923
11 | NO | 10.1474 | 32.1667 | 0.012004 | 0.013733 | -0.0047124 | -0.018445
11 | NO | 5.1436 | 32.107 | 0.00098445 | -0.015846 | -0.004709 | 0.011137
12 | YES | 1000.47 | 34.3309 | 0.00028956 | -0.0067062 | -0.0053706 | 0.0013357
13 | NO | 149.923 | 33.8288 | 0.00024579 | -0.0056763 | -0.0048053 | 0.00087101
14 | YES | 1000.4 | 34.3488 | 0.00025448 | -0.0078545 | -0.0053705 | 0.0024839
14 | YES | 750.108 | 34.2175 | 0.00025987 | -0.0059815 | -0.0052042 | 0.00077728
14 | NO | 500.364 | 34.049 | 0.00041188 | -0.012318 | -0.0050382 | 0.0072795 |  |  |  | x
14 | YES | 400.565 | 34.0213 | 0.00044524 | -0.0046005 | -0.0049718 | -0.00037133
14 | YES | 300.179 | 33.9269 | 0.00027663 | -0.0048027 | -0.0049051 | -0.00010243
14 | NO | 250.425 | 33.9158 | 0.00045813 | -0.0053596 | -0.0048721 | 0.00048759
14 | NO | 200.581 | 33.8882 | 0.00067811 | 0.0035934 | -0.0048389 | -0.0084324
14 | NO | 175.276 | 33.8466 | 0.00093793 | -0.0018272 | -0.0048221 | -0.0029949
14 | NO | 150.856 | 33.8104 | 0.0017596 | 0.00029755 | -0.0048059 | -0.0051034
14 | NO | 125.379 | 33.6041 | 0.001459 | -0.0017395 | -0.0047889 | -0.0030494
14 | NO | 100.221 | 33.0463 | 0.0031701 | -0.068809 | -0.0047722 | 0.064037
14 | NO | 74.8455 | 32.4002 | 0.00063469 | -0.003849 | -0.0047554 | -0.00090633
14 | NO | 50.4387 | 32.3113 | 0.0016072 | -0.0037613 | -0.0047391 | -0.00097785
14 | NO | 29.9999 | 32.2547 | 0.008314 | -0.018372 | -0.0047256 | 0.013646
14 | NO | 20.697 | 32.0689 | 0.0095363 | -0.012543 | -0.0047194 | 0.0078233
14 | NO | 10.1986 | 31.8572 | 0.0014768 | -0.012337 | -0.0047124 | 0.0076243
14 | NO | 5.0495 | 31.8626 | 0.00026149 | -0.0052566 | -0.004709 | 0.00054768
15 | YES | 1000.63 | 34.3479 | 0.00026039 | -0.0056648 | -0.0053707 | 0.00029415
15 | YES | 749.547 | 34.2219 | 0.00025055 | -0.0058899 | -0.0052038 | 0.00068609
15 | YES | 500.284 | 34.0709 | 0.00026253 | -0.0047035 | -0.0050381 | -0.00033461
15 | NO | 5.2877 | 32.1233 | 0.0033439 | -0.010834 | -0.0047091 | 0.0061246
16 | NO | 4.8691 | 31.4388 | 0.0014692 | -0.0076485 | -0.0047088 | 0.0029396 | [5 empty] | Bottle said to be from 350db has salinity 30.7361psu
17 | NO | 351.694 | 30.7464 | 0.00042067 | 0.01034 | -0.0049394 | -0.015279 | 30.7361 |  |  | x |  | CTD salinity at 351.7db is 30.7464 (primary) and 30.7461 (secondary
18 | YES | 350.153 | 31.2936 | 0.00019437 | -0.0060196 | -0.0049383 | 0.0010812 | [5 empty] | CTD salinity at 300db is 30.73831 and 30.7382psu
18 | NO | 1.8659 | 28.1991 | 0.09372 | 0.18123 | -0.0047069 | -0.18593 | [5 empty] | So this bottle appears more likely to be from 300db as CTD salinity is higher by 0.002.
19 | YES | 315.259 | 31.1856 | 0.00027272 | -0.0042057 | -0.0049151 | -0.00070945 |  |  |  | x |  | However, even at 300db it would still be an outlier, so little is achieved by reasigning it to that level.
19 | NO | 10.341 | 28.0903 | 0.0025939 | -0.10505 | -0.0047125 | 0.10034
20 | NO | 166.969 | 31.0698 | 0.00025252 | -0.0058289 | -0.0048166 | 0.0010123
20 | NO | 4.9782 | 27.1433 | 0.0047869 | -0.12496 | -0.0047089 | 0.12025
NO
NO
YES
YES
YES
NO
NO
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
NO
YES
NO
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
NO
YES
YES
NO
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
NO
NO
NO
YES
NO
YES
NO
NO
NO
FIT vs. Pressure
FIT vs. Sal0
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_2
Calculation of Linear Least Squares Fit For
(Sal1 - Sal_Bot)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal1 - Sal_Bot), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.058539 |  | X | CALC Y
m= | -9.4506e-07 | 6.9195e-07 |  | Std Dev of Resid |  | 0.0012283 |  | 300.179 | -0.0057402
b= | -0.0054565 | 0.00053819 |  | Resid Sum of Squares |  | 4.5264e-05 |  | 1499.79 | -0.0068739
# of Values Extracted = |  |  | 84
# of Values Used in Fit = |  |  | 32
File | Use | Pressure | Sal1 | Sal1 | Diff (Sal1 - Sal_Bot) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 11.666 | 31.7574 | 0.0018748 | 0.081078 | -0.0054675 | -0.086545 |  | 999.676 | -0.0079346
2 | NO | 5.0843 | 32.2382 | 0.023162 | 0.00074768 | -0.0054613 | -0.006209 |  | 750.422 | -0.0066071
3 | YES | 999.676 | 34.3591 | 0.00026898 | -0.0079346 | -0.0064013 | 0.0015333 |  | 499.412 | -0.0068359
3 | YES | 750.422 | 34.2373 | 0.00024923 | -0.0066071 | -0.0061657 | 0.00044136 |  | 1250.46 | -0.006218
3 | YES | 499.412 | 34.0844 | 0.00025273 | -0.0068359 | -0.0059285 | 0.00090746 |  | 1000.98 | -0.0062218
4 | NO | 4.9436 | 31.6195 | 0.0025438 | -0.0090866 | -0.0054612 | 0.0036254 |  | 749.707 | -0.0049248
5 | NO | 5.2706 | 32.2115 | 0.0013583 | -0.012135 | -0.0054615 | 0.0066731 |  | 500.565 | -0.0064735
6 | YES | 1250.46 | 34.4449 | 0.00025478 | -0.006218 | -0.0066383 | -0.00042031 |  | 1499.79 | -0.0062904
6 | YES | 1000.98 | 34.3581 | 0.00041323 | -0.0062218 | -0.0064025 | -0.00018072 |  | 1250.06 | -0.0073051
6 | YES | 749.707 | 34.2408 | 0.00023937 | -0.0049248 | -0.006165 | -0.0012403 |  | 1000.44 | -0.0070496
6 | YES | 500.565 | 34.1048 | 0.00039938 | -0.0064735 | -0.0059296 | 0.00054397 |  | 750.726 | -0.0063705
6 | NO | 5.4573 | 32.1937 | 0.0006505 | -0.007782 | -0.0054617 | 0.0023203 |  | 501.063 | -0.006012
7 | YES | 1499.79 | 34.4927 | 0.00025133 | -0.0062904 | -0.0068739 | -0.00058346 |  | 400.164 | -0.0061302
7 | YES | 1250.06 | 34.443 | 0.00026968 | -0.0073051 | -0.0066379 | 0.00066726 |  | 300.306 | -0.0077209
7 | YES | 1000.44 | 34.3681 | 0.00041156 | -0.0070496 | -0.006402 | 0.00064758 |  | 750.513 | -0.0050087
7 | YES | 750.726 | 34.2635 | 0.00023466 | -0.0063705 | -0.006166 | 0.00020455 |  | 500.707 | -0.0070343
7 | YES | 501.063 | 34.0996 | 0.00022741 | -0.006012 | -0.00593 | 8.1916e-05 |  | 750.621 | -0.0046539
7 | YES | 400.164 | 34.0474 | 0.00024505 | -0.0061302 | -0.0058347 | 0.00029553 |  | 1000.03 | -0.0031014
7 | YES | 300.306 | 33.9949 | 0.0003874 | -0.0077209 | -0.0057403 | 0.0019806 |  | 750.314 | -0.0053596
7 | NO | 250.891 | 33.9418 | 0.00049475 | -0.00010681 | -0.0056936 | -0.0055868 |  | 499.693 | -0.0034599
7 | NO | 200.259 | 33.8802 | 0.00037385 | -0.0084419 | -0.0056458 | 0.0027962 |  | 400.239 | -0.0064583
7 | NO | 175.129 | 33.8736 | 0.0010013 | -0.0065613 | -0.005622 | 0.00093926 |  | 300.205 | -0.0054131
7 | NO | 150.334 | 33.805 | 0.00032338 | -0.007679 | -0.0055986 | 0.0020804 |  | 1000.47 | -0.0077477
7 | NO | 125.834 | 33.6893 | 0.0004873 | -0.0087547 | -0.0055754 | 0.0031793 |  | 1000.4 | -0.0087395
7 | NO | 100.062 | 33.3931 | 0.00054894 | -0.009388 | -0.0055511 | 0.0038369 |  | 750.108 | -0.0067482
7 | NO | 75.4766 | 32.8573 | 0.0010916 | -0.017464 | -0.0055278 | 0.011936 |  | 400.565 | -0.0056877
7 | NO | 50.2732 | 32.3853 | 0.00029743 | -0.010761 | -0.005504 | 0.0052573 |  | 300.179 | -0.0056191
7 | NO | 30.2248 | 32.3205 | 0.00094629 | -0.014145 | -0.0054851 | 0.0086598 |  | 1000.63 | -0.0065346
7 | NO | 20.0443 | 32.2944 | 0.0073216 | -0.011272 | -0.0054755 | 0.0057969 |  | 749.547 | -0.0067368
7 | NO | 10.8838 | 32.2855 | 0.0023495 | -0.0068054 | -0.0054668 | 0.0013386 |  | 500.284 | -0.005455
7 | NO | 4.686 | 32.2666 | 0.0046113 | -0.0034561 | -0.0054609 | -0.0020048 |  | 350.153 | -0.0062943
8 | YES | 750.513 | 34.2271 | 0.00028306 | -0.0050087 | -0.0061658 | -0.0011571 |  | 315.259 | -0.0039844 | -0.0061291 | 0.0012454
8 | YES | 500.707 | 34.0525 | 0.00055271 | -0.0070343 | -0.0059297 | 0.0011046
9 | NO | 5.291 | 32.56 | 0.014032 | -0.00023651 | -0.0054615 | -0.005225
10 | YES | 750.621 | 34.2145 | 0.00027849 | -0.0046539 | -0.0061659 | -0.001512
10 | NO | 5.4539 | 32.11 | 0.00044047 | -0.0059815 | -0.0054617 | 0.00051978
11 | YES | 1000.03 | 34.351 | 0.0003092 | -0.0031014 | -0.0064016 | -0.0033002
11 | YES | 750.314 | 34.2097 | 0.00025944 | -0.0053596 | -0.0061656 | -0.00080595
11 | YES | 499.693 | 34.0391 | 0.00022959 | -0.0034599 | -0.0059287 | -0.0024688
11 | YES | 400.239 | 34.001 | 0.00027684 | -0.0064583 | -0.0058348 | 0.00062352
11 | YES | 300.205 | 33.976 | 0.00023457 | -0.0054131 | -0.0057402 | -0.00032716
11 | NO | 249.8 | 33.9573 | 0.00094731 | -0.0027504 | -0.0056926 | -0.0029422
11 | NO | 200.026 | 33.9111 | 0.00038194 | -0.01059 | -0.0056455 | 0.0049441
11 | NO | 175.796 | 33.8888 | 0.00040257 | -0.0079498 | -0.0056226 | 0.0023272
11 | NO | 150.197 | 33.8437 | 0.00043273 | -0.0053215 | -0.0055985 | -0.00027696
11 | NO | 124.924 | 33.7182 | 0.002569 | -0.0092621 | -0.0055746 | 0.0036875
11 | NO | 99.9529 | 33.3789 | 0.001222 | -0.0071793 | -0.005551 | 0.0016283
11 | NO | 74.4902 | 32.8135 | 0.00098871 | 0.003315 | -0.0055269 | -0.0088419
11 | NO | 50.1177 | 32.361 | 0.00065488 | -0.022293 | -0.0055039 | 0.016789
11 | NO | 29.9946 | 32.324 | 0.00031868 | -0.014023 | -0.0054849 | 0.0085379
11 | NO | 20.3109 | 32.2734 | 0.00075513 | -0.015293 | -0.0054757 | 0.0098174
11 | NO | 10.1474 | 32.1624 | 0.010007 | 0.0093803 | -0.0054661 | -0.014846
11 | NO | 5.1436 | 32.1061 | 0.0008251 | -0.016655 | -0.0054614 | 0.011194
12 | YES | 1000.47 | 34.3299 | 0.00033513 | -0.0077477 | -0.006402 | 0.0013456
13 | NO | 149.923 | 33.8284 | 0.0002484 | -0.0061188 | -0.0055982 | 0.00052057
14 | YES | 1000.4 | 34.348 | 0.00024017 | -0.0087395 | -0.0064019 | 0.0023375
14 | YES | 750.108 | 34.2168 | 0.00026273 | -0.0067482 | -0.0061654 | 0.00058279
14 | NO | 500.364 | 34.048 | 0.00038845 | -0.013313 | -0.0059294 | 0.0073839 | 34.0613
14 | YES | 400.565 | 34.0202 | 0.0002972 | -0.0056877 | -0.0058351 | -0.00014736 | 34.0259
14 | YES | 300.179 | 33.9261 | 0.0002769 | -0.0056191 | -0.0057402 | -0.00012115 | 33.9317
14 | NO | 250.425 | 33.915 | 0.00041516 | -0.0062256 | -0.0056932 | 0.00053241 | 33.9212
14 | NO | 200.581 | 33.8873 | 0.00075273 | 0.0027313 | -0.0056461 | -0.0083774 | 33.8846
14 | NO | 175.276 | 33.8455 | 0.00063247 | -0.0028763 | -0.0056222 | -0.0027459 | 33.8484
14 | NO | 150.856 | 33.8099 | 0.0016631 | -0.0001564 | -0.0055991 | -0.0054427 | 33.8101
14 | NO | 125.379 | 33.6033 | 0.0012771 | -0.002533 | -0.005575 | -0.003042 | 33.6058
14 | NO | 100.221 | 33.0443 | 0.004068 | -0.070789 | -0.0055512 | 0.065238 | 33.1151
14 | NO | 74.8455 | 32.3995 | 0.00055841 | -0.0044632 | -0.0055272 | -0.001064 | 32.404
14 | NO | 50.4387 | 32.3107 | 0.00098834 | -0.0044022 | -0.0055042 | -0.001102 | 32.3151
14 | NO | 29.9999 | 32.2571 | 0.0078733 | -0.015972 | -0.0054849 | 0.010487 | 32.2731
14 | NO | 20.697 | 32.0697 | 0.0076251 | -0.011661 | -0.0054761 | 0.0061854 | 32.0814
14 | NO | 10.1986 | 31.8559 | 0.0015664 | -0.013613 | -0.0054661 | 0.0081466 | 31.8695
14 | NO | 5.0495 | 31.8616 | 0.00030624 | -0.0062752 | -0.0054613 | 0.0008139 | 31.8679
15 | YES | 1000.63 | 34.3471 | 0.0002776 | -0.0065346 | -0.0064022 | 0.00013242 | 34.3536
15 | YES | 749.547 | 34.2211 | 0.00029332 | -0.0067368 | -0.0061649 | 0.00057188 | 34.2278
15 | YES | 500.284 | 34.0701 | 0.00027699 | -0.005455 | -0.0059293 | -0.00047429 | 34.0756
15 | NO | 5.2877 | 32.122 | 0.0029561 | -0.012081 | -0.0054615 | 0.0066196 | 32.1341
16 | NO | 4.8691 | 31.4385 | 0.0012721 | -0.0079441 | -0.0054611 | 0.002483 | 31.4464
17 | NO | 351.694 | 30.7461 | 0.0004233 | 0.0099716 | -0.0057889 | -0.015761 | 30.7361
18 | YES | 350.153 | 31.2933 | 0.00023073 | -0.0062943 | -0.0057874 | 0.00050682
18 | NO | 1.8659 | 28.1858 | 0.09072 | 0.16787 | -0.0054583 | -0.17333
19 | YES | 315.259 | 31.1858 | 0.00024471 | -0.0039844 | -0.0057544 | -0.00177
19 | NO | 10.341 | 28.0885 | 0.0028468 | -0.10692 | -0.0054663 | 0.10145
20 | NO | 166.969 | 31.0696 | 0.0004757 | -0.0059528 | -0.0056143 | 0.00033853
20 | NO | 4.9782 | 27.1455 | 0.0063467 | -0.12278 | -0.0054612 | 0.11732
NO
NO
YES
YES
YES
NO
NO
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
NO
YES
NO
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
NO
YES
YES
NO
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
NO
NO
NO
YES
NO
YES
NO
NO
NO
FIT vs. Pressure
FIT vs. Sal1
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_3
Calculation of Linear Least Squares Fit For
(Sal0 - Sal_Bot)
3
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal0 - Sal_Bot), Prop=File Pair #
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.0060224 |  | X | CALC Y
m= | 2.1856e-05 | 5.1265e-05 |  | Std Dev of Resid |  | 0.0012277 |  | 3 | -0.0053268 |  |  | 2023-088 Corrections
b= | -0.0053924 | 0.00054638 |  | Resid Sum of Squares |  | 4.5214e-05 |  | 19 | -0.0049771 |  |  |  | Salinity:T0:C0_Corrected = Salinity:T0 :C0 + 0.00003 * Event_Number +0.0024
Salinity:T1:C1_Corrected = Salinity:T1 :C1 + 0.00008 * Event_Number +0.0009
# of Values Extracted = |  |  | 84
# of Values Used in Fit = |  |  | 32 | [9 empty] | Trendline Sal0 |  | 2*10-5 * file pair number - 0.0054
[13 empty] | Trendline Sal1 |  | 7*10-5 * file pair number - 0.0068
File | Use | Pressure | Sal0 | Sal0 | Diff (Sal0 - Sal_Bot) |  |  | Sal_Bot | Fit vs. | File Pair # | [6 empty] | Corrections if 2023-088 drift continued
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y | Sal0 | file pair | Sal0-bot | Sal1-bot | Sal1-Sal | Event# | Sal0-Salbot | Sal1-Salbot
1 | NO | 11.666 | 31.7592 | 0.0021347 | 0.082893 | -0.0053706 | -0.088264 | 31.6763 | 3 | -0.006958 |  | 1 | -0.00538 | -0.00673 | -0.00135 | 9 | 0.00487 | 0.00842
2 | NO | 5.0843 | 32.2374 | 0.019249 | -9.1553e-05 | -0.0053487 | -0.0052571 | 32.2375 | 3 | -0.0054893 |  | 2 | -0.00536 | -0.00666 | -0.0013 | 27 | 0.00541 | 0.00986
3 | YES | 999.676 | 34.36 | 0.00027517 | -0.006958 | -0.0053268 | 0.0016312 | 34.367 | 3 | -0.0057716 |  | 3 | -0.00534 | -0.00659 | -0.00125 | 36 | 0.00568 | 0.01058
3 | YES | 750.422 | 34.2384 | 0.00023329 | -0.0054893 | -0.0053268 | 0.00016251 | 34.2439 | 6 | -0.0050087 |  | 4 | -0.00532 | -0.00652 | -0.0012 | 47 | 0.00601 | 0.01146
3 | YES | 499.412 | 34.0854 | 0.00022847 | -0.0057716 | -0.0053268 | 0.0004448 | 34.0912 | 6 | -0.0050278 |  | 5 | -0.0053 | -0.00645 | -0.00115 | 58 | 0.00634 | 0.01234
4 | NO | 4.9436 | 31.6203 | 0.0029989 | -0.0082645 | -0.005305 | 0.0029596 | 31.6286 | 6 | -0.0035667 |  | 6 | -0.00528 | -0.00638 | -0.0011 | 66 | 0.00658 | 0.01298
5 | NO | 5.2706 | 32.2118 | 0.0012103 | -0.011814 | -0.0052831 | 0.006531 | 32.2236 | 6 | -0.0052834 |  | 7 | -0.00526 | -0.00631 | -0.00105 | 71 | 0.00673 | 0.01338
6 | YES | 1250.46 | 34.4461 | 0.00023507 | -0.0050087 | -0.0052613 | -0.00025258 | 34.4511 | 7 | -0.0053673 |  | 8 | -0.00524 | -0.00624 | -0.001 | 78 | 0.00694 | 0.01394
6 | YES | 1000.98 | 34.3593 | 0.00037026 | -0.0050278 | -0.0052613 | -0.00023351 | 34.3643 | 7 | -0.0062294 |  | 9 | -0.00522 | -0.00617 | -0.00095 | 81 | 0.00703 | 0.01418
6 | YES | 749.707 | 34.2421 | 0.00023353 | -0.0035667 | -0.0052613 | -0.0016945 | 34.2457 | 7 | -0.0060043 |  | 10 | -0.0052 | -0.0061 | -0.0009 | 98 | 0.00754 | 0.01554
6 | YES | 500.565 | 34.106 | 0.0003891 | -0.0052834 | -0.0052613 | 2.2085e-05 | 34.1113 | 7 | -0.0052414 |  | 11 | -0.00518 | -0.00603 | -0.00085 | 117 | 0.00811 | 0.01706
6 | NO | 5.4573 | 32.1947 | 0.00083066 | -0.0068207 | -0.0052613 | 0.0015594 | 32.2015 | 7 | -0.0049591 |  | 12 | -0.00516 | -0.00596 | -0.0008 | 120 | 0.0082 | 0.0173
7 | YES | 1499.79 | 34.4936 | 0.0002311 | -0.0053673 | -0.0052394 | 0.00012786 | 34.499 | 7 | -0.0051041 |  | 13 | -0.00514 | -0.00589 | -0.00075 | 123 | 0.00829 | 0.01754
7 | YES | 1250.06 | 34.4441 | 0.00024556 | -0.0062294 | -0.0052394 | 0.00098998 | 34.4503 | 7 | -0.006649 |  | 14 | -0.00512 | -0.00582 | -0.0007 | 135 | 0.00865 | 0.0185
7 | YES | 1000.44 | 34.3692 | 0.00042606 | -0.0060043 | -0.0052394 | 0.00076491 | 34.3752 | 8 | -0.0040588 |  | 15 | -0.0051 | -0.00575 | -0.00065 | 142 | 0.00886 | 0.01906
7 | YES | 750.726 | 34.2647 | 0.00020386 | -0.0052414 | -0.0052394 | 1.971e-06 | 34.2699 | 8 | -0.0059013 |  | 16 | -0.00508 | -0.00568 | -0.0006 | 165 | 0.00955 | 0.0209
7 | YES | 501.063 | 34.1006 | 0.00021715 | -0.0049591 | -0.0052394 | -0.00028031 | 34.1056 | 10 | -0.0034561 |  | 17 | -0.00506 | -0.00561 | -0.00055 | 183 | 0.01009 | 0.02234
7 | YES | 400.164 | 34.0484 | 0.00022571 | -0.0051041 | -0.0052394 | -0.00013536 | 34.0535 | 11 | -0.0021172 |  | 18 | -0.00504 | -0.00554 | -0.0005 | 198 | 0.01054 | 0.02354
7 | YES | 300.306 | 33.996 | 0.00033204 | -0.006649 | -0.0052394 | 0.0014096 | 34.0026 | 11 | -0.004528 |  | 19 | -0.00502 | -0.00547 | -0.00045 | 201 | 0.01063 | 0.02378
7 | NO | 250.891 | 33.9431 | 0.00055836 | 0.0012054 | -0.0052394 | -0.0064449 | 33.9419 | 11 | -0.0025749 |  | 20 | -0.005 | -0.0054 | -0.0004 | 203 | 0.01069 | 0.02394
7 | NO | 200.259 | 33.8811 | 0.00025031 | -0.007473 | -0.0052394 | 0.0022336 | 33.8886 | 11 | -0.0055199
7 | NO | 175.129 | 33.8745 | 0.00090458 | -0.0057259 | -0.0052394 | 0.00048644 | 33.8802 | 11 | -0.0044708 |  |  | Trendline Sal0 |  | equals -5*10-5 * file pair number - 0.0049
7 | NO | 150.334 | 33.8062 | 0.00038394 | -0.0064964 | -0.0052394 | 0.001257 | 33.8127 | 12 | -0.0067062 |  |  | Trendline Sal1 |  | equals -1*10-5 * file pair number - 0.0002
7 | NO | 125.834 | 33.6902 | 0.00050522 | -0.0079117 | -0.0052394 | 0.0026723 | 33.6981 | 14 | -0.0078545 |  | Applying the correction based on the trendline for bottles below 500
7 | NO | 100.062 | 33.3938 | 0.0007654 | -0.0087357 | -0.0052394 | 0.0034962 | 33.4025 | 14 | -0.0059815 |  | file pair | Sal0-bot | Sal1-bot | Sal1-Sal | Event#
7 | NO | 75.4766 | 32.8583 | 0.00074282 | -0.016472 | -0.0052394 | 0.011232 | 32.8748 | 14 | -0.0046005 |  | 1 | -0.00492 | -0.00621 | -0.00129 | 9
7 | NO | 50.2732 | 32.3862 | 0.0002663 | -0.0099335 | -0.0052394 | 0.0046941 | 32.3961 | 14 | -0.0048027 |  | 2 | -0.00494 | -0.00622 | -0.00128 | 27
7 | NO | 30.2248 | 32.3212 | 0.0022966 | -0.013374 | -0.0052394 | 0.0081349 | 32.3346 | 15 | -0.0056648 |  | 3 | -0.00496 | -0.00623 | -0.00127 | 36
7 | NO | 20.0443 | 32.2995 | 0.0061859 | -0.0061646 | -0.0052394 | 0.00092513 | 32.3057 | 15 | -0.0058899 |  | 4 | -0.00498 | -0.00624 | -0.00126 | 47
7 | NO | 10.8838 | 32.2867 | 0.0014092 | -0.0055924 | -0.0052394 | 0.00035293 | 32.2923 | 15 | -0.0047035 |  | 5 | -0.005 | -0.00625 | -0.00125 | 58
7 | NO | 4.686 | 32.2671 | 0.0067263 | -0.0029755 | -0.0052394 | -0.002264 | 32.2701 | 18 | -0.0060196 |  | 6 | -0.00502 | -0.00626 | -0.00124 | 66
8 | YES | 750.513 | 34.228 | 0.00023241 | -0.0040588 | -0.0052176 | -0.0011587 | 34.2321 | 19 | -0.0042057 |  | 7 | -0.00504 | -0.00627 | -0.00123 | 71
8 | YES | 500.707 | 34.0536 | 0.00059172 | -0.0059013 | -0.0052176 | 0.00068378 | 34.0595 |  |  |  | 8 | -0.00506 | -0.00628 | -0.00122 | 78
9 | NO | 5.291 | 32.562 | 0.0143 | 0.0018082 | -0.0051957 | -0.0070039 | 32.5602 |  |  |  | 9 | -0.00508 | -0.00629 | -0.00121 | 81
10 | YES | 750.621 | 34.2157 | 0.00026536 | -0.0034561 | -0.0051738 | -0.0017177 | 34.2192 |  |  |  | 10 | -0.0051 | -0.0063 | -0.0012 | 98
10 | NO | 5.4539 | 32.1111 | 0.00047203 | -0.0048866 | -0.0051738 | -0.00028722 | 32.116 |  |  |  | 11 | -0.00512 | -0.00631 | -0.00119 | 117
11 | YES | 1000.03 | 34.352 | 0.00028992 | -0.0021172 | -0.005152 | -0.0030348 | 34.3541 |  |  |  | 12 | -0.00514 | -0.00632 | -0.00118 | 120
11 | YES | 750.314 | 34.2106 | 0.00028172 | -0.004528 | -0.005152 | -0.00062394 | 34.2151 |  |  |  | 13 | -0.00516 | -0.00633 | -0.00117 | 123
11 | YES | 499.693 | 34.04 | 0.00023868 | -0.0025749 | -0.005152 | -0.0025771 | 34.0426 |  |  |  | 14 | -0.00518 | -0.00634 | -0.00116 | 135
11 | YES | 400.239 | 34.002 | 0.0002181 | -0.0055199 | -0.005152 | 0.00036788 | 34.0075 |  |  |  | 15 | -0.0052 | -0.00635 | -0.00115 | 142
11 | YES | 300.205 | 33.9769 | 0.00022431 | -0.0044708 | -0.005152 | -0.00068116 | 33.9814 |  |  |  | 16 | -0.00522 | -0.00636 | -0.00114 | 165
11 | NO | 249.8 | 33.9579 | 0.0011243 | -0.002243 | -0.005152 | -0.002909 | 33.9601 |  |  |  | 17 | -0.00524 | -0.00637 | -0.00113 | 183
11 | NO | 200.026 | 33.9121 | 0.00037352 | -0.0096321 | -0.005152 | 0.0044801 | 33.9217 |  |  |  | 18 | -0.00526 | -0.00638 | -0.00112 | 198
11 | NO | 175.796 | 33.8896 | 0.00056285 | -0.0071068 | -0.005152 | 0.0019548 | 33.8967 |  |  |  | 19 | -0.00528 | -0.00639 | -0.00111 | 201
11 | NO | 150.197 | 33.8446 | 0.00038875 | -0.0043983 | -0.005152 | -0.00075364 | 33.849 |  |  |  | 20 | -0.0053 | -0.0064 | -0.0011 | 203
11 | NO | 124.924 | 33.7168 | 0.0020446 | -0.010666 | -0.005152 | 0.0055139 | 33.7275
11 | NO | 99.9529 | 33.3798 | 0.0012324 | -0.0062828 | -0.005152 | 0.0011308 | 33.3861
11 | NO | 74.4902 | 32.8139 | 0.00079395 | 0.003746 | -0.005152 | -0.008898 | 32.8102
11 | NO | 50.1177 | 32.362 | 0.00059948 | -0.021328 | -0.005152 | 0.016176 | 32.3833
11 | NO | 29.9946 | 32.3249 | 0.00034622 | -0.01313 | -0.005152 | 0.0079782 | 32.338
11 | NO | 20.3109 | 32.2747 | 0.0010996 | -0.014011 | -0.005152 | 0.0088594 | 32.2887
11 | NO | 10.1474 | 32.1667 | 0.012004 | 0.013733 | -0.005152 | -0.018885 | 32.153
11 | NO | 5.1436 | 32.107 | 0.00098445 | -0.015846 | -0.005152 | 0.010694 | 32.1228
12 | YES | 1000.47 | 34.3309 | 0.00028956 | -0.0067062 | -0.0051301 | 0.0015761 | 34.3376
13 | NO | 149.923 | 33.8288 | 0.00024579 | -0.0056763 | -0.0051083 | 0.00056799 | 33.8345
14 | YES | 1000.4 | 34.3488 | 0.00025448 | -0.0078545 | -0.0050864 | 0.002768 | 34.3567
14 | YES | 750.108 | 34.2175 | 0.00025987 | -0.0059815 | -0.0050864 | 0.00089503 | 34.2235
14 | NO | 500.364 | 34.049 | 0.00041188 | -0.012318 | -0.0050864 | 0.0072313 | 34.0613
14 | YES | 400.565 | 34.0213 | 0.00044524 | -0.0046005 | -0.0050864 | -0.0004859 | 34.0259
14 | YES | 300.179 | 33.9269 | 0.00027663 | -0.0048027 | -0.0050864 | -0.00028372 | 33.9317
14 | NO | 250.425 | 33.9158 | 0.00045813 | -0.0053596 | -0.0050864 | 0.00027323 | 33.9212
14 | NO | 200.581 | 33.8882 | 0.00067811 | 0.0035934 | -0.0050864 | -0.0086799 | 33.8846
14 | NO | 175.276 | 33.8466 | 0.00093793 | -0.0018272 | -0.0050864 | -0.0032592 | 33.8484
14 | NO | 150.856 | 33.8104 | 0.0017596 | 0.00029755 | -0.0050864 | -0.005384 | 33.8101
14 | NO | 125.379 | 33.6041 | 0.001459 | -0.0017395 | -0.0050864 | -0.0033469 | 33.6058
14 | NO | 100.221 | 33.0463 | 0.0031701 | -0.068809 | -0.0050864 | 0.063723 | 33.1151
14 | NO | 74.8455 | 32.4002 | 0.00063469 | -0.003849 | -0.0050864 | -0.0012374 | 32.404
14 | NO | 50.4387 | 32.3113 | 0.0016072 | -0.0037613 | -0.0050864 | -0.0013251 | 32.3151
14 | NO | 29.9999 | 32.2547 | 0.008314 | -0.018372 | -0.0050864 | 0.013285 | 32.2731
14 | NO | 20.697 | 32.0689 | 0.0095363 | -0.012543 | -0.0050864 | 0.0074563 | 32.0814
14 | NO | 10.1986 | 31.8572 | 0.0014768 | -0.012337 | -0.0050864 | 0.0072503 | 31.8695
14 | NO | 5.0495 | 31.8626 | 0.00026149 | -0.0052566 | -0.0050864 | 0.00017023 | 31.8679
15 | YES | 1000.63 | 34.3479 | 0.00026039 | -0.0056648 | -0.0050646 | 0.00060026 | 34.3536
15 | YES | 749.547 | 34.2219 | 0.00025055 | -0.0058899 | -0.0050646 | 0.00082532 | 34.2278
15 | YES | 500.284 | 34.0709 | 0.00026253 | -0.0047035 | -0.0050646 | -0.00036105 | 34.0756
15 | NO | 5.2877 | 32.1233 | 0.0033439 | -0.010834 | -0.0050646 | 0.0057691 | 32.1341
16 | NO | 4.8691 | 31.4388 | 0.0014692 | -0.0076485 | -0.0050427 | 0.0026058 | 31.4464
17 | NO | 351.694 | 30.7464 | 0.00042067 | 0.01034 | -0.0050209 | -0.015361 | 30.7361
18 | YES | 350.153 | 31.2936 | 0.00019437 | -0.0060196 | -0.004999 | 0.0010206 | 31.2996
18 | NO | 1.8659 | 28.1991 | 0.09372 | 0.18123 | -0.004999 | -0.18623 | 28.0179
19 | YES | 315.259 | 31.1856 | 0.00027272 | -0.0042057 | -0.0049771 | -0.00077144 | 31.1898
19 | NO | 10.341 | 28.0903 | 0.0025939 | -0.10505 | -0.0049771 | 0.10008 | 28.1954
20 | NO | 166.969 | 31.0698 | 0.00025252 | -0.0058289 | -0.0049553 | 0.00087357 | 31.0756
20 | NO | 4.9782 | 27.1433 | 0.0047869 | -0.12496 | -0.0049553 | 0.12 | 27.2683
NO
NO
YES
YES
YES
NO
NO
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
NO
YES
NO
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
NO
YES
YES
NO
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
NO
NO
NO
YES
NO
YES
NO
NO
NO
FIT vs. Pressure
FIT vs. Sal0
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_4
Calculation of Linear Least Squares Fit For
(Sal1 - Sal_Bot)
3
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal1 - Sal_Bot), Prop=File Pair #
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.0595 |  | X | CALC Y
m= | 7.0628e-05 | 5.1267e-05 |  | Std Dev of Resid |  | 0.0012277 |  | 3 | -0.006608
b= | -0.0068199 | 0.00054641 |  | Resid Sum of Squares |  | 4.5218e-05 |  | 19 | -0.005478
# of Values Extracted = |  |  | 84
# of Values Used in Fit = |  |  | 32
File | Use | Pressure | Sal1 | Sal1 | Diff (Sal1 - Sal_Bot) |  |  |  | Fit vs. | File Pair #
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 11.666 | 31.7574 | 0.0018748 | 0.081078 | -0.0067493 | -0.087827 |  | 3 | -0.0079346
2 | NO | 5.0843 | 32.2382 | 0.023162 | 0.00074768 | -0.0066787 | -0.0074263 |  | 3 | -0.0066071
3 | YES | 999.676 | 34.3591 | 0.00026898 | -0.0079346 | -0.006608 | 0.0013265 |  | 3 | -0.0068359
3 | YES | 750.422 | 34.2373 | 0.00024923 | -0.0066071 | -0.006608 | -9.7166e-07 |  | 6 | -0.006218
3 | YES | 499.412 | 34.0844 | 0.00025273 | -0.0068359 | -0.006608 | 0.00022791 |  | 6 | -0.0062218
4 | NO | 4.9436 | 31.6195 | 0.0025438 | -0.0090866 | -0.0065374 | 0.0025492 |  | 6 | -0.0049248
5 | NO | 5.2706 | 32.2115 | 0.0013583 | -0.012135 | -0.0064668 | 0.0056678 |  | 6 | -0.0064735
6 | YES | 1250.46 | 34.4449 | 0.00025478 | -0.006218 | -0.0063961 | -0.00017819 |  | 7 | -0.0062904
6 | YES | 1000.98 | 34.3581 | 0.00041323 | -0.0062218 | -0.0063961 | -0.00017438 |  | 7 | -0.0073051
6 | YES | 749.707 | 34.2408 | 0.00023937 | -0.0049248 | -0.0063961 | -0.0014714 |  | 7 | -0.0070496
6 | YES | 500.565 | 34.1048 | 0.00039938 | -0.0064735 | -0.0063961 | 7.7392e-05 |  | 7 | -0.0063705
6 | NO | 5.4573 | 32.1937 | 0.0006505 | -0.007782 | -0.0063961 | 0.0013858 |  | 7 | -0.006012
7 | YES | 1499.79 | 34.4927 | 0.00025133 | -0.0062904 | -0.0063255 | -3.508e-05 |  | 7 | -0.0061302
7 | YES | 1250.06 | 34.443 | 0.00026968 | -0.0073051 | -0.0063255 | 0.00097963 |  | 7 | -0.0077209
7 | YES | 1000.44 | 34.3681 | 0.00041156 | -0.0070496 | -0.0063255 | 0.00072404 |  | 8 | -0.0050087
7 | YES | 750.726 | 34.2635 | 0.00023466 | -0.0063705 | -0.0063255 | 4.502e-05 |  | 8 | -0.0070343
7 | YES | 501.063 | 34.0996 | 0.00022741 | -0.006012 | -0.0063255 | -0.00031356 |  | 10 | -0.0046539
7 | YES | 400.164 | 34.0474 | 0.00024505 | -0.0061302 | -0.0063255 | -0.0001953 |  | 11 | -0.0031014
7 | YES | 300.306 | 33.9949 | 0.0003874 | -0.0077209 | -0.0063255 | 0.0013954 |  | 11 | -0.0053596
7 | NO | 250.891 | 33.9418 | 0.00049475 | -0.00010681 | -0.0063255 | -0.0062187 |  | 11 | -0.0034599
7 | NO | 200.259 | 33.8802 | 0.00037385 | -0.0084419 | -0.0063255 | 0.0021164 |  | 11 | -0.0064583
7 | NO | 175.129 | 33.8736 | 0.0010013 | -0.0065613 | -0.0063255 | 0.00023576 |  | 11 | -0.0054131
7 | NO | 150.334 | 33.805 | 0.00032338 | -0.007679 | -0.0063255 | 0.0013535 |  | 12 | -0.0077477
7 | NO | 125.834 | 33.6893 | 0.0004873 | -0.0087547 | -0.0063255 | 0.0024292 |  | 14 | -0.0087395
7 | NO | 100.062 | 33.3931 | 0.00054894 | -0.009388 | -0.0063255 | 0.0030624 |  | 14 | -0.0067482
7 | NO | 75.4766 | 32.8573 | 0.0010916 | -0.017464 | -0.0063255 | 0.011138 |  | 14 | -0.0056877
7 | NO | 50.2732 | 32.3853 | 0.00029743 | -0.010761 | -0.0063255 | 0.0044358 |  | 14 | -0.0056191
7 | NO | 30.2248 | 32.3205 | 0.00094629 | -0.014145 | -0.0063255 | 0.0078194 |  | 15 | -0.0065346
7 | NO | 20.0443 | 32.2944 | 0.0073216 | -0.011272 | -0.0063255 | 0.0049469 |  | 15 | -0.0067368
7 | NO | 10.8838 | 32.2855 | 0.0023495 | -0.0068054 | -0.0063255 | 0.0004799 |  | 15 | -0.005455
7 | NO | 4.686 | 32.2666 | 0.0046113 | -0.0034561 | -0.0063255 | -0.0028694 |  | 18 | -0.0062943
8 | YES | 750.513 | 34.2271 | 0.00028306 | -0.0050087 | -0.0062549 | -0.0012462 |  | 19 | -0.0039844
8 | YES | 500.707 | 34.0525 | 0.00055271 | -0.0070343 | -0.0062549 | 0.00077941
9 | NO | 5.291 | 32.56 | 0.014032 | -0.00023651 | -0.0061843 | -0.0059478
10 | YES | 750.621 | 34.2145 | 0.00027849 | -0.0046539 | -0.0061136 | -0.0014597
10 | NO | 5.4539 | 32.11 | 0.00044047 | -0.0059815 | -0.0061136 | -0.00013219
11 | YES | 1000.03 | 34.351 | 0.0003092 | -0.0031014 | -0.006043 | -0.0029417
11 | YES | 750.314 | 34.2097 | 0.00025944 | -0.0053596 | -0.006043 | -0.00068336
11 | YES | 499.693 | 34.0391 | 0.00022959 | -0.0034599 | -0.006043 | -0.0025831
11 | YES | 400.239 | 34.001 | 0.00027684 | -0.0064583 | -0.006043 | 0.00041527
11 | YES | 300.205 | 33.976 | 0.00023457 | -0.0054131 | -0.006043 | -0.00062995
11 | NO | 249.8 | 33.9573 | 0.00094731 | -0.0027504 | -0.006043 | -0.0032926
11 | NO | 200.026 | 33.9111 | 0.00038194 | -0.01059 | -0.006043 | 0.0045466
11 | NO | 175.796 | 33.8888 | 0.00040257 | -0.0079498 | -0.006043 | 0.0019068
11 | NO | 150.197 | 33.8437 | 0.00043273 | -0.0053215 | -0.006043 | -0.00072151
11 | NO | 124.924 | 33.7182 | 0.002569 | -0.0092621 | -0.006043 | 0.0032191
11 | NO | 99.9529 | 33.3789 | 0.001222 | -0.0071793 | -0.006043 | 0.0011363
11 | NO | 74.4902 | 32.8135 | 0.00098871 | 0.003315 | -0.006043 | -0.009358
11 | NO | 50.1177 | 32.361 | 0.00065488 | -0.022293 | -0.006043 | 0.01625
11 | NO | 29.9946 | 32.324 | 0.00031868 | -0.014023 | -0.006043 | 0.0079798
11 | NO | 20.3109 | 32.2734 | 0.00075513 | -0.015293 | -0.006043 | 0.0092501
11 | NO | 10.1474 | 32.1624 | 0.010007 | 0.0093803 | -0.006043 | -0.015423
11 | NO | 5.1436 | 32.1061 | 0.0008251 | -0.016655 | -0.006043 | 0.010612
12 | YES | 1000.47 | 34.3299 | 0.00033513 | -0.0077477 | -0.0059724 | 0.0017753
13 | NO | 149.923 | 33.8284 | 0.0002484 | -0.0061188 | -0.0059018 | 0.00021702
14 | YES | 1000.4 | 34.348 | 0.00024017 | -0.0087395 | -0.0058311 | 0.0029083
14 | YES | 750.108 | 34.2168 | 0.00026273 | -0.0067482 | -0.0058311 | 0.00091707
14 | NO | 500.364 | 34.048 | 0.00038845 | -0.013313 | -0.0058311 | 0.0074822
14 | YES | 400.565 | 34.0202 | 0.0002972 | -0.0056877 | -0.0058311 | -0.00014342
14 | YES | 300.179 | 33.9261 | 0.0002769 | -0.0056191 | -0.0058311 | -0.00021208
14 | NO | 250.425 | 33.915 | 0.00041516 | -0.0062256 | -0.0058311 | 0.00039446
14 | NO | 200.581 | 33.8873 | 0.00075273 | 0.0027313 | -0.0058311 | -0.0085624
14 | NO | 175.276 | 33.8455 | 0.00063247 | -0.0028763 | -0.0058311 | -0.0029548
14 | NO | 150.856 | 33.8099 | 0.0016631 | -0.0001564 | -0.0058311 | -0.0056747
14 | NO | 125.379 | 33.6033 | 0.0012771 | -0.002533 | -0.0058311 | -0.0032982
14 | NO | 100.221 | 33.0443 | 0.004068 | -0.070789 | -0.0058311 | 0.064958
14 | NO | 74.8455 | 32.3995 | 0.00055841 | -0.0044632 | -0.0058311 | -0.0013679
14 | NO | 50.4387 | 32.3107 | 0.00098834 | -0.0044022 | -0.0058311 | -0.001429
14 | NO | 29.9999 | 32.2571 | 0.0078733 | -0.015972 | -0.0058311 | 0.010141
14 | NO | 20.697 | 32.0697 | 0.0076251 | -0.011661 | -0.0058311 | 0.0058304
14 | NO | 10.1986 | 31.8559 | 0.0015664 | -0.013613 | -0.0058311 | 0.0077816
14 | NO | 5.0495 | 31.8616 | 0.00030624 | -0.0062752 | -0.0058311 | 0.00044405
15 | YES | 1000.63 | 34.3471 | 0.0002776 | -0.0065346 | -0.0057605 | 0.00077408
15 | YES | 749.547 | 34.2211 | 0.00029332 | -0.0067368 | -0.0057605 | 0.00097626
15 | YES | 500.284 | 34.0701 | 0.00027699 | -0.005455 | -0.0057605 | -0.00030548
15 | NO | 5.2877 | 32.122 | 0.0029561 | -0.012081 | -0.0057605 | 0.0063206
16 | NO | 4.8691 | 31.4385 | 0.0012721 | -0.0079441 | -0.0056899 | 0.0022542
17 | NO | 351.694 | 30.7461 | 0.0004233 | 0.0099716 | -0.0056192 | -0.015591
18 | YES | 350.153 | 31.2933 | 0.00023073 | -0.0062943 | -0.0055486 | 0.00074564
18 | NO | 1.8659 | 28.1858 | 0.09072 | 0.16787 | -0.0055486 | -0.17342
19 | YES | 315.259 | 31.1858 | 0.00024471 | -0.0039844 | -0.005478 | -0.0014935
19 | NO | 10.341 | 28.0885 | 0.0028468 | -0.10692 | -0.005478 | 0.10144
20 | NO | 166.969 | 31.0696 | 0.0004757 | -0.0059528 | -0.0054074 | 0.00054548
20 | NO | 4.9782 | 27.1455 | 0.0063467 | -0.12278 | -0.0054074 | 0.11737
NO
NO
YES
YES
YES
NO
NO
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
NO
YES
NO
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
NO
YES
YES
NO
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
NO
NO
NO
YES
NO
YES
NO
NO
NO
FIT vs. Pressure
FIT vs. Sal1
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_5
Calculation of Linear Least Squares Fit For
(Sal0 - Sal_Bot)
3
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal0 - Sal_Bot), Prop=File Pair #
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.012375 |  | X | CALC Y
m= | -3.9593e-05 | 8.8424e-05 |  | Std Dev of Resid |  | 0.0014148 |  | 3 | -0.0050486
b= | -0.0049298 | 0.00086286 |  | Resid Sum of Squares |  | 3.2026e-05 |  | 15 | -0.0055237
# of Values Extracted = |  |  | 84
# of Values Used in Fit = |  |  | 18
File | Use | Pressure | Sal0 | Sal0 | Diff (Sal0 - Sal_Bot) |  |  |  | Fit vs. | File Pair #
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 11.666 | 31.7592 | 0.0021347 | 0.082893 | -0.0049694 | -0.087863 |  | 3 | -0.006958
2 | NO | 5.0843 | 32.2374 | 0.019249 | -9.1553e-05 | -0.005009 | -0.0049174 |  | 3 | -0.0054893
3 | YES | 999.676 | 34.36 | 0.00027517 | -0.006958 | -0.0050486 | 0.0019095 |  | 6 | -0.0050087
3 | YES | 750.422 | 34.2384 | 0.00023329 | -0.0054893 | -0.0050486 | 0.00044079 |  | 6 | -0.0050278
3 | NO | 499.412 | 34.0854 | 0.00022847 | -0.0057716 | -0.0050486 | 0.00072308 |  | 6 | -0.0035667
4 | NO | 4.9436 | 31.6203 | 0.0029989 | -0.0082645 | -0.0050881 | 0.0031764 |  | 7 | -0.0053673
5 | NO | 5.2706 | 32.2118 | 0.0012103 | -0.011814 | -0.0051277 | 0.0066864 |  | 7 | -0.0062294
6 | YES | 1250.46 | 34.4461 | 0.00023507 | -0.0050087 | -0.0051673 | -0.00015863 |  | 7 | -0.0060043
6 | YES | 1000.98 | 34.3593 | 0.00037026 | -0.0050278 | -0.0051673 | -0.00013956 |  | 7 | -0.0052414
6 | YES | 749.707 | 34.2421 | 0.00023353 | -0.0035667 | -0.0051673 | -0.0016006 |  | 8 | -0.0040588
6 | NO | 500.565 | 34.106 | 0.0003891 | -0.0052834 | -0.0051673 | 0.00011603 |  | 10 | -0.0034561
6 | NO | 5.4573 | 32.1947 | 0.00083066 | -0.0068207 | -0.0051673 | 0.0016533 |  | 11 | -0.0021172
7 | YES | 1499.79 | 34.4936 | 0.0002311 | -0.0053673 | -0.0052069 | 0.00016035 |  | 11 | -0.004528
7 | YES | 1250.06 | 34.4441 | 0.00024556 | -0.0062294 | -0.0052069 | 0.0010225 |  | 12 | -0.0067062
7 | YES | 1000.44 | 34.3692 | 0.00042606 | -0.0060043 | -0.0052069 | 0.0007974 |  | 14 | -0.0078545
7 | YES | 750.726 | 34.2647 | 0.00020386 | -0.0052414 | -0.0052069 | 3.4463e-05 |  | 14 | -0.0059815
7 | NO | 501.063 | 34.1006 | 0.00021715 | -0.0049591 | -0.0052069 | -0.00024782 |  | 15 | -0.0056648
7 | NO | 400.164 | 34.0484 | 0.00022571 | -0.0051041 | -0.0052069 | -0.00010287 |  | 15 | -0.0058899 | -0.0052861
7 | NO | 300.306 | 33.996 | 0.00033204 | -0.006649 | -0.0052069 | 0.0014421
7 | NO | 250.891 | 33.9431 | 0.00055836 | 0.0012054 | -0.0052069 | -0.0064124
7 | NO | 200.259 | 33.8811 | 0.00025031 | -0.007473 | -0.0052069 | 0.0022661
7 | NO | 175.129 | 33.8745 | 0.00090458 | -0.0057259 | -0.0052069 | 0.00051893
7 | NO | 150.334 | 33.8062 | 0.00038394 | -0.0064964 | -0.0052069 | 0.0012895
7 | NO | 125.834 | 33.6902 | 0.00050522 | -0.0079117 | -0.0052069 | 0.0027048
7 | NO | 100.062 | 33.3938 | 0.0007654 | -0.0087357 | -0.0052069 | 0.0035287
7 | NO | 75.4766 | 32.8583 | 0.00074282 | -0.016472 | -0.0052069 | 0.011265
7 | NO | 50.2732 | 32.3862 | 0.0002663 | -0.0099335 | -0.0052069 | 0.0047265
7 | NO | 30.2248 | 32.3212 | 0.0022966 | -0.013374 | -0.0052069 | 0.0081674
7 | NO | 20.0443 | 32.2995 | 0.0061859 | -0.0061646 | -0.0052069 | 0.00095762
7 | NO | 10.8838 | 32.2867 | 0.0014092 | -0.0055924 | -0.0052069 | 0.00038542
7 | NO | 4.686 | 32.2671 | 0.0067263 | -0.0029755 | -0.0052069 | -0.0022315
8 | YES | 750.513 | 34.228 | 0.00023241 | -0.0040588 | -0.0052465 | -0.0011877
8 | NO | 500.707 | 34.0536 | 0.00059172 | -0.0059013 | -0.0052465 | 0.00065482
9 | NO | 5.291 | 32.562 | 0.0143 | 0.0018082 | -0.0052861 | -0.0070943
10 | YES | 750.621 | 34.2157 | 0.00026536 | -0.0034561 | -0.0053257 | -0.0018696
10 | NO | 5.4539 | 32.1111 | 0.00047203 | -0.0048866 | -0.0053257 | -0.00043907
11 | YES | 1000.03 | 34.352 | 0.00028992 | -0.0021172 | -0.0053653 | -0.0032481
11 | YES | 750.314 | 34.2106 | 0.00028172 | -0.004528 | -0.0053653 | -0.00083725
11 | NO | 499.693 | 34.04 | 0.00023868 | -0.0025749 | -0.0053653 | -0.0027904
11 | NO | 400.239 | 34.002 | 0.0002181 | -0.0055199 | -0.0053653 | 0.00015457
11 | NO | 300.205 | 33.9769 | 0.00022431 | -0.0044708 | -0.0053653 | -0.00089447
11 | NO | 249.8 | 33.9579 | 0.0011243 | -0.002243 | -0.0053653 | -0.0031223
11 | NO | 200.026 | 33.9121 | 0.00037352 | -0.0096321 | -0.0053653 | 0.0042668
11 | NO | 175.796 | 33.8896 | 0.00056285 | -0.0071068 | -0.0053653 | 0.0017415
11 | NO | 150.197 | 33.8446 | 0.00038875 | -0.0043983 | -0.0053653 | -0.00096695
11 | NO | 124.924 | 33.7168 | 0.0020446 | -0.010666 | -0.0053653 | 0.0053006
11 | NO | 99.9529 | 33.3798 | 0.0012324 | -0.0062828 | -0.0053653 | 0.00091751
11 | NO | 74.4902 | 32.8139 | 0.00079395 | 0.003746 | -0.0053653 | -0.0091113
11 | NO | 50.1177 | 32.362 | 0.00059948 | -0.021328 | -0.0053653 | 0.015963
11 | NO | 29.9946 | 32.3249 | 0.00034622 | -0.01313 | -0.0053653 | 0.0077649
11 | NO | 20.3109 | 32.2747 | 0.0010996 | -0.014011 | -0.0053653 | 0.0086461
11 | NO | 10.1474 | 32.1667 | 0.012004 | 0.013733 | -0.0053653 | -0.019098
11 | NO | 5.1436 | 32.107 | 0.00098445 | -0.015846 | -0.0053653 | 0.010481
12 | YES | 1000.47 | 34.3309 | 0.00028956 | -0.0067062 | -0.0054049 | 0.0013014
13 | NO | 149.923 | 33.8288 | 0.00024579 | -0.0056763 | -0.0054445 | 0.00023179
14 | YES | 1000.4 | 34.3488 | 0.00025448 | -0.0078545 | -0.0054841 | 0.0023704
14 | YES | 750.108 | 34.2175 | 0.00025987 | -0.0059815 | -0.0054841 | 0.00049738
14 | NO | 500.364 | 34.049 | 0.00041188 | -0.012318 | -0.0054841 | 0.0068336
14 | NO | 400.565 | 34.0213 | 0.00044524 | -0.0046005 | -0.0054841 | -0.00088355
14 | NO | 300.179 | 33.9269 | 0.00027663 | -0.0048027 | -0.0054841 | -0.00068137
14 | NO | 250.425 | 33.9158 | 0.00045813 | -0.0053596 | -0.0054841 | -0.00012442
14 | NO | 200.581 | 33.8882 | 0.00067811 | 0.0035934 | -0.0054841 | -0.0090775
14 | NO | 175.276 | 33.8466 | 0.00093793 | -0.0018272 | -0.0054841 | -0.0036568
14 | NO | 150.856 | 33.8104 | 0.0017596 | 0.00029755 | -0.0054841 | -0.0057816
14 | NO | 125.379 | 33.6041 | 0.001459 | -0.0017395 | -0.0054841 | -0.0037446
14 | NO | 100.221 | 33.0463 | 0.0031701 | -0.068809 | -0.0054841 | 0.063325
14 | NO | 74.8455 | 32.4002 | 0.00063469 | -0.003849 | -0.0054841 | -0.001635
14 | NO | 50.4387 | 32.3113 | 0.0016072 | -0.0037613 | -0.0054841 | -0.0017228
14 | NO | 29.9999 | 32.2547 | 0.008314 | -0.018372 | -0.0054841 | 0.012888
14 | NO | 20.697 | 32.0689 | 0.0095363 | -0.012543 | -0.0054841 | 0.0070586
14 | NO | 10.1986 | 31.8572 | 0.0014768 | -0.012337 | -0.0054841 | 0.0068526
14 | NO | 5.0495 | 31.8626 | 0.00026149 | -0.0052566 | -0.0054841 | -0.00022742
15 | YES | 1000.63 | 34.3479 | 0.00026039 | -0.0056648 | -0.0055237 | 0.00014116
15 | YES | 749.547 | 34.2219 | 0.00025055 | -0.0058899 | -0.0055237 | 0.00036622
15 | NO | 500.284 | 34.0709 | 0.00026253 | -0.0047035 | -0.0055237 | -0.00082015
15 | NO | 5.2877 | 32.1233 | 0.0033439 | -0.010834 | -0.0055237 | 0.00531
16 | NO | 4.8691 | 31.4388 | 0.0014692 | -0.0076485 | -0.0055633 | 0.0020852
17 | NO | 351.694 | 30.7464 | 0.00042067 | 0.01034 | -0.0056029 | -0.015943
18 | NO | 350.153 | 31.2936 | 0.00019437 | -0.0060196 | -0.0056424 | 0.00037715
18 | NO | 1.8659 | 28.1991 | 0.09372 | 0.18123 | -0.0056424 | -0.18687
19 | NO | 315.259 | 31.1856 | 0.00027272 | -0.0042057 | -0.005682 | -0.0014763
19 | NO | 10.341 | 28.0903 | 0.0025939 | -0.10505 | -0.005682 | 0.099373
20 | NO | 166.969 | 31.0698 | 0.00025252 | -0.0058289 | -0.0057216 | 0.00010723
20 | NO | 4.9782 | 27.1433 | 0.0047869 | -0.12496 | -0.0057216 | 0.11924
NO
NO
YES
YES
NO
NO
NO
YES
YES
YES
NO
NO
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
NO
NO
YES
NO
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
NO
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
FIT vs. Pressure
FIT vs. Sal0
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_6
Calculation of Linear Least Squares Fit For
(Sal1 - Sal_Bot)
3
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal1 - Sal_Bot), Prop=File Pair #
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.0015845 |  | X | CALC Y
m= | -1.3679e-05 | 8.5844e-05 |  | Std Dev of Resid |  | 0.0013735 |  | 3 | -0.0062264
b= | -0.0061853 | 0.00083768 |  | Resid Sum of Squares |  | 3.0184e-05 |  | 15 | -0.0063905
# of Values Extracted = |  |  | 84
# of Values Used in Fit = |  |  | 18
File | Use | Pressure | Sal1 | Sal1 | Diff (Sal1 - Sal_Bot) |  |  |  | Fit vs. | File Pair #
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 11.666 | 31.7574 | 0.0018748 | 0.081078 | -0.006199 | -0.087277 |  | 3 | -0.0079346
2 | NO | 5.0843 | 32.2382 | 0.023162 | 0.00074768 | -0.0062127 | -0.0069604 |  | 3 | -0.0066071
3 | YES | 999.676 | 34.3591 | 0.00026898 | -0.0079346 | -0.0062264 | 0.0017082 |  | 6 | -0.006218
3 | YES | 750.422 | 34.2373 | 0.00024923 | -0.0066071 | -0.0062264 | 0.00038069 |  | 6 | -0.0062218
3 | NO | 499.412 | 34.0844 | 0.00025273 | -0.0068359 | -0.0062264 | 0.00060957 |  | 6 | -0.0049248
4 | NO | 4.9436 | 31.6195 | 0.0025438 | -0.0090866 | -0.0062401 | 0.0028466 |  | 7 | -0.0062904
5 | NO | 5.2706 | 32.2115 | 0.0013583 | -0.012135 | -0.0062537 | 0.0058809 |  | 7 | -0.0073051
6 | YES | 1250.46 | 34.4449 | 0.00025478 | -0.006218 | -0.0062674 | -4.9453e-05 |  | 7 | -0.0070496
6 | YES | 1000.98 | 34.3581 | 0.00041323 | -0.0062218 | -0.0062674 | -4.5643e-05 |  | 7 | -0.0063705
6 | YES | 749.707 | 34.2408 | 0.00023937 | -0.0049248 | -0.0062674 | -0.0013426 |  | 8 | -0.0050087
6 | NO | 500.565 | 34.1048 | 0.00039938 | -0.0064735 | -0.0062674 | 0.00020613 |  | 10 | -0.0046539
6 | NO | 5.4573 | 32.1937 | 0.0006505 | -0.007782 | -0.0062674 | 0.0015146 |  | 11 | -0.0031014
7 | YES | 1499.79 | 34.4927 | 0.00025133 | -0.0062904 | -0.0062811 | 9.3481e-06 |  | 11 | -0.0053596
7 | YES | 1250.06 | 34.443 | 0.00026968 | -0.0073051 | -0.0062811 | 0.0010241 |  | 12 | -0.0077477
7 | YES | 1000.44 | 34.3681 | 0.00041156 | -0.0070496 | -0.0062811 | 0.00076847 |  | 14 | -0.0087395
7 | YES | 750.726 | 34.2635 | 0.00023466 | -0.0063705 | -0.0062811 | 8.9448e-05 |  | 14 | -0.0067482
7 | NO | 501.063 | 34.0996 | 0.00022741 | -0.006012 | -0.0062811 | -0.00026913 |  | 15 | -0.0065346
7 | NO | 400.164 | 34.0474 | 0.00024505 | -0.0061302 | -0.0062811 | -0.00015087 |  | 15 | -0.0067368 | -0.0063085
7 | NO | 300.306 | 33.9949 | 0.0003874 | -0.0077209 | -0.0062811 | 0.0014399
7 | NO | 250.891 | 33.9418 | 0.00049475 | -0.00010681 | -0.0062811 | -0.0061743
7 | NO | 200.259 | 33.8802 | 0.00037385 | -0.0084419 | -0.0062811 | 0.0021608
7 | NO | 175.129 | 33.8736 | 0.0010013 | -0.0065613 | -0.0062811 | 0.00028019
7 | NO | 150.334 | 33.805 | 0.00032338 | -0.007679 | -0.0062811 | 0.0013979
7 | NO | 125.834 | 33.6893 | 0.0004873 | -0.0087547 | -0.0062811 | 0.0024736
7 | NO | 100.062 | 33.3931 | 0.00054894 | -0.009388 | -0.0062811 | 0.0031069
7 | NO | 75.4766 | 32.8573 | 0.0010916 | -0.017464 | -0.0062811 | 0.011183
7 | NO | 50.2732 | 32.3853 | 0.00029743 | -0.010761 | -0.0062811 | 0.0044802
7 | NO | 30.2248 | 32.3205 | 0.00094629 | -0.014145 | -0.0062811 | 0.0078638
7 | NO | 20.0443 | 32.2944 | 0.0073216 | -0.011272 | -0.0062811 | 0.0049913
7 | NO | 10.8838 | 32.2855 | 0.0023495 | -0.0068054 | -0.0062811 | 0.00052433
7 | NO | 4.686 | 32.2666 | 0.0046113 | -0.0034561 | -0.0062811 | -0.002825
8 | YES | 750.513 | 34.2271 | 0.00028306 | -0.0050087 | -0.0062948 | -0.0012861
8 | NO | 500.707 | 34.0525 | 0.00055271 | -0.0070343 | -0.0062948 | 0.00073953
9 | NO | 5.291 | 32.56 | 0.014032 | -0.00023651 | -0.0063085 | -0.0060719
10 | YES | 750.621 | 34.2145 | 0.00027849 | -0.0046539 | -0.0063221 | -0.0016682
10 | NO | 5.4539 | 32.11 | 0.00044047 | -0.0059815 | -0.0063221 | -0.00034068
11 | YES | 1000.03 | 34.351 | 0.0003092 | -0.0031014 | -0.0063358 | -0.0032345
11 | YES | 750.314 | 34.2097 | 0.00025944 | -0.0053596 | -0.0063358 | -0.00097616
11 | NO | 499.693 | 34.0391 | 0.00022959 | -0.0034599 | -0.0063358 | -0.0028759
11 | NO | 400.239 | 34.001 | 0.00027684 | -0.0064583 | -0.0063358 | 0.00012247
11 | NO | 300.205 | 33.976 | 0.00023457 | -0.0054131 | -0.0063358 | -0.00092275
11 | NO | 249.8 | 33.9573 | 0.00094731 | -0.0027504 | -0.0063358 | -0.0035854
11 | NO | 200.026 | 33.9111 | 0.00038194 | -0.01059 | -0.0063358 | 0.0042538
11 | NO | 175.796 | 33.8888 | 0.00040257 | -0.0079498 | -0.0063358 | 0.001614
11 | NO | 150.197 | 33.8437 | 0.00043273 | -0.0053215 | -0.0063358 | -0.0010143
11 | NO | 124.924 | 33.7182 | 0.002569 | -0.0092621 | -0.0063358 | 0.0029263
11 | NO | 99.9529 | 33.3789 | 0.001222 | -0.0071793 | -0.0063358 | 0.00084345
11 | NO | 74.4902 | 32.8135 | 0.00098871 | 0.003315 | -0.0063358 | -0.0096508
11 | NO | 50.1177 | 32.361 | 0.00065488 | -0.022293 | -0.0063358 | 0.015957
11 | NO | 29.9946 | 32.324 | 0.00031868 | -0.014023 | -0.0063358 | 0.007687
11 | NO | 20.3109 | 32.2734 | 0.00075513 | -0.015293 | -0.0063358 | 0.0089573
11 | NO | 10.1474 | 32.1624 | 0.010007 | 0.0093803 | -0.0063358 | -0.015716
11 | NO | 5.1436 | 32.1061 | 0.0008251 | -0.016655 | -0.0063358 | 0.010319
12 | YES | 1000.47 | 34.3299 | 0.00033513 | -0.0077477 | -0.0063495 | 0.0013982
13 | NO | 149.923 | 33.8284 | 0.0002484 | -0.0061188 | -0.0063632 | -0.0002444
14 | YES | 1000.4 | 34.348 | 0.00024017 | -0.0087395 | -0.0063768 | 0.0023626
14 | YES | 750.108 | 34.2168 | 0.00026273 | -0.0067482 | -0.0063768 | 0.00037135
14 | NO | 500.364 | 34.048 | 0.00038845 | -0.013313 | -0.0063768 | 0.0069365
14 | NO | 400.565 | 34.0202 | 0.0002972 | -0.0056877 | -0.0063768 | -0.00068914
14 | NO | 300.179 | 33.9261 | 0.0002769 | -0.0056191 | -0.0063768 | -0.0007578
14 | NO | 250.425 | 33.915 | 0.00041516 | -0.0062256 | -0.0063768 | -0.00015126
14 | NO | 200.581 | 33.8873 | 0.00075273 | 0.0027313 | -0.0063768 | -0.0091082
14 | NO | 175.276 | 33.8455 | 0.00063247 | -0.0028763 | -0.0063768 | -0.0035006
14 | NO | 150.856 | 33.8099 | 0.0016631 | -0.0001564 | -0.0063768 | -0.0062204
14 | NO | 125.379 | 33.6033 | 0.0012771 | -0.002533 | -0.0063768 | -0.0038439
14 | NO | 100.221 | 33.0443 | 0.004068 | -0.070789 | -0.0063768 | 0.064412
14 | NO | 74.8455 | 32.3995 | 0.00055841 | -0.0044632 | -0.0063768 | -0.0019136
14 | NO | 50.4387 | 32.3107 | 0.00098834 | -0.0044022 | -0.0063768 | -0.0019747
14 | NO | 29.9999 | 32.2571 | 0.0078733 | -0.015972 | -0.0063768 | 0.0095953
14 | NO | 20.697 | 32.0697 | 0.0076251 | -0.011661 | -0.0063768 | 0.0052847
14 | NO | 10.1986 | 31.8559 | 0.0015664 | -0.013613 | -0.0063768 | 0.0072359
14 | NO | 5.0495 | 31.8616 | 0.00030624 | -0.0062752 | -0.0063768 | -0.00010167
15 | YES | 1000.63 | 34.3471 | 0.0002776 | -0.0065346 | -0.0063905 | 0.00014405
15 | YES | 749.547 | 34.2211 | 0.00029332 | -0.0067368 | -0.0063905 | 0.00034623
15 | NO | 500.284 | 34.0701 | 0.00027699 | -0.005455 | -0.0063905 | -0.00093551
15 | NO | 5.2877 | 32.122 | 0.0029561 | -0.012081 | -0.0063905 | 0.0056906
16 | NO | 4.8691 | 31.4385 | 0.0012721 | -0.0079441 | -0.0064042 | 0.0015399
17 | NO | 351.694 | 30.7461 | 0.0004233 | 0.0099716 | -0.0064179 | -0.01639
18 | NO | 350.153 | 31.2933 | 0.00023073 | -0.0062943 | -0.0064316 | -0.00013731
18 | NO | 1.8659 | 28.1858 | 0.09072 | 0.16787 | -0.0064316 | -0.1743
19 | NO | 315.259 | 31.1858 | 0.00024471 | -0.0039844 | -0.0064452 | -0.0024608
19 | NO | 10.341 | 28.0885 | 0.0028468 | -0.10692 | -0.0064452 | 0.10047
20 | NO | 166.969 | 31.0696 | 0.0004757 | -0.0059528 | -0.0064589 | -0.00050608
20 | NO | 4.9782 | 27.1455 | 0.0063467 | -0.12278 | -0.0064589 | 0.11632
NO
NO
YES
YES
NO
NO
NO
YES
YES
YES
NO
NO
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
NO
NO
YES
NO
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
NO
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
FIT vs. Pressure
FIT vs. Sal1
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE