#### 2023-066

CCGS John P Tully

#### Dimethylsulfide (DMS) & Dimethylsulfoniopropionate (DMSP) Report

#### April 30-May 16, 2023

Prepared by Michael Arychuk

- Sample Collection

Samples were collected from stations  P2, P4, P12, P16, P20 and P26 for DMS, DMSPD (dissolved) & DMSPT (total).

  - DMS

The sampling depths were adjusted on this cruise.  Eleven water samples (100m, 75m, 50m, 40m, 30m, 25m, 20m, 15m, 10m, 5m, surface) were collected at station P2.  Fourteen samples (600m, 200m, 150m, 100m, 75m, 50m, 40m 30m, 25m, 20m, 15m, 10m, 5m, surface) were collected at P4,  P12, P16, P20 & P26. Duplicates were taken at 20m.  In all cases, samples were collected in 250mL ground glass stoppered bottles and stored in a fridge, in the dark and removed one at a time before analysis.

Additional samples were taken in 125mL ground glass stoppered bottles at stations P12, P16 & P26 to run alongside a secondary “new” system.

  - DMSP

Seven samples (400m, 200m, 150m, 30m, 20m, 5m, surface) for both DMSPD and DMSPT were collected at P4, P12, P16, P20 & P26.  Four samples (100m, 20m, 5m, surface) were collected at P2. Duplicates were taken at 20m.

- Analysis

7890 DMS System was used for all stations except P26 for which the new 8890 DMS System was used.

2.1 DMS

A sample was loaded onto the stripper and purged with UHP Helium for 10 minutes at ~100mL/min.  The DMS was extracted from the water and absorbed onto a Tenax TA trap kept at -80oC.  The trap was subsequently desorbed at 100oC (with a dewar containing boiling water) onto a Chromasil 330 column which eluted onto a Flame Photometric Detector (FPD).  All samples were run as soon as possible after being collected.

A secondary DMS system (8890 DMS System) was taken out on this cruise to field test the equipment and compare the data with that of the older system.  The DMS was extracted from the seawater in the same way but was absorbed onto a 1/16” Teflon tubing “trap” immersed in liquid nitrogen.  After 7 minutes the trap is desorbed at 100oC (with a dewar containing boiling water) onto a HP-1 capillary column which eluted into a Flame Photometric Detector (FPD).  Like the older system, all samples were run as soon as possible after being collected.

2.2 DMSPD

Approximately 50-75mL of seawater was allowed to flow directly from the niskin into a filtration funnel containing a 0.7m GF/F filter.  The first 3.5mL was collected in a 15mL or 5mL polypropylene tube.  The entire 3.5mL was then transferred into a 5ml, glass, serum bottle and 50L of a 50% sulphuric acid/water solution was added.    The sample was then crimp sealed and stored in the dark and at 4˚C where it would be analysed back at IOS at a later date.

2.3 DMSPT

Exactly 3.5mL of seawater was collected directly from the niskin into a 15mL or 5mL polypropylene tube. The entire 3.5mL was then transferred into a 5ml, glass, serum bottle and 50L of a 50% sulphuric acid/water solution was added.    The sample was then crimp sealed and stored in the dark and at 4˚C where it would be analysed back at IOS at a later date.

- Calibration

3.1 DMS

A four to six level calibration table was used for calculating the concentrations of DMS.  The standards were prepared in MQ water (brought from the lab and stored in 4L amber glass bottles) for stations P2 & P4 and then prepared in 2000m or deeper seawater (also stored in amber glass bottles kept at 4˚C) for stations P12, P16, P20 & P26.  Standards were run under the same conditions, as described above, for the samples and the calibration curve was valid for 12 hours.  If analysis exceeded 12 hours, a continuing calibration standard was run to ensure the calibration curve was still within acceptable limits.

- Quality Control

4.1 DMS

System blanks and duplicates were run approximately every 13 samples to ensure the system remained free of contamination and had acceptable reproducibility.

4.2 DMSP

Blanks and duplicates were collected at every station.  Blanks were done by either using MQ water, some of the bottled SQ water brought from the lab or using some deep sea water (2000 m or deeper) and treating them as actual samples.    For example, in the case of DMSPD,   3.5mL was collected from a separate funnel and for DMSPT    3.5mL  was added directly to the polypropylene tube.  Like the samples, they were then transferred into a 5ml, glass, serum bottle and 50L of a 50% sulphuric acid/water solution was added.    The blank was then crimp sealed and stored in the dark and at 4˚C where it would be analysed back at IOS at a later date.

- Data & Results

  - DMS

Data from this cruise was compared between two systems for stations P12, P16 & P26.  The newer system performed well and is faster than the older system but is not as precise and reproducibility is not as good as on the old system.  For all stations except P26 the DMS data was reported from the old system.  At P26 the old system suffered a software error and data from that station is from the new system.

  - DMSP

The DMSP samples were collected for all stations.  Samples were run back in the lab in late May 2023.

- Conclusions

6.1 DMS

Additional refinements are required on the new DMS system.  In an emergency it can be used but until the reproducibility can be improved the older DMS system should continue to be used.

  - DMSP

Sample analysis went well without any issues.