## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2023-078

Agency: OSD, Ocean Sciences Division, Sidney, B.C.

Party Chief: Boldt J.				Location: Strait of Georgia

Project: Strait of Georgia Juvenile Herring	Platform: Walker Rock

Date: 5 September 2023 – 19 September 2023

Processed by: Germaine Gatien

Date of Processing: 19 July 2024 –   22 July 2024

Number of original HEX files: 20		Number of CTD files: 20

# INSTRUMENT SUMMARY

SeaBird Model SBE-19+ CTD (s/n 8118) was used for all casts. There were no external sensors attached.

# SUMMARY OF QUALITY AND CONCERNS

A spreadsheet with header information was provided which included start times and positions.

The file names were non-standard. It used to be easy to edit such names as a group as long as they contained the event number, but that tool is no longer available; each file had to be renamed separately.  In future it would make processing go faster if the following simple format is used for raw file names:

YYYY-###-EEEE plus extension    (ex. 2023-078-0001.hex)

### is the 3-digit cruise number and EEEE is the 4-digit event number.

The descent rate of the CTD was generally very high, around 1.5m/s, except for very shallow casts. A speed of 1m/s is recommended. Higher speeds limit the amount of data collected, so the 4Hz setting only yields about 2.5 records per metre.

There were surface soaks of at least 1 minute after the pump turned on before casts began; most soaks were much longer than that.

There was no calibration sampling for salinity. The calibrations were from February 2021.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea.

The file names were non-standard and were renamed after conversion to CNV files.

##### Preliminary Steps

A spreadsheet with positions and comments for casts was provided.  The event numbers were 1 to 20.

The cruise summary sheet was completed.

A single configuration file was used for all casts.

There were no parameters on file at IOS for the temperature, conductivity and pressure calibrations, since this CTD belongs to ESD at PBS. It was used during cruise 2021-071 and 2022-078 which was processed by OSD. No changes have been made to the configuration since 2021-071 was run.

The configuration file was saved as file 2023-078-ctd.xmlcon.

##### Conversion of Raw Data

All casts were converted using file 2023-078-ctd.xmlcon.

For all casts there was a soak of between 1 and 3.5 minutes with the majority between 1 and 2 minutes.

The pumps were set to come on after a minimum conductivity was reached which took from 6s to 51s.

There was no subsurface soak.

##### FILTER

A low-pass filter with a 0.5s time constant was applied to the conductivity to force it to have the same response as the temperature. Pressure was very smooth, so no filter was applied..

##### ALIGNCTD

Temperature and conductivity were advanced by +0.5s relative to pressure.

##### DERIVE

Program DERIVE was run to calculate salinity and dissolved oxygen concentration.

Tests were run to compare casts with no filter and/or no align step to ensure they improved the data and they did.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert the CNV files to IOS Headers.

Profiles of temperature and salinity and temperature versus salinity plots were examined to see that data look reasonable. There are some unstable features so some further alignment looks appropriate.

##### Checking Headers

Some header information was missing: latitude, longitude, station name and water depth. Spreadsheet “Boldt_Thompson_JuvHerringSOG_CTD_2023.xlsx” provided that information.

The file was saved as 2023-078-header-mrg.csv and position formats were adjusted slightly to ones suitable for IOS SHELL.

IOS SHELL program “MERGE:CSV file to headers” was used to add information from the csv file to the IOS files with output MRH.

The start times from the files taken from the instrument’s time stamp are within a few minutes of those in the spreadsheet provided, so those were not included in the merge process..

Times are local (evening sampling for all casts), so 7 hours need to be added to get UTC.

ADD TIME CHANNEL was run to add 7 hours to all files.

A cross-reference listing was produced and no problems were found.

A preliminary header check was also run and no errors were found.

Pressures at the end of casts varied from -0.02db to +0.05db, so well within specifications.

Track plots look good and were added to the end of this report.

##### CLIP

Plots were made to determine how many scans should be removed from the beginning of casts. The results were highly variable and it looks like picking a minimum depth would work better. There were soaks between 1db and 1.6db.

CLIP was run to remove records until the full cast began; choices of 1.2, 1.3, 1.4 or 1.6db were used, as appropriate, to get a full profile.

##### SHIFT

Conductivity

Small shifts to conductivity were tested to see if they improved the salinity, as judged by T-S plots. There were very few unstable features at depth so differences were slight, but -2 records was found to be the best choice.

SHIFT was run to apply an advance of –2 records to conductivity and salinity was recalculated.

##### DELETE and CTDEDIT

DELETE was run on all casts using the following parameters:

Remove Surface Records in the top 10m.

Surface Record Removal: Last Press Min.	Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 5 points) was deleted from 10db to 10db above the maximum pressure.

COMMENTS ON WARNINGS: There were no warnings.

All DEL files were copied to *.EDT so that there will be a complete set of EDT files after editing.

CTDEDIT was used to edit all casts except from events #17 and #18.

Most editing involved light cleaning of salinity and removal of a few records at the top and/or bottom of cast. Cast #16 required heavier cleaning of salinity near the bottom.

##### Inter-comparisons

Sensor History – The CTD was used for cruise 2021-071 in September 2021 and 2022-078 in September 2022. There was no calibration sampling.

Comparison of repeat casts –None. Nearby casts too shallow to be suitable for comparison.

Historic Ranges – Most casts were located very close to shore and the 3-standard deviation climatology is not suitable for such areas. Nonetheless, all salinity and temperatures values were within the local climatology, where  available.

Post-cruise calibrations – None available.

##### CALIBRATE

There was no calibration sampling. No recalibration was applied.

##### Bin Average and REMOVE

The files were bin-averaged using 1db bins.

T-S plots were examined. No problems were found.

REMOVE was run to remove Scan_Number, Descent Rate and Flag channels.

##### HEADER EDIT and final checks.

Header Edit was used to fix headers, fix formats and channel names and to add some comments to the headers:

A cross-reference listing was produced.

A header check was run on the CTD files and no errors were found.

The sensor history was updated.

## CRUISE SUMMARY

| Cruise ID#: 2023-078 |  |  |  |  |
| --- | --- | --- | --- | --- |
| Dates: Start: 5 September 2023 End: 19 September 2023 |  |  |  |  |
| Location: Strait of Georgia |  |  |  |  |
| Chief Scientist: Boldt J. |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? |
| 1 | SEABIRD | 19+ | 8118 | No |

CTD CALIBRATION INFORMATION

Make/Model/Serial#:	SEABIRD/SBE19+/8118  Cruise ID#:		2023-078

| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| --- | --- | --- | --- | --- | --- |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 8118 | 18Feb2021 | Factory |  |  |
| Conductivity | 8118 | 18Feb2021 | Factory |  |  |
| Pressure | 8118 | 21Feb2021 | Factory |  |  |