## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2023-080

Agency: Ocean Sciences Division (OSD)

Location: Salish Sea

Project:  Geochemical monitoring

Party Chief: Johannessen, S.

Platform: Vector

Date: 1-6 November 2023

Processed by: Jeannette Bedard

Date of Processing: 11-31 Oct 2024

Number of original HEX files: 27

Number of CTD files:  	27	Number of CTD casts processed: 27

Number of bottle files: 	27	Number of bottle casts processed: 10

Number of TSG files:   0	Number of TSG files processed:  0

# INSTRUMENT SUMMARY

SBE911+ CTD 0443 was mounted in a rosette and attached were Wetlabs CSTAR transmissometers 1185DR and 1883DG, SBE 43 DO sensor 4372, SeaPoint fluorometer 4186, and altimeter 76341.

Seasave version 7.26.7.121 was used for acquisition.

There was no thermosalinograph data.

# SUMMARY OF QUALITY AND CONCERNS

The Science Log was in digital format. The header page had complete information about participants and equipment. There were good notes in the comments sections of both the CTD and Sampling logs about problems encountered.

# PROCESSING SUMMARY

GENERAL

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

- The digital daily log and rosette log sheets were obtained and checked for comments.

- Chlorophyll, Salinity, Nutrients and Dissolved Oxygen data were obtained in QF spreadsheets.

- The history of the pressure sensor, conductivity and dissolved oxygen sensors were checked.

- The calibration control files were obtained – as they stayed the same through out the cruise 2023-080-0001.xmlcon was used for processing.

##### BOTTLE FILE PREPARATION

The ROS files were created using 2023-080-0001.xmlcon. The hysteresis correction was not selected since there were no very deep casts; the Tau correction was selected.

The ROS files were converted to IOS format

The files were put through CLEAN to create BOT files. Temperature and salinity were plotted for all BOT files to check for outliers, none were found.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv.

- For Event 1 – from the Daily Science Log: ‘no samples taken, 24 bottles tripped to test’ and there is no sampling log – this cast was removed from further bottle data processing.

- In Event 48 there was the following comment in the Sampling Log: ‘TOC sampling experiment – not part of regular sampling – gave 800 series numbers to differentiate’

- Sample numbers were added based on the rosette logs.

The ADDSAMP file was then sorted on event number & then sample number. Converted to CST files using Convert Spreadsheet.

It was used to add sample numbers to the BOT files – output *.SAM.

The SAM files were bin-averaged on bottle # and called SAMAVG.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

Next, the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were saved in file 2023-080-bot-hdr.txt.

SALINITY

Salinity analysis was obtained in file QF2023-080_SAL_21Nov2023.xlsx. The analyses were carried out in a temperature-controlled lab within 30-34 days of collection. The files were simplified and saved as 2023-080_SAL.csv. That file was then converted to individual SAL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2023-080_OXY_18Jan2024.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and saved as 2023-080_OXY.csv. That file was converted into individual *.OXY files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2023_080_NUTS_2023-11-22.xlsx which includes flags, comments and a precision study. The spreadsheet page with the final data was simplified and saved as 2023-080_NUTS.csv. That file was converted into individual *.NUTS files.

CHLOROPHYL

The chlorophyl data were obtained in spreadsheet QF2023_080_CHL_19Jan2024.xlsx which includes flags, comments and a precision study. The spreadsheet page with the final data was simplified and saved as 2023-080_CHL.csv. That file was converted into individual *.CHL files.

The SAL, CHL, OXY, and NUTS files were merged with CST files in 4 steps. The files were then put through CLEAN to reduce the headers to File and Comment sections only. These files were ordered on sample number, but the SAMAVG files are ordered on bottle number, so the MRGCLN1 files were reordered on Bottle_Number and saved as *. MRGCLN1s. The MRGCLN1s files were then merged with SAMAVG files using merge channel Bottle_Number. There were a few issues:

- Cast 43 and 46 both had duplicates, they were averaged together and given a flag of 6

- Cast 48 did not have any samples associated with it and was removed bringing the total of files down to 10.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. None were found.

A header check was run and no problems were found.

##### Compare

Salinity

This cruise had salinity sampling between 3-714 m. In the top 10 m there are larger differences between the bottle and measured salinity reaching a maximum difference of 1.5. The only other outlier is in cast 13 where the difference at 50 m is 4.7. For the remaining data the fit against pressure is reasonably flat. For more detail see file 2023-080-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run between measured and bottle dissolved oxygen with pressure as the reference channel. There was more variability at shallower depths, otherwise the fit is reasonably flat. Note, the cast were from a coastal environment. For more detail see file 2023-080-oxy-comp1.xls.

Fluorescence

COMPARE was run between measured fluorescence and bottle chlorophyl. There were larger differences near the surface. For more detail see file 2023-080-chl-fl-comp1.xls.

##### Conversion of Full CTD Files from Raw Data

All files were converted using files 2023-080-0001.xmlcon. The hysteresis correction was not selected since there were no deep casts; the Tau correction was selected. A few casts were examined and all expected channels are present.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the

full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

CELLTM was run using the default settings (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SeaBird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number. There are no negative fluorescence values.

##### Checking Headers

A header check and cross-reference list was run.

- There was no sounder depth for casts 33, 35, 47, 48, 49 – manually added a max water depth of max sampling depth + altimeter.

- Water depth was missing from cast 41, however in there was a sounder depth in the science log of 481m – manually added to file

Surface check – mean surface pressure was 0.996 db, with a minimum of 0.58 db and a maximum of 2.72 db, which is within the expected range from the Vector.

The track plots were added to the end of this report, one with station names and the other with cast number.

A check was made of the water depths entered in the file headers by adding the maximum depth sampled plus the altimetry reading when the CTD was at the bottom of the cast.

Five out of the 26 casts (32, 40, 41, 44 and 46) had values above ± 5m but all were below ± 7m. After cross-checking with the log, it was determined the difference was most likely due to the complex bathymetry found in coastal environments.

##### Shift

Fluorescence

SHIFT was run on the fluorescence channel with an advance of 24 records. Plots showed this setting was appropriate.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier using Seabird software. No further alignment is needed for the DO concentration channel.

Conductivity

Tests were run on 3 casts to determine the best setting to align conductivity and temperature by judging the effect on salinity as seen in T-S space. SHIFT was run on all casts using -0.65 records for the primary conductivity and -0.2 records for the secondary. Salinity was recalculated for both channels.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

##### Other Comparisons

Previous experience with these sensors – These sensors were used on 2023-029 shortly before this cruise (there was one cruise in between).

Historic ranges – These casts were all taken in coastal sites and fit within climatology values temperature for the most part. The salinity values do not agree as well.. There is no indication of calibration problems.

Post-Cruise Calibration – There were no post-cruise calibrations available.

##### DETAILED EDITING

The primary T-S pair were chosen for editing and eventual archiving. For this cruise both primary and secondary were very similar.

All DEL files were uploaded into the CTD-QC File Processor to obtain AI generated data quality flags.

CTDEDIT was used to remove records near the top and bottom of the casts where the descent rate was below 0.2 m/s. The occasional salinity spike was also removed.

The edited files were copied to *.EDT.

After editing T-S plots were examined for all casts; there remain many small unstable features but they are in areas when they may well reflect real conditions. No further editing was applied.

##### Fluorescence Processing

A median filter with a width that varies with pressure was applied to the fluorescence channel using:

- Up to pressure of 10 db, filter width of 3

- Up to pressure of 40 db, filter width of 5

- Up to pressure of 100 db, filter width of 7

##### Recalibration

The minimum CTD Salinity was ~10 psu, below the <25 psu threshold that requires correction to silicate samples. Silicate was calibrated using 2024-080-recal-SIL.ccf.

##### Remove unwanted channels

The following channels were removed from the bottle files:

- Scan_Number

- Salinity:T1:C1

- Temperature:Secondary

- Conductivity:Secondary

- Altimeter

- Status:Pump

- Descent_Rate

- Flag

- Oxygen:Voltage:SBE

- Sample_Number:2nd

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined.

Profile plots were examined to see if there any problems. No problems were noted.

##### Final CTD File Steps (REMOVE and HEADEDIT)

- For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Secondary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T1:C1, Flag and Prediction_Flag.

- A second SBE DO channel (with umol/kg units) was derived.

- HEADER EDIT was used to fix formats and channel names and to add comments to the headers.

- The Standards Check routine was run and no problems were found.

- The Header Check was run; no problems were found.

- Profile and T-S plots were examined and no problems were found.

- The sensor history was updated.

##### Final Bottle Files

- The MRGCOR1 files were put through SORT to order on increasing pressure.

- A second SBE DO channel with mass units was added for both the CTD DO and REORDER was run to get the pair of DO channels together.

- HEADER EDIT was run to ensure formats and units are correct

- Standards check was run and no problems were found.

- Bottle data were exported to a spreadsheet and no problems were found.

- A header check were run. No problems were found.

- A cross-reference listing was produced for the CHE files.

CRUISE SUMMARY – CTD

| CTD# | Make | Model | Serial# |  | Used with Rosette? |  | CTD Calibration Sheet Competed? |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0443 |  | Yes |  | Yes |  |  |  |
| Calibration Information - 0443 |  |  |  |  |  |  |  |  |  |  |
| Sensor |  |  |  | Pre-Cruise |  |  |  | Post Cruise |  |  |
| Name |  | S/N |  | Date |  | Location |  | Date | Location |  |
| Temperature |  | 4700 |  | 5Jan2022 |  | Factory |  |  |  |  |
| Conductivity |  | 3531 |  | 8Feb2022 |  | Factory |  |  |  |  |
| Secondary Temp. |  | 4888 |  | 14Jan2022 |  | Factory |  |  |  |  |
| Secondary Cond. |  | 4513 |  | 8Feb2022 |  | Factory |  |  |  |  |
| Transmissometer |  | 1185DR |  | 20 Feb 2023 |  | Factory |  |  |  |  |
| Transmissometer |  | 1883DG |  | 20 Feb 2023 |  | Factory |  |  |  |  |
| SBE 43 DO sensor |  | 4372 |  | 27Mar2023 |  | Factory |  |  |  |  |
| SeaPoint Fluor. |  | 4186 |  | 17Feb2022 |  | Factory |  |  |  |  |
| Pressure Sensor |  | 0443 |  | 23Mar2022 |  | Factory |  |  |  |  |
| Altimeter |  | 76341 |  | 10Feb2021 |  | Factory |  |  |  |  |