#### 2024-002

CCGS John P Tully

#### Dimethylsulfide (DMS) & Dimethylsulfoniopropionate (DMSP) Report

#### August 08 - August 24, 2023

Prepared by Michael Arychuk

- Sample Collection

Samples were collected from stations  P2, P4, P12, P16, P20 and P26 for DMS, DMSPD (dissolved) & DMSPT (total).

  - DMS

Eleven water samples (100m, 75m, 50m, 40m, 30m, 25m, 20m, 15m, 10m, 5m, surface) were collected at station P2.  Fifteen samples (600m, 300m, 200m, 150m, 100m, 75m, 50m, 40m 30m, 25m, 20m, 15m, 10m, 5m, surface) were collected at P4,  P12, P16, P20 & P26. Duplicates were taken at 20m at all stations.  In all cases, samples were collected in 150mL ground glass stoppered bottles and stored in a fridge, in the dark and removed one at a time before analysis.

  - DMSP

Seven samples (600m, 200m, 150m, 30m, 20m, 5m, surface) for both DMSPD and DMSPT were collected at P4, P12, P16, P20 & P26.  Four samples (100m, 20m, 5m, surface) were collected at P2. Duplicates were taken at 20m.

- Analysis

2.1 DMS

Approximately 10mL of sample was loaded onto the stripper and purged with UHP Helium for 5 minutes at ~120mL/min. The DMS is extracted from the seawater and absorbed onto a 1/16” Teflon tubing “trap” immersed in liquid nitrogen.  After 5 minutes the trap is desorbed at ~100oC (with a dewar containing hot water) onto a HP-1 capillary column which eluted into a Flame Photometric Detector (FPD).  All samples were run as soon as possible after being collected.

2.2 DMSPD

Approximately 50-75mL of seawater was allowed to flow directly from the niskin into a filtration funnel containing a 0.7m GF/F filter.  The first 3.5mL was collected in a 15mL or 5mL polypropylene tube.  The entire 3.5mL was then transferred into a 5ml, glass, serum bottle and 50L of a 50% sulphuric acid/water solution was added.    The sample was then crimp sealed and stored in the dark and at 4˚C where it would be analysed back at IOS at a later date.

2.3 DMSPT

Exactly 3.5mL of seawater was collected directly from the niskin into a 15mL or 5mL polypropylene tube. The entire 3.5mL was then transferred into a 5ml, glass, serum bottle and 50L of a 50% sulphuric acid/water solution was added.    The sample was then crimp sealed and stored in the dark and at 4˚C where it would be analysed back at IOS at a later date.

- Calibration

3.1 DMS

A four to six level calibration table was used for calculating the concentrations of DMS.  The standards were prepared in MQ water (brought from the lab and stored in 4L amber glass bottles) for stations P2 & P4 and then prepared in 2000m or deeper seawater (also stored in amber glass bottles kept at 4˚C) for stations P12, P16, P20 & P26.  Standards were run under the same conditions, as described above, for the samples and the calibration curve was valid for 12 hours.  If analysis exceeded 12 hours, a continuing calibration standard was run to ensure the calibration curve was still within acceptable limits.

- Quality Control

4.1 DMS

System blanks and duplicates were run approximately every 15 samples to ensure the system remained free of contamination and had acceptable reproducibility.

4.2 DMSP

Blanks and duplicates were collected at every station.  Blanks were done by either using MQ water, some of the bottled SQ water brought from the lab or using some deep sea water (2000 m or deeper) and treating them as actual samples.    For example, in the case of DMSPD,   3.5mL was collected from a separate funnel and for DMSPT    3.5mL  was added directly to the polypropylene tube.  Like the samples, they were then transferred into a 5ml, glass, serum bottle and 50L of a 50% sulphuric acid/water solution was added.    The blank was then crimp sealed and stored in the dark and at 4˚C where it would be analysed back at IOS at a later date.

- Data & Results

  - DMS

The sample data for this cruise was very typical of a Winter cruise.  Results were very low at all stations and there was no DMS below 75m in any of the stations.

  - DMSP

The DMSP samples were collected for all stations.  Samples were run back in the lab in early February 2024.

- Conclusions

6.1 DMS

All samples were only run on the 8890 GC system this cruise and the old 6890 GC system has been officially retired.

  - DMSP

Sample analysis went well without any issues.