## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2024-004

Agency: Ocean Sciences Division (OSD)

Location: Salish Sea

Project:  Salish Sea Biophysical Survey

Chief Scientist: Bluteau C.

Platform: Vector

Date: 10 April 2024 – 16 April 2024

Processed by: Germaine Gatien

Date of Processing: 27 May 2025 –17 June 2025

Number of original HEX files: 77

Number of CTD files:  	77	Number of CTD casts processed: 77

Number of bottle files: 	27	Number of bottle casts processed: 27

Number of TSG files:   1	Number of TSG files processed:  1

# INSTRUMENT SUMMARY

An SBE911+ CTD #0443 was mounted in a rosette and attached were 2 Wetlabs CSTAR transmissometers (#1185DR and #1883DG), a SBE 43 DO sensor (#4372), a SeaPoint fluorometer (#4186), a Wetlabs ECO fluorometer (#2216), a Biospherical PAR sensor (#4565), a reference SPAR sensor (#20518)  and an altimeter (#76341).

The data logging computer was a Lenova ThinkCentre.

Seasave version 7.26.7.121 was used for acquisition.

The salinometer used at IOS was a Guildline model 8400B Portasal, serial # 73274.

An IOS rosette with 24 10L bottles was used.

A thermosalinograph (SeaBird 45 S/N 3411) was used.

# SUMMARY OF QUALITY AND CONCERNS

The Science Log was in digital format. The header page was missing information about the Oxygen kit and TSG and there was no indication that a Surface PAR was in use. There were a number of errors in the rosette sheets. This cruise was run during a transition period in the leadership of this program.

The sounder readings were generally reported at the bottom as well as at the beginning of casts. This was very helpful as the readings at the bottom can be compared with the difference between the maximum depth sampled by CTD plus altimetry reading at the bottom. The water depth is of most interest when the CTD is at the bottom. For this cruise there were 10 cases where that difference was >5m; for 9 of those cases using the sounder reading at the bottom compared better than the reading at the surface, while in 1 case it looked better when the end reading was used. It is likely that these differences are due primarily to shoaling during the cast, though in one case the water depth in the log looks like it was not changed from the previous cast.

Recalibration of dissolved oxygen and salinity was based on results of cruise 2024-006 which had deeper sampling, but a larger offset was applied to dissolved oxygen fit based on comparison with titrated samples. This dissolved oxygen sensor has an unusual fit against titrated samples, but the same behaviour has been noted during other cruises that used the sensor in 2024, and the factory service in late 2024 also showed an unusual fit before service. The pressure recalibration is based found for this cruise is in agreement with results from several other cruises using the same CTD.

The SBE DO sensor has a fairly long response time, so data accuracy is not as high when it is in motion as it is during stops for bottles. This will be especially true when vertical DO gradients are large. To get an estimate of the accuracy of the SBE DO data during downcasts (after recalibration) a rough comparison was made between downcast SBE DO and upcast titrated samples. Some of the difference will be due to problems with flushing of Niskin bottles and/or analysis errors, so the following statement likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data for this cruise are considered, very roughly, to be:

±0.20 mL/L from 0-100db

±0.12 mL/L from 100db to 200db

±0.03 mL/L below 300db.

For details see file 2024-004-dox-comp3.xls.

The Thermosalinograph performed reasonably well for most of the cruise, but on April 13th there was a section in the southern Georgia Strait where salinity was spiky and then dropped to very low levels. From about 13:00 UTC there are spiky sections, but given the location near the Fraser River, those data were left unedited. Between 14:51 UTC and 20:00 UTC salinity is extremely low and it rose abruptly by 10psu between 2 records at 20:00. Salinity and conductivity data were removed from 14:51 to 20:00. No problems were noted in the temperature data.

The TSG salinity was generally lower than that from the CTD salinity and the 1 loop sample taken, but the differences vary too much to provide a reliable basis for recalibration.

The comparison of temperature with CTD data enabled an estimate of heating in the loop. Channel Temperature:Primary was prepared as a proxy for intake temperature by adding 0.20C to Channel Temperature:Lab.

# PROCESSING SUMMARY

GENERAL

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

- The digital daily log and rosette log sheets were obtained and checked for comments.

- Salinity, Nutrients, Extracted Chlorophyll and Dissolved Oxygen data were obtained in QF spreadsheets.

- The cruise summary sheet was completed.

- The history of the pressure sensor, conductivity and dissolved oxygen sensors were checked.

- The calibration control files were obtained; since they stayed the same through out the cruise 2024-004-0001.xmlcon was used for processing.

- Initial tests showed no signal in the PAR channel, so tests were run to make sure the right PAR sensor parameters were used. These tests confirm that the problem is not due to mis-identification of the PAR in use.

##### BOTTLE FILE PREPARATION

The ROS files were created using 2024-004-ctd.xmlcon. The hysteresis correction was not selected since there were no deep casts; the Tau correction was selected.

The PAR channel was converted in case there are some later casts with a signal.

The ROS files were converted to IOS format

The ROS files were put through CLEAN to create BOT files.

The station name format was changed in casts #2 and 58 to match the log entries.

Temperature and salinity were plotted for all BOT files to check for outliers, none were found.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. Sample numbers were added based on the rosette logs.

A problem was found in cast #91. The samples identified as #311 should have been #310 since Niskin #4 was not fired. This affects DO and NUTS samples. The analysts were notified and updated QF reports.

The ADDSAMP file was then sorted on event number & sample number and converted to CST files.

The CST files will form the framework for the bottle files.

The ADDSAMP file was also used to add sample numbers to the BOT files – output *.SAM.

The SAM files were bin-averaged on bottle # and called *.SAMAVG.

Next, the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were saved in file 2023-004-bot-hdr.txt.

SALINITY

Salinity analysis was obtained in file QF2024-004_SAL_30Nov2023.xlsx. The analyses were carried out in a temperature-controlled lab within 10-14 days of collection. The files were simplified and saved as 2023-069_SAL.csv. That file was then converted to individual SAL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2024-004_OXY_13Nov2024.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and saved as 2024-004_OXY.csv. That file was converted into individual *.OXY files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2024_004_NUTS_2024-04-20.xlsx which includes flags, comments and a precision study. The spreadsheet page with the final data was simplified and saved as 2024-004_NUTS.csv. That file was converted into individual *.NUTS files.

CHLOROPHYL

The chlorophyl data were obtained in spreadsheet QF2024_004_CHL_28Nov2024.xlsx which includes flags, comments and a precision study. The spreadsheet page with the final data was simplified and saved as 2024-004_CHL.csv. That file was converted into individual *.CHL files.

The SAL, CHL, OXY, and NUTS files were merged with CST files in 4 steps. The files were then put through CLEAN to reduce the headers to File and Comment sections only. These files were ordered on sample number, but the SAMAVG files are ordered on bottle number, so the MRGCLN1 files were reordered on Bottle_Number and saved as *. MRGCLN1s. The MRGCLN1s files were then merged with SAMAVG files using merge channel Bottle_Number.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. The sampling file was very useful in this process.

On the rosette log there are said to be salinity samples from event #1 but none were found and cast #91 does not match exactly as noted before. Otherwise, the exported file contains all samples indicated in the rosette log, plus at least one extra.

At this point samples flagged by the oxygen analyst with comment “ALL:” were examined to see if other variables needed flag 3 or 9 attached. Only nutrients were affected. The NUT files were updated and merges rerun from that point.

For cast #18 since the first planned bottle was not actually fired, that record was removed from the cast file.

After those changes, the complete MERGE process was repeated.

The spreadsheet was checked again and looks ok.

A header check was run and no problems were found.

##### Compare

Salinity

This cruise had 22 salinity samples taken between 3m and 351m. There is a lot of noise in the fit, not unusual for this region, given active mixing at some sites and poor flushing of Niskin bottles expected at others. The fit against pressure below 100db (excluding points with standard deviation in the CTD salinity >0.0008psu) is reasonably flat, with the primary low by 0.001psu and the secondary low by 0.0027psu.

When a different approach is taken by gradually removing points to achieve a flat fit, the results are similar for the primary which was low by an average of 0.0009psu. The secondary looks better using this approach, being low by 0.0019psu. There was some suspicion of pressure dependence in the secondary conductivity when it was used during 2024-069, which might explain this difference. However, there are too few samples and they are too shallow to confirm that.

These results are not considered very reliable because there are few samples, they are quite shallow and there is a very large scatter in the fits. The results are quite different from those of 2024-006 when there was a lot of deep sampling; for that cruise the primary was found to be high by 0.0015psu and the secondary fairly close to the primary. It is likely that incomplete flushing of bottles accounts for the apparent picture of CTD salinity reading lower than it did during 2024-006.

For more detail see file 2024-004-sal-comp1.xls.

None of the outliers were sufficiently out of line to recommend assigning quality flags.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

There was one severe outlier that had been flagged 4 by the analyst. Flag 5 and a pad value were recommended for sample #175 as this is clearly too high a value to be found at 50m. (Event #54.) The analyst agreed and updated the QF spreadsheet for dissolved oxygen.

The fit of differences between SBE DO and titrated samples versus SBE DO was examined. The fit is reasonably tight when a few outliers are excluded, but there are no DO with values  <1.84 mL/L, so the offset to choose for the fit is far from obvious. Seabird report that the offsets for fits of CTD versus bottle DO are not expected to drift, while the slope does drift. The offsets found for some recent cruises were 0.03mL/L and 0.05mL/L. The 0.03mL/L choice came from a Line P cruise with a lot of deep sampling, including samples in the OMZ.

For this cruise, the lowest DO values are mostly from the bottom of casts. Bottles from the bottom of casts often read higher than the CTD DO due to poor Niskin flushing, if DO is decreasing with pressure.. However, 7 of these casts had very low DO gradients near the bottom, so flushing errors are likely not significant. For those 7 the CTD read low by an average of 0.026 (std dev 0.012mL/L), while for 7 casts where the vertical gradients were high, the average was 0.057mL/L (std dev 0.043mL/L). We would expect those differences to be slightly larger as DO approaches 0 mL/L. The choice of 0.03mL/L as offset looks like a good choice.

The fit when the offset was forced to 0.03mL/L and outliers were excluded based on residuals was:

SBE DO Corrected = 0.924*SBE DO original + 0.03

The slope is significantly lower than found during 2024-006, likely due to the fact that most of the bottles were fired during upcasts and in significant vertical gradients; incomplete flushing leads to bottle samples coming from lower in the profile where DO values are generally lower than those seen by the CTD DO sensor. So using the 2024-006 choice for slope looks most appropriate.

The best choice for recalibration appears to be to use the 2024-006 result:

SBE DO Corrected = 0.987*SBE DO original + 0.03

Fluorescence

COMPARE was run with pressure as the reference channel.

Uranine tests in February 2024 found a voltage threshold of 0.001 for the SeaPoint Fluorometer and 0.058 for the ECO.

It was discovered late in processing that the offset in the configuration file was from an older calibration with a value of 0.044. This is very close to the 2024 value, and given the jitter found in ECO fluorometers, increasing the offset would likely lead to a lot of negative values. So no correction was applied.

Dark values in these data are both ~0.05ug/L for both fluorometers but they are noisy and in this region we do not expected 0.

Tests were done to see how well the ECO and Seapoint fluorometers compared and differences were found to be slight.

As noted during 2024-006 using dark values for offsets will have to be done carefully with the ECOs because values are often negative in very deep water. Deep water may affect the ECO in some way that is not relevant to shallow water. Since the technicians now do uranine tests and can set the configuration files for ECOs appropriately, it is not surprising to find the two data sets are close without correction. However, the uranine test values may lead to many negative results for the ECO, so a slightly smaller correction may be wise; the correction applied to these data was -0.044 rather than -0.058 and worked well.

The ECO is slightly closer to CHL samples though still significantly lower than CHL.

The 2 types of fluorometers continue to perform in a similar fashion, as found in previous studies.

For more details see document 2024-004-fl-chl-comp1.xls.

##### Conversion of Full CTD Files from Raw Data

All files were converted using files 2024-004-0001.xmlcon. The hysteresis correction was not selected since there were no deep casts; the Tau correction was selected. A few casts were examined and all expected channels are present.

NMEA was not working for events 1 and 2; details were added to the headers based on the log entries.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the

full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

CELLTM was run using the default settings (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

Tests were run to ensure the previous 3 steps worked properly.

- A few casts were checked and no spikes were found in downcast P, T, C.

- The DO profile is well aligned with temperature.

- CELLTM is expected to lower salinity where temperature is decreasing (cell warms water raising conductivity but temperature sensor reading is correct, so salinity is overestimated). Conversely, it raises salinity where temperature is increasing (conductivity too low). This brings downcast and upcast traces closer together, so T-S downcast and upcast plots are in better agreement after this step. For this cruise CELLTM clearly worked well  on the well-mixed casts, but some of the casts have very complex gradients and/or noisy upcasts, so it is hard to judge.

DERIVE was run a second time on a few of the deeper casts to find the differences between the pairs of temperature, conductivity and salinity channels.

The sensors were used during cruise 2024-001 and 2024-006. Differences were found to be small during 2024-006. For this cruise only shallow casts are available so results are not as reliable.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2024-001-0030 | 550 | -0.0003 | -0.00038 | -0.0038 | High, V. Steady |
| 2024-001-0050 | 550 | -0.0004 | -0.00033 | -0.0037 | High, F. Steady |
| 2024-001-0074 | 550 | -0.0003 | -0.00032 | -0.0031 | High, F. Steady |
| 2024-004-0054 | 360 | -0.0003 | -0.00015 | -0.0013 | High, Moderate |
| 2024-004-0070 | 310 | -0.0004 | -0.00016 | -0.0013 | High, Steady |
| 2024-004-0081 | 340 | -0.0003 | -0.00015 | -0.0013 | High, Steady |
| 2024-069-0025 | 500 | -0.0010 | -0.00014 | -0.0008 | High, Moderate |
|  | 1000 | -0.0006 | -0.00017 | -0.0015 | “ |
|  | 1500 | -0.0003 | -0.00020 | -0.0020 | “ |
|  | 2000 | -0.0001 | -0.00021 | -0.0024 | “ |
| 2024-069-0036 | 500 | -0.0005 | -0.00014 | -0.0011 | High, Moderate |
|  | 1000 | -0.0004 | -0.00016 | -0.0016 | “ |
|  | 1500 | -0.0002 | -0.00017 | -0.0019 | ‘ |
|  | 2000 | +0.0001 | -0.00018 | -0.0023 | ‘ |
| 2024-069-0069 | 500 | -0.0006 | -0.00014 | -0.0011 | High, V. Steady |
| 2024-069-0180 | 500 | -0.0008 | -0.00024 | -0.0018 | High, V. Steady |

Temperature differences are small and there is no indication of temporal drift.

Conductivity and salinity differences are smaller than those of 2024-001 and close to the 500m values for 2024-069. There was some evidence during 2024-069 that there might be slight pressure dependence in the secondary conductivity.

Cruise data processing is not complete for 2024-006 which was run immediately after 2024-004 and used CTD #0443 for part of the cruise. Differences between channels are reported to be small.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SeaBird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

A header check and cross-reference list was run.

The station name format was changed in casts #2. 32, 58 and 66 to match the log entries.

There are some negative values in the ECO fluorometer channel; there is a jitter that can sometimes lead to such results, but for this cruise the only negative values found were in surface spikes. This will be checked again after bin-averaging and changed to 0, as necessary.

A surface check was run and found a mean surface pressure was 2.8db, with a minimum of 1.6db and a maximum of 3.89db, which is a little deeper than usual for the Vector..

The lowest pressure found during the cruise was at the end of event #18. The pumps were running with pressure at 0.687db. All channels, especially conductivity, transmissivity and fluorescence, suggest that the CTD was moving out of water at about 1db. This is in line with observations during 2024-006 and 2024-069 when pressure was recalibrated by subtracting -1.1db.

The track plots were added to the end of this report.

Checks were made that the water depth entries are as accurate as possible. The sounder readings were generally reported at the bottom as well as at the beginning of casts. This was very helpful as the readings at the bottom can be compared with the difference between the maximum depth sampled by CTD plus altimetry reading at the bottom (referred to as the check value).  Most casts did go close enough to the bottom to do this check. The water depth reading of most interest in the data files is for the time when the CTD is at the bottom. For this cruise there were 10 cases where the check value was >5m; for 9 of those cases the sounder reading at the bottom gave a smaller check value than with the reading at the surface, while in 1 case it looked better left unchanged. It is likely that these differences are due primarily to shoaling during the cast, but in one case the water depth in the log just looks like it was not changed from the previous cast.

The water depth headers were adjusted in 9 profile files and 2 bottle files (#49 and 69). Most changes were small indicating that the sounder was working well.

##### Shift

Fluorescence

SHIFT was run on the fluorescence channel with an advance of 24 records. Plots showed this setting was appropriate.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier using Seabird software. No further alignment is needed for the DO concentration channel.

Conductivity

Tests were run on 3 casts to determine the best setting to align conductivity and temperature by judging the effect on salinity as seen in T-S space. SHIFT was run on all casts using -1.0 records for the primary conductivity and -0.8 records for the secondary. Salinity was recalculated for both channels.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

##### Other Comparisons

Previous experience with these sensors – The CTD was used on many cruises before this one, though the secondary conductivity and the DO sensor were used on only 1 previous cruise since last factory service. 2 cruises after this one have been processed. The best calibrations came from cruise 2024-006 in June; results from that cruise were:

- The pressure was recalibrated by subtracting 1.1db for 2024-006 in late April and 2024-069 in mid June.

- The primary salinity was recalibrated by subtracting -.0015psu. The secondary was used for only one cast and was recalibrated by subtracting 0.008psu to bring it into line with the primary.

- DO values were found to be slightly high in the 2024 cruises already processed, with the best estimate being that it was high by about 1.5%.

A later cruise suggested that the secondary conductivity sensor might have some pressure dependence.

Historic ranges – Almost all the temperature and salinity fell within the local climatology. Temperature was slightly high around 100m for cast #82. Salinity was slightly low at depth at the mouth of Juan de Fuca on the south side. Two casts had salinity that looked a little low from 200m down compared to the Southern Strait of Georgia climatology, but well within the climatology for the Gulf Islands. This includes station 56 which often looks closer to the Gulf Islands range.

Post-Cruise Calibration – Factory checks were available for conductivity from December 2024 and conductivity from November 2024. There were significant errors in dissolved oxygen (reading high by roughly 5%) while salinity values looked to be within 0.002psu. There were many cruises between 2024-004 and those factory checks.

##### DETAILED EDITING

The primary T-S pair were chosen for editing and eventual archiving. For this cruise both primary and secondary channels were very similar, but there is some suggestion from a later cruise that there may be some drift and/or pressure dependence in the secondary conductivity sensor.

All DEL files were uploaded into the CTD-QC File Processor to obtain AI generated data quality flags.

The descent rate of the CTD was noisier than usual for this region, presumably due to weather conditions. This led to heavier editing than usual for some casts..

CTDEDIT was used to remove records corrupted by wakes where the descent rate was low; salinity was cleaned. Most casts required some editing.

Casts that required no editing were: 32, 35, 86.

The edited files were copied to *.EDT.

After editing T-S plots were examined for all casts; there remain many small unstable features but they are in areas when they may well reflect real conditions. No further editing was applied.

The EDT files were zipped and sent to the CTD-QC site.

##### Fluorescence Processing

A median filter with width 11 was applied to the fluorescence channel.

##### Recalibration

Silicate does not require recalibration in bottle files since there Salinity was >25psu in all bottle files.

Pressure looked consistently low by 1.1db during cruises before and after this one.

Comparisons with bottles from this cruise were limited. The primary salinity was found to be low by about 0.001psu while in later cruises it was found to be high by about 0.0015psu. This is likely due to incomplete flushing of Niskin bottles during this cruise. Similarly, the secondary was found to be low by 0.0027psu whereas it was thought to be low by about 0.0008psu during 2024-069.

Based on later cruises there may be more drift in the secondary salinity and the secondary conductivity sensor may have some pressure dependence.

First, recalibrations were applied to Pressure, Salinity;T0:C0 and Dissolved Oxygen recalibration based on those applied to 2024-006 data. No correction is necessary for Salinity:T1:C1 because it is reading lower than the primary by 0.0017psu, so is very close to the primary after the primary is recalibrated.

File 2024-004-recal_test.ccf was prepared to apply the following corrections:

Pressure Corrected = Pressure -1.1

SBE DO Corrected = 0.987*SBE DO original + 0.03

Salinity:T0:C0 Corrected = Salinity:T0:C0 – 0.0015psu

These corrections were first applied to the SAM and MRGCLN2 files.

Pressure was found to have been corrected appropriately.

COMPARE was rerun on salinity and dissolved oxygen to check recalibration results.

Salinity appears to be low by an average of 0.0025psu, but this is presumed to be a measure of the effect of incomplete flushing of bottles.

The fit for dissolved oxygen is puzzling. The fit is much flatter after recalibration which is good, but the offset looks inappropriate as it leads to CTD reading lower than bottles. That is not expected except in areas where gradients reverse. In particular, in deep water we expect the CTD to be quite close to bottles as flushing errors are small in low gradient areas. While there is no clear explanation for why the offset should be different for this cruise, it produces more reasonable results.

File 2024-004-dox-comp2x.xls shows results using offset +0.03mL/L.

File 2024-004-recal1.ccf was created to apply the following corrections:

Pressure Corrected = Pressure -1.1

SBE DO Corrected = 0.987*SBE DO original + 0.06

Salinity:T0:C0 Corrected = Salinity:T0:C0 – 0.0015psu

While the offset in the instrument calibration is considered to be constant, there may be a further effect due to incomplete flushing of bottles that also contributes to the offset for this case.

File 2024-004-dox-comp2.xls shows results using offset +0.06mL/L.

Using this correction, the SBE DO looks like it is reading slightly high, as expected due to incomplete flushing.

The overall fit for all cruises using this oxygen sensor are odd, so unusual results are not surprising especially for cruises with relatively shallow sampling. During servicing in November 2024 the sensor was found to be reading high by up ~5%, while it is usually found to read low. However, the sensor was used during many cruises between this one and that service visit, so some drift may well have occurred after 2024-004.

Calibrate was then run on the EDT files.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and delayed response and noise in CTD data.

Downcast recalibrated files (bin-averaged to 0.5m-bins) were thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles. Outliers were removed based on residuals in a fit against DO. No recalibration appears necessary.

Plots were then made of fits of differences versus pressure using the same data to judge accuracy in various depth ranges. This is a very rough comparison and likely does not distinguish sensor errors from other error sources such as those in analysis and incomplete flushing of bottles. So errors are likely overestimated.

Downcast (CTD files) Oxygen:Dissolved:SBE data for this cruise are considered, very roughly, to be:

±0.20 mL/L from 0-100db

±0.12 mL/L from 100db to 200db

±0.03 mL/L below 300db.

For details see files 2024-004-dox-comp3.xls

No further recalibration is necessary.

##### Bin Average of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined. Small unstable features remain in a few files, but they are in places where unstable features may well be real.

Profile plots were examined to see if there any problems. None were found. While PAR:Reference appeared to have no signal for a few night-time plots, but there is a small signal right at the surface that was removed in processing.

##### Final CTD File Steps (REMOVE and HEADEDIT)

A second SBE DO channel (with umol/kg units) was added.

Dissolved Oxygen saturation was calculated and near-surface values ranged from 75% to 125%; this is typical of this area where active mixing causes low values in Juan de Fuca Strait and narrow channels, while plankton blooms raise value in other areas. This is not helpful in evaluating the accuracy of the DO sensor.

REMOVE was run to remove the following channels from all but cast #132:

Scan_Number, Temperature:Secondary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T1:C1, PAR and Flag and Prediction_Flag.

REMOVE was run to remove the following channels from cast #132:

Scan_Number, Temperature:Primary, Conductivity:Primary, Oxygen:Voltage:SBE, Oxygen:Dissolved:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T0:C0, PAR and Flag and Prediction_Flag.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the comments about processing.

The Standards Check routine was run and no problems were found.

The Header Check was run; no problems were found.

Profile and T-S plots were examined. No problems were found.

The sensor history was updated.

##### Final Bottle Files

- The MRGCOR1 files were put through SORT to order on increasing pressure.

- REMOVE was run to remove channels: Scan_Number, Temperature:Secondary, Conductivity:Secondary, PAR, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T1:C1, Flag.

- A second SBE DO channel with mass units was added for both the CTD DO and Titrated DO.

- REORDER was run to get the pair of DO channels together.

- HEADER EDIT was run to ensure formats and units are correct

- Standards check was run and no problems were found.

- Bottle data were exported to a spreadsheet and a few problems were found in event #1. Those were fixed and the data were exported and no further issues found.

- A header check were run. No problems were found.

- A cross-reference listing was produced for the CHE files.

##### Thermosalinograph processing

There was 1 thermosalinograph file.

There was 1 loop sample. There was no flow meter or intake thermistor. The intake is at about 2m.

a.) Checking calibrations

The configuration file parameters were checked and are correct.

b.) Conversion of Files

The file was converted to CNV format.

The file was converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels.

A time-series plot was produced. There is a noisy patch on April 13th that lasted a few hours. This corresponds to positions in the southern Strait of Georgia. There is no corresponding noise in temperature.

The track plot looks fine and was added to the end of this report.

c.)  Comparing TSG to near-surface CTD cast data

CTD files were thinned to 2db and exported to a spreadsheet. There were 72 casts that overlapped with TSG data and had data available at 2db. This file was named 2024-004-CTD-TSG-Comparison.xlsx.

The ATC file with the TSG data was opened in EXCEL. Unnecessary columns were removed.

Median and standard deviations of T and S were calculated over 5 points (2 minutes). Those values were copied and Paste Special was used to save the calculations before records are removed.

Those data were then thinned to the times of the CTD files in the spreadsheet.

The thinned ATC file data were added to the XLSX file above. Checks were made occasionally that the matches were correct by monitoring latitude and longitude differences.

Comparisons were then made between the 2 data sets.

|  | Ttsg-Tcrd | Stsg-Sctd |
| --- | --- | --- |
| average | 0.2025 | -1.4371 |
| median | 0.1980 | -0.5427 |
| stdev | 0.0388 | 2.8535 |
| max | 0.2771 | 0.0014 |
| min | -0.0100 | -13.5343 |

The TSG temperature is high by a median of 0.198Cº and standard deviation of 0.039Cº. That is typical of the amount of warming expect in the loop for the Vector. We expect that the TSG salinity will read a little low, but this is lower than usual and variability is extremely high. A plot of differences versus event # shows that the early part of the cruise looks as expected.

Given the unusual pattern of differences in salinity, the data were sorted on standard deviation in the TSG temperature; limiting the selection to the 40 casts with the lowest standard deviation in TSG temperature or the 20 lowest both gave a difference of 0.199Cº. So a proxy for intake temperature can be created by subtracting 0.20Cº.

The same process was used to try to assess the pattern in salinity but it was not enlightening, with high variability.

| Median Differences | Ttsg-Tctd | Stsg-Sctd |
| --- | --- | --- |
| 40 casts with lowest std dev in Temp | 0.1989 | -0.4168 |
| 20 casts with lowest std dev in Temp | 0.1989 | -0.2871 |
| 40 casts with lowest std dev in Sal | 0.1966 | -0.7652 |
| 20 casts with lowest std dev in Sal | 0.1976 | -0.4236 |

Reducing the selection had little effect on temperature differences. However, it had an unexpected effect on salinity with differences getting larger when 40 casts were used and slightly smaller if 20 casts were selected. So the casts were divided into regions ignoring the standard deviations in the TSG data.

| Range | # of casts | SALtsg-SALctd | Std Dev |
| --- | --- | --- | --- |
| Casts 1-45 | 33 | -0.2965 | 0.100139 |
| Casts 46-70 | 18 | -1.5661 | 5.006459 |
| Casts 71-98 | 21 | -1.1115 | 0.016244 |

The middle group look out of line and salinity looks bad in the time-series plot.

Two possibilities may explain the differences in these sections.

- The first is that we are comparing with CTD downcast salinity which may be a little higher due to wakes from the soak period. This is most likely to be a significant issue in stormy seas. We can guess where those are by looking at the descent rate of the CTD. It was particularly noisy near the surface between casts 50 and 53.

- The second is that the loop may be drawing water from a little higher in the water column than the data in the CTD files. Bottle stops at the end of casts provide another source of salinity data and the 30s wait before firing should give the CTD salinity time to equilibrate. So salinity was extracted from the CTD files between 1 and 2m in the bottle files and compared to the TSG values. They are lower than what the CTD measures. There is a mis-match in time but it does show there is a significant salinity gradient near the surface.

| Events | # of casts | TSG Sal -Upcast CTD Sal | Std Dev |
| --- | --- | --- | --- |
| 1-45 | 10 | 0.0060 | 0.0236 |
| 46-70 | 4 | -0.0933 | 0.0621 |
| 71-98 | 8 | -0.0019 | 0.0053 |

The median value for the whole group is -0.006psu.

The data from the middle group look unreliable and an examination of the salinity plot shows a drop in values between 08:00 and 20:00 on April 11th. The early part may well be real, but between 14:00 and 20:00 clearly looks bad and values return to high values very abruptly, between 2 records 30s apart.

d) Comparison of TSG and 1 Loop Sample

1 salinity samples was taken from the loop as Event #39.

The sample was saved as sample #5004; the analyst changed it to #5039.

The time was 3:33:33 on April 13th. Position:48 54.733N 122 59.989W  48.9062N 122.999W

Salinity Sample =29.7911psu

The TSG reading at that time was S=29.4818 Position:  48.91546N 123.0000W

So the TSG is reading lower than the loop sample by 0.307psu.

e) Calibration History

This TSG was used on 3 previous cruises since calibration. On cruise 2024-001, in early March, temperature was believed to be high by ~0.026Cº. There were few points of comparison and differences between TSG and CTD salinity, while large, were too variable to enable an estimate for recalibration. Most casts were in inlets.

f.) Calibration

Temperature in the lab is higher than CTD temperatures by about 0.2Cº. This is a typical value for heating in the loop in spring. This value will be used to derive a proxy for intake temperature.

Salinity will not be recalibrated due to insufficient and contradictory evidence.

g) Editing

Data was relatively clean (i.e. no spikes), except for a section on April 13th where salinity was very low. While this seemed possible since it was fairly close to the Fraser River, it ended very abruptly between two readings 30s apart with salinity that increased by 10psu. The cause is unknown.

2024/04/13 19:59:54   9.6221   9.8221  14.6088   1.71228 104.83327  49.14276  -123.61372

2024/04/13 20:00:24   9.6230   9.8230  24.3551   2.73285 104.83361  49.14282  -123.61380

CTDEDIT was used to remove salinity and conductivity data between 14:51UTC and 19:59 UTC April 13, 2024. There are data in the hour before that section that are also questionable, but given the ship was near a very active region and possible plume, and the differences are not as large as seen later, it was left unedited.

REMOVE was used to remove the following channels: Scan Number and Flag channels.

g.) Recalibration

Add Channels was used to add Channel Temperature:Lab with values set equal to Temperature:Primary.

Calibrate was run using file 2024-004-tsg-recal1.ccf to subtract 0.20Cº from Temperature:Primary.

h.) Preparing Final Files

HEADER EDIT was used to change the DATA DESCRIPTION to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats.

The format for salinity was changed from the usual 4 decimal places to 3, to indicate lower quality than usual.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and time-series and all look fine.

The cruise plot was added to the end of this report.

h.) Preparing Final Files

HEADER EDIT was used to change the DATA DESCRIPTION to THERMOSALINOGRAPH and to change channel names to standard names and formats.

The TSG sensor history was updated.

PARTICULARS

1 & 2. NMEA not available. Positions/time added based on log.

11. Garburator – lots of trash in water.

15. Forgot to turn off pump.

36. Lots of currents, well mixed so shorter flushing time

43. The CHL on bottle 10 was missed. Sample number 145A.

60. Stayed 100m off the bottom.

91. Niskin #4 was not fired. There are no samples #311.

2024-004 CRUISE SUMMARY – CTD & TSG

| CTD# | Make | Model |  | Serial# |  | Used with Rosette? |  | CTD Calibration Sheet Competed? |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ |  | 0443 |  | Yes |  | Yes |  |  |
| Calibration Information - 0443 |  |  |  |  |  |  |  |  |  |  |
| Sensor |  |  |  |  | Pre-Cruise |  |  |  | Post Cruise |  |
| Name |  |  | S/N |  | Date |  | Location |  | Date | Location |
| Temperature |  |  | 2106 |  | 15Feb2023 |  | Factory |  |  |  |
| Conductivity |  |  | 2754 |  | 14Feb2023 |  | Factory |  |  |  |
| Secondary Temp. |  |  | 5130 |  | 18Mar2023 |  | Factory |  |  |  |
| Secondary Cond. |  |  | 1766 |  | 18Jan2023 |  | Factory |  |  |  |
| Transmissometer |  |  | 1185DR |  | 11Jan2024 |  | IOS |  |  |  |
| Transmissometer |  |  | 1883DG |  | 11Jan2024 |  | IOS |  |  |  |
| SBE 43 DO sensor |  |  | 4372 |  | 27Mar2023 |  | Factory |  |  |  |
| PAR |  |  | 70613 |  | 24Feb2024 |  | Factory |  |  |  |
| SPAR |  |  | 20518 |  | 16Feb2024 |  | Factory |  |  |  |
| SeaPoint Fluor |  |  | 4186 |  | Aug 2023 |  | Factory |  |  |  |
| WetLabs ECO Fluor |  |  | 2216 |  | 8Mar2017 |  | Factory |  |  |  |
| Pressure Sensor |  |  | 0443 |  | 23Mar2022 |  | Factory |  |  |  |
| Valeport Altimeter |  |  | 76341 |  | 10Feb2021 |  | Factory |  |  |  |

| Calibration Information – TSG 3411 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 3411 | 5Feb2023 | Factory |  |  |
| Conductivity | 3411 | 5Feb2023 | Factory |  |  |