
REVISION NOTICE TABLE

	DATE
	DESCRIPTION OF REVISION

	
	

	
	


PROCESSING NOTES
Cruise: 2024-005
Agency: IOS, Ocean Sciences Division, Sidney BC
Chief Scientist: Young K.  

Platform: Alta
Location: Barkley Sound

Project: Barkley Sound Euphausiid Monitoring Survey

Date: 8 February 2024 –9 November 2024
Processed by: Germaine Gatien

Date of Processing: 14 January 2026 – 3 February 2025
Number of original HEX files: 50 
Number of CTD files: 
49

Number of BOT files: 50
1 INSTRUMENT SUMMARY

SeaBird Model SBE-19 CTD (S/N  8241) was mounted with dissolved oxygen sensor #4354 and ECO fluorometer #7495.  

A Niskin bottle was mounted between 1 and 3m above the CTD to collect near-bottom samples; a separate Niskin was lowered from the vessel to collect surface samples.
SUMMARY OF QUALITY AND CONCERNS
This program involved repeat visits to 5 sites in February, March, April, May, June, July, August, October and November 2024. At 2 of the sites there was near-bottom salinity and nutrient sampling during a CTD cast as well as surface sample, and at some sites only a surface CHL sample was taken. 
The positions, station names and water depths were entered in the headers in a way that enabled easy entry into IOS Header format.
Bottle files were created by combining analysis results with CTD data selected from the approximate levels of the samples. There were some log entries indicating the distance between the CTD and Niskin bottle. Since this distance varied from 1m to 3m, and near-bottom gradients can be significant, it would be helpful to provide the information for all bottom sampling, or agree to a standard to be used for all casts. This will lead to better matches for salinity calibration. 

The CTD was lowered to ~10m and soaked for 1 to 2 minutes. The CTD was then brought to the surface and the full cast was run. The wait at the surface varied greatly from a few seconds to a minute. It is recommended that the wait be at least 30s to allow the wake effects from the rising CTD to dissipate. This wait is especially helpful in quiet areas where the boat is not moving away from the disturbance.

Calibration samples were available for comparison with CTD salinity and fluorescence. The CTD salinity at the bottom appears to be reading low by about 0.01psu and the post-cruise calibration from early 2025 confirms that it was reading significantly low.
The comparison of CTD fluorescence and extracted chlorophyll indicates that the fluorometer read too high when CHL was low, dropped to about the same as CHL at about 2ug/L and steadily fell to about 60% of CHL when CHL was >5ug/L. This is typical performance for this type of sensor.
While there was no dissolved oxygen calibration sampling, there was a post-cruise calibration. When differences between the CTD reading using the 2 calibrations was plotted against DO the usual sort of distribution was found and suggests that the CTD DO was reading low by about 3.5%. While some drift may have occurred after this cruise, the correction is typical of those found for this type of sensor, so all data were corrected using that result.

There were no CTD data acquired for event #69, but a bottle file was prepared with pad values entered for CTD variables. 
PROCESSING SUMMARY
1. Seasave 
This step was completed at sea.   
2. Preliminary Steps
All casts were converted using file 2024-005-ctd.xmlcon.

Cast #69 contained only a few data points. 

3. Conversion of Raw Data
The configuration file used at sea matched that used in the previous year. It was saved as 2024-005-ctd.xmlcon.

All files were converted using that file.
Test plots were made. All channels have normal profiles.

No spikes were noted except in fluorescence where they are expected.

The deployment scheme included a 10m soak for 1 to 2 minutes with a 20m soak on one occasion.
The CTD was then brought to the surface and the full cast was run. The wait at the surface varied greatly from a few seconds to a minute. It is recommended that the wait be at least 30s to allow the wake effects from the rising CTD to dissipate. This wait is especially helpful in quiet areas where the boat is not moving away from the disturbance.
Surface bottles were generally fired shortly after the beginning of the CTD cast. An exception is event #15 when it was taken while the CTD was at the bottom.
4. WILDEDIT

No spikes were noted in the data except at the beginning or end of the casts or in fluorescence, so WILDEDIT was not run. 
5. FILTER

The resolution of this instrument appears to be good and the pressure does not obviously need filtering. However, a test showed a slight improvement in salinity, so program FILTER was run using a low-pass filter, size 1s, on pressure and depth.
The temperature and conductivity were examined and the usual approach of applying a cosine filter size 8 in routine WFILTER did a good job of removing small reversals.
6. ALIGNCTD

Based on tests run for other cruises using similar equipment, ALIGNCTD was run on all casts to advance the DO channel by 2.5s. Plots were examined after this step and the results looked poor, with the dissolved oxygen clearly overcorrected. Similar findings for the same CTD from cruise 2023-064 and 2023-010 led to a choice of 0.5s for this step, which worked well.
ALIGNCTD was run using a 0.5s advance on all casts.  

7. CELLTM   INVESTIGATE 
CELLTM was run on all casts using the SeaBird recommended parameters, (α, 1/β) = (0.04, 8).
Tests were run to ensure this step worked well. It generally worked as expected, lowering salinity where temperature is decreasing (cell warms water raising conductivity but temperature sensor reading is correct, so salinity is overestimated). Conversely, it raises salinity where temperature is increasing (conductivity too low). 

8. DERIVE

Program DERIVE was run to calculate salinity and dissolved oxygen concentration (tau correction included). 
9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert the CNV files to IOS Headers. 
CLEAN was run to add event numbers (though most are present already) and replace pad values in the pressure channel with interpolated values. 
Time was checked in a few casts and all were in UTC.
10. Checking Headers

A cross reference list was checked and all expected information was present.
Header Check was run 

Surface Check was run and the average was -0.06db with a range from -0.19 to 0 to +0.05db. 
This is in good agreement with the 2023-064 and 2023-010 results.
The pressure sensor has a resolution of ~1db so these readings are excellent. 
No recalibration will be applied.
HEADER CHECK was run. There were some negative values in pressure, conductivity and fluorescence that are likely from surface spikes. This will be checked again after CLIP removes soak data.  
11. CLIP and CALIBRATE
The next step is to remove the data collected during soaks either at the surface or at 10m. 
CLIP was run to remove 600 records from each file. Plots were checked and the number had to be adjusted for some files to include more or fewer records to insure DELETE would choose only data from the full cast but not delete too much surface data. 
File clip.csv was prepared with the appropriate # of records to remove for each cast. 

12. SHIFT 
Conductivity  
Tests were run during processing of  2023-064 to fine-tune the shift to conductivity to best improve the stability in T-S space. A shift of -2.2 records was found best. That setting was tested on these data and also proved to make a significant improvement. That setting was applied to all casts.

Fluorescence

The fluorometer was not pumped, so a shift in alignment is expected to be small or unnecessary. Profile plots of temperature and fluorescence were examined and confirm that the alignment is ok. 

Dissolved Oxygen

This channel was aligned earlier, but checks were made by examining plots of temperature and dissolved oxygen. No further adjustment was made.

13. DELETE

DELETE was run on all casts using the following parameters: 

Surface Record Removal: Last Press Min.
Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 5 points) was deleted from 10db to 10db above the maximum pressure.

COMMENTS ON WARNINGS: There were no warnings.
Plots were made the results look good, but most casts need editing at the surface. 
14. DETAILED EDITING

The DEL files were submitted to CTD-QC and DELPRED files were returned.

All DEL files were copied to *.EDT so there will be a complete set of files even if some need no editing.

CTDEDIT was used to do some light editing of 35 files; this was limited to removal of records for which the pumps were off or had not had time to equilibrate, and records corrupted by shed wakes near the bottom as well as light editing of salinity for 1 cast. 
Notes of editing details were made in the headers. 

The output files were zipped and submitted to CTD-QC.

T-S plots were examined after this step and the results look good.
15. Initial Bottle Data Steps
There was no rosette in use during this program. There were surface bottles fired at all sites at about 1db and deep bottles at sites Sarita and Swales. 
Each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2024-005-bot-hdr.txt which will be updated as needed during processing. 
Spreadsheet 2024-005-bottles_plus_CTD_6linehdr. csv was created using entries from the event log relevant to Niskin sampling plus bottle data. Space was made for CTD data. 
Sample data from the analysts’ spreadsheets were entered.

It is useful to find CTD data to accompany the analysis results for general use, and to enable comparisons between bottles and CTD salinity. The first step is to obtain CTD data that comes from approximately the depth at which bottles were closed. For surface bottles this is assumed to be at 1m. 
For the bottles fired at the bottom we usually have a consistent vertical offset between the CTD and Niskin bottle. That was not the case for these data. The log entries indicate the height of the Niskin above the CTD varying from 1 to 3m. For many events there is no log entry to indicate the distance. These matches are always imperfect since the Niskin bottles are likely to contain water from above the firing level. For this cruise there is the added issue of not always knowing the depth at which the bottles were fired. Where bottom gradients are small this is not a big issue, but that is not the case in the project area.
An estimate was made for each cast to find the depth from which CTD data should be extracted. An initial attempt using edited and binned data provided highly variable results, especially for station Swale. Individual files at the last stage before DELETE were examined to pick out a depth that looks appropriate for the deep samples. The levels were based on log entries and/or expectations of particular CTD watch keepers. 
A THIN file was prepared to extract data from 1m, 2m and the various levels estimated near the bottom. 

For the first attempt the files thinned were those that had been through DELETE and CTDEDIT and in one run data that had also been bin averaged. Those data were exported to a spreadsheet and lines removed that are not appropriate for each cast. The data were sorted on decreasing pressure since that is the sample # order.
Those data were added to the 6line header file and a comparison done with salinity bottle data. 
The comparison between bottles fired near the bottom and CTD salinity was noisy with the results at Swale very different from those at Saria. This indicated that this way of finding appropriate CTD data to accompany samples was not satisfactory for the deeper bottles. It is appropriate for the surface bottles since they are believed to have been taken using a consistent method. And the lower vertical gradient at Sarita made results there more consistent than at Swale.
Since the first comparison run had large differences of random signs, a second attempt was made by extracting the CTD salinity from files before DELETE, CTDEDIT and BIN-AVERAGING were run. The distance of the bottle above the CTD did vary, so where that distance is clearly not 2m based on log notes or general habits of the watch keeper, that was taken into account.

This brought the comparison of the Swale casts into much better correspondence with the Sarita casts; The near-bottom gradients are larger at Swale.

Headers and CTD data were downloaded from the thinned CTD files with the exception of cast #69 for which no CTD data were acquired though samples were taken at the surface. The log entries were used for time/position information for that cast, pressure was estimated to be 1db and pad values were entered for the other CTD properties. 
The spreadsheet was converted to individual BOT files. 

CLEAN was used to add start time and positions. This also added Stop times and positions which just repeat the start entries, so those will be removed later.
16. Compare  

Data from the 6-line header file were exported to file 2024-005-compare.xlsx. Separate sheets were used to compare salinity and extracted chlorophyll samples with CTD salinity and fluorescence.

Salinity Comparison

There were 40 salinity samples with both CTD and bottle salinity available.
If the CTD sensor and bottle analysis were both perfectly accurate we would expect the CTD to possibly read a little higher salinity at the bottom due to the height difference and incomplete flushing of bottles. At the surface the bottles might read a little low if not completely flushed and the CTD could read a little high if it has carried deeper water up from the 10m soak. So again, we might expect that the CTD would appear to be reading higher when it is actually accurate.

There is tremendous variability in the near-surface differences with the CTD reading high by a median of 0.23psu and a standard deviation of 1.23psu. This may be due to the CTD being affected by wakes as it rose from the 10m soak.

As mentioned above, the first near-bottom comparison run had large differences and very different results for the two sites.
[image: image1.png]CTD Sal - Bottle Sal

0.06

0.05

0.04

0.03

0.02

0.01

-0.01

-0.02

2024-005 (Sal:CTD-Sal-BOT) vs Pressure

50

100

Pressure

150

“200
L]

250




A second attempt was made by extracting the CTD salinity from files before DELETE, CTDEDIT and BIN-AVERAGE were run. The distance of the bottle above the CTD did vary, so where that distance is clearly not 2m based on log notes or general habits of the watch keeper, that was taken into account.

This brought the shallower Swale casts into much better correspondence with the deeper Sarita casts. The near-bottom gradients are larger at Swale.
[image: image2.png]CTD Sal (pre-delete) - Bottle Sal

0.06

0.04

0.02

-0.02

-0.04

CTD-BOT using pre-edit & bin vs Pressure

[} 50

Pressure

150

“200
L]

250




COMPARE was rerun and it showed similar results for both sites, though there is still a wider spread at Swale than at Sarita. That is to be expected based on the higher vertical salinity gradient at Swale. 
The bottle contents are likely from slightly higher in the water column than the firing level due to incomplete flushing, so the bottle salinity is likely a little lower than ambient conditions. So we expect that the CTD will appear to be higher than the bottles. But the comparison shows the CTD reading lower by about 0.01psu. This is a large difference.
	Near-bottom bottles using binned data

	
	average
	stdev
	median

	All
	-0.0005
	0.0159
	-0.0072

	Swale
	0.0090
	0.0181
	0.0079

	Sarita
	-0.0099
	0.0025
	-0.0097

	Near-bottom bottles using pre-DELETE files

	
	average
	stdev
	median

	All
	-0.0116
	0.0060
	-0.0105

	Swale
	-0.0132
	0.0080
	-0.0131

	Sarita
	-0.0100
	0.0025
	-0.0101


The CTD appears to be reading lower by about 0.01psu at both stations when the pre-DELETE data are used. This may mean that the CTD is reading low or bottles are high. The latter seems unlikely as it is a result given the different times analyses were run and samples collected. No such problem was found during other cruises using the same salinometer. And there is no obvious temporal variability.

[image: image3.png]CTD Sal - Bottle Sal

-0.005

-0.01

-0.015

-0.02

-0.025

-0.03

CTD Salinity - BOTTLE Salinity

20g 40 60

Event#

80

100




It is likely that the CTD salinity really is low and the factory service record from early 2025 confirms that the salinity was reading low. One file was converted using the post-cruise calibrations and shows salinity lower by >0.01psu. Given that further drift may have occurred after this program ended, the comparisons just gives a rough guide. But it does support recalibration by adding 0.01psu to all CTD salinity data.

Fluorescence vs extracted chlorophyll

There were 49 CHL samples, all from the surface. 
The comparison of extracted chlorophyll with CTD fluorescence looks typical of these fluorometers. The fluorescence is mostly higher than CHL when CHL is ≤ 1ug/L and gradually drops to about 60% of CHL as CHL rises. This is typical of these sensors. The CHL samples are likely taken slightly higher than sampled by the CTD data.
[image: image4.png]2024-005 CTD ECO FL/CHL vs CHL

.
o.
.
L
S
O
ouoo
se
Y4
Q30

12

10

Extracted CHL




[image: image5.png]CTD ECOFluorescence

MW AU e N ®

2024-005 CTD ECO Fluor vs Extracted CHL

Y=0.6746x .

2 4 6 8 10

Extracted CHL




17. Other calibration checks
Sensor History – Previous cruises were 2023-010 and 2023-0064 with limited sampling. Salinity seemed close to samples in 2023-010. There was no calibration sampling during 2023-064. Sensors were recalibrated in Jan. 2025.
Historic Ranges – There was no local climatology available. A climatology that covers a large area and the full year was available but not very informative. All data fell into that except for some lower salinity at top and bottom and higher temperature near the bottom. 
Post-cruise calibrations – One of the deeper files was converted using the post-cruise calibrations from early 2025. 
· Plots show salinity lower by 0.012psu than when the 2025 calibration is used.
· Pressures differ by about 0.02db with the pre-cruise result being slightly higher. 

· There is a larger change in dissolved oxygen data. Data from 2db to the end of the downcast were compared using the pre-cruise and post-cruise calibrations; this suggests that DO is low by about 3.5%. That is a reasonable estimate for expected drift from these sensors, though some drift may have occurred during this cruise and after the project concluded.
[image: image6.png]Difference in CTD DO

CTD DO -CTD DO with post-cruise cal) vs CTD

y=-0.0351x

CTD DO usingpre-cruise calibration




18 CALIBRATE
Pressure does not require recalibration. File 2024-005-recal1.ccf was run to correct Salinity;T0:C0 by adding 0.01psu and multiplying Oxygen:Dissolved:SBE values by 1.035 in all EDT files.
19. Bin Average, Remove, Derive DO in mass units, Reorder
The files were bin averaged using 1db bins.
REMOVE was run to remove Scan_Number, Oxygen:Voltage, Descent Rate, Prediction Flag and Flag channels. 
Dissolved Oxygen was derived in mass units and that was used to calculate DO saturation. 
Plots of near-surface saturation show a range of 85% to 150%, with values close 100% in February, rising gradually to a maximum in August and then dropping, reaching values as low as 85% in November.
REORDER was used to get the 2 dissolved oxygen channels together.
20. HEADER EDIT and final checks of CTD files. 

Header Edit was used to fix headers, fix formats, correct headers and to add comments about processing. 
A cross-reference listing was produced.

A header check and standards check were run on the CTD files and no errors were found.

The sensor history was updated.

Plots of CTD casts were examined and no problems were found.

2 1. Final BOT file preparation

CLEAN was run to add flag 0 to blank flag channels and to create a start time and end time to the headers. The End Time will be removed later.
CALIBRATE was run using file 2024-005-recalsil.ccf to correct SILICATE where salinity is <25psu.

[Formula 120 : coefficients 0.904/0.003/25; source channel: Salinity*]
CALIBRATE was run using file 2024-005-recal1.ccf was run to correct Salinity;T0:C0 & Oxygen:Dissolved:SBE. (COR1)
Change Units was run to derive mass units for the Oxygen:Dissolved:SBE channel.
REORDER was run to get the 2  DO channels together.
REMOVE was run to remove the DATE, TIME:UTC. Latitude and Longitude channels.

SORT was run to get data in pressure order.
Header Edit was run to add comments and to remove END TIME (same as START TIME), TIME ZERO, Latitude2 and Longitude2.
Plots were made of all BOT casts. With just 1 or 2 levels these were not very useful but did show that the recalibration of salinity brought bottles and CTD closer together, especially near the bottom.
Finally all data from BOT files were extracted to a spreadsheet and compared to the Bottle Summary spreadsheet; no problems were found. 

A track plot was produced with station positions and no problems were found.

The standards check was run until all errors had been corrected.

A cross-reference list and header check were run on the BOT files. No errors were found.
Particulars
2. 20m soak.
7. Lost GPS feed at end of cast.

9. Wouldn’t load from rosette sheet.

19. Bottom sample ~2m above CTD.

28. Niskin ~1m above CTD.
32, 37, 41, 46. Bottom sample ~3m above CTD.

34. GPS keeps dropping feed. Used ship nav position

36. GPS didn’t work. No sounder. Chart says 200m. 

44. GPS dropped out.
45. Positions from vessel.

46, 50. Bottom sample ~2m above CTD.

47. Chlorophylls took forever to filter, reduced volumes
55, 56, 59. surface bottle snapped during CTD cast
64. Bottom bottles @188m

69. No data from CTD? Has 'low batt' note but batteries were fresh. Bottle file created with estimated pressure of 1db.
73. surface bottle snapped during CTD soak-forgot to log MR
82. Bottom bottle ~1-2m above CTD?

83. surface bottle Niskin leak from spigot
CRUISE SUMMARY

	Cruise ID#:    2024-005

	Dates:   Start: 9 February 2024                 End: 8 November 2024

	Location: Barkley Sound

	Chief Scientist: Young K.

	CTD#
	Make
	Model
	Serial#
	Used with Rosette?
	CTD Calibration Sheet Competed?

	1
	SEABIRD
	19+
	8241
	No
	Yes


CTD CALIBRATION INFORMATION

Make/Model/Serial#:
SEABIRD/SBE19+/8241
Cruise ID#:

2023-064


	Calibration Information

	Sensor
	Pre-Cruise
	Post Cruise

	Name
	S/N
	Date
	Location
	Date
	Location

	Temperature
	8241
	9Dec2022
	Factory
	
	

	Conductivity
	8241
	9Dec2022
	Factory
	
	

	ECO Fluorometer
	7495
	4Dec2022
	Factory
	
	

	SBE43 Oxygen
	4354
	20Dec2022
	Factory
	
	

	Press


	8241
	3Dec2022
	Factory
	
	


Station positions are approximate as they varied slightly among different legs of the program.
[image: image7.png]North Latitude

125,30

2024-005 Stn Names
152

125,10
I

125.00

8.9 ©

IS

&

o

i<
I

48,85

Conster

o

Swtle

@

v
§\\ Q Mockenz ie

48,90

- 4B.88

- 48.86

48,84

48.84
125.30

1B
‘West Longitude

125408




PAGE  
8

