## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2024-008				Agency: OSD

Location: North-East Pacific			Project: Line P

Chief Scientist: Robert M.			Platform: John P. Tully

Date: 14 – 30 Sept 2024

Processed by: Jeannette Bedard 		Date of Processing: Feb – July 2026

Number of original HEX files:    85		Number of processed CTD files: 85

Number of rosette casts: 44			Number of processed CHE files: 44

Number of original TSG txt files: 5		Number of processed TOB files: (1 per day) 14

# INSTRUMENT SUMMARY

CTD #0443 was mounted in a rosette and attached were 2 Wetlabs CSTAR transmissometer (1185DR & #1883DG), a SBE 43 DO sensor on the primary pump (#4372), SeaPoint Fluorometer on the secondary pump (#4186), a WetLabs fluorometer (#2216), a Biospherical QSP-400 PAR sensor (#70613) and a Valeport altimeter (#6341).

A thermosalinograph (SeaBird 45 S/N 0620) was mounted with a Wetlabs WETStar fluorometer (#1656), intake thermometer (SeaBird 38 #603) and flow meter; sampling interval was 5s.

Seasave version 7.26.7.121 was used for acquisition. The data logging computer was TULLY. The deck unit was a Seabird model 11+ #508. An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log Book was in excellent order with comments about problems encountered and a detailed list of equipment. The sampling logs were in good order. There were reports from the Chief Scientist outlining sampling methods and problems

The standard deployment procedure for this cruise as follows (based on Sampling notes):

For most casts, the rosette was brought to the surface.  Pumps were turned ON.  The rosette was brought down to 10m and kept there for 30 seconds.  Once back at the surface, the data started to be archived, with the rosette at the surface for 30 seconds longer.  Then the cast would start.

For the inlet work (Rivers, Smith, Knight) the rosette was brought and kept at the surface for a one minute soak, then the cast would start.  No “10-m dip”.  Note that for two casts (0116 and 0129) the archiving command did not work at the beginning of the cast so the water column was “stirred” before the archived cast was done.

On “PAR casts”, one file was started with the CTD in the water but the PAR sensor still out of the water. The CTD was lowered slowly to 10 m, and when passing 5 the pumps were turned on. Then 30 sec wait at 10, back to surface.  A new file was started at the surface then the regular cast was done. No Niskin bottles were closed on the short first cast.

During this cruise the wait time before closing the Niskin bottles was 60 seconds for pressures 400 (included) and shallower for casts that have oxygen sampling.  All casts without OXY, and pressures deeper than 400, had a 30 sec soak time.

Casts 125 to 0138 were done with the Niskin bottles closed, filled with water.

The sounder seemed to be roughly “30” too shallow on deep (>2000) casts.

There were 2 WetLabs CStar transmissometers in use during this cruise:

- Channel Transmissometer refers to sensor #1185DR (650nm - red)

- Channel Transmissometer:Green refers to sensor #1883DG (530nm - green)

For comparison with other Institute of Ocean Sciences cruises, note that the transmissometer wavelength is 650nm unless otherwise stated.

The SBE DO sensor has a fairly long response time so data accuracy is not as high when it is in motion as it is during stops for bottles. This will be especially true when vertical DO gradients are large. To get an estimate of the accuracy of the SBE DO data during downcasts (after recalibration) a rough comparison was made between downcast SBE DO and upcast titrated samples. Some of the difference will be due to problems with flushing of Niskin bottles and/or analysis errors and small mismatches in depth in the presence of large DO gradients, so the following statement likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data for this cruise are considered, very roughly, to be:

<0.05mL/L for pressure <400db

<0.03mL/L for pressure >400db

Despite the flow meter for the TSG stopping reporting flow for the last few days, the data appeared to be fine (suggesting the flow meter was faulty). There was good agreement between the different data types.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

- The Log Book and rosette log sheets were obtained.

- Nutrients, extracted chlorophyll, dissolved oxygen, salinity and DMS/DMSP data were obtained in QF spreadsheet format from the analysts.

- The cruise summary sheet was completed.

The CTD configuration was not changed during the cruise – 2024-008start.xmlcon was used throughout.

##### BOTTLE FILE PREPARATION

- ROS files were created using 2024-008-ctd*.xmlcon. The hysteresis correction and tau corrections were selected.

- Cast 33 was mislabelled—it should have been cast 32, change was made

- Note: a lot of the rosette sheets didn’t have bottle firing order filled in. There were no notes in the Daily Science Log that any were fired out of order, the log was assumed to be correct.

  - Cast 51 was fired out of order and there was no note in the log.

- Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create the header comments file which was updated during processing.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2024-008_OXY*.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and saved as OXY.csv.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2024-008_CHL QF*.xlsx. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared and saved as CHL.csv.

SALINITY

Salinity analysis was obtained in file QF2024-008_SAL.xlsx which included a precision study. The analyses were carried out in a temperature-controlled lab 11 to 25 days after collection. The files were simplified and saved as SAL.csv.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2024-006_NUTS*.xlsx. This includes a precision study. The file was simplified, saved as NUTS.csv.

DMS

DMS data were obtained in spreadsheet DMS Summary (2024-008).xls which includes duplicate analysis. Details on analysis are in file 2024-008 DMS report.doc. Only 2 figures are considered significant. Event #s were added to the file. All stations were run on an 8890 gas chromatograph.

- In Cast 51 there were three DMS values with the same sample number—these were averaged together and given a flag of 2

DMSP

DMSP-D and DMSP-T data were obtained in file 2024-008 DMSP QF file.xls. Details on analysis are in file 2024-008 DMS report.doc. The data were converted into DMSP files. Only 2 figures are considered significant.

- Bottle data was merged with chemistry data using BottleData_Processing.py.

- Cast 121 and 123 were removed as no samples were taken—total files is now 44

- Error file was examined.

CTD salinity was checked and there are values <25psu, so silicate correction was applied using 2024-008-recal-SIL.ccf

##### Compare and Corrections

Salinity

Compare was run with pressure as reference channel on the primary sensors. When a few outliers were excluded based on residuals the fit was:

CTD SAL Corrected = CTD SAL - 0.0003

For full details for the COMPARE run see file 2024-008-sal-comp1.xls.

Dissolved Oxygen

There was a significant amount of variability in near the surface. One the top 100m were excluded the fit was:

CTD DO Corrected = CTD DO - 0.0191

Upon examination a second DO calibration was needed of:

CTD DO Corrected = CTD DO - 0.0282

This still didn’t result in a great comparison, however it was an improvement. It was determined best to leave the corrections here.

For full details for the COMPARE run see file 2024-008-oxy-comp1.xls and 2024-008-oxy-comp2.xls

Calibration files used: 2024-008-recal-CTD.ccf and 2024-008-recal-CTD2.ccf

##### Final Bottle File Steps

- SORT was run to arrange casts in pressure order.

- For all casts REMOVE was run to remove the following channels:

  - Temperature:Secondary, Conductivity:Secondary, Salinity:T1:C1, Oxygen:Dissolved:SBE (mass units)

  - Removed PAR from cast 53 (cast where PAR sensor removed).

- A second SBE DO channel with mass units was added for both the CTD DO and titrated DO and REORDER was run to get the pairs of DO channels together.

- EDIT HEADERS was run to fix formats and channel names and to add comments about analyses and CTD processing.

- Standards check was run. No problems were found.

- Data were exported from the CHE files to file 2024-008-bottles-final.xlsx. The entries were compared with the rosette log sheets and no problems were found.

- The track plot looks fine.

- Plots of each file were examined and no problems were found.

- A cross-reference listing and header check were produced for the CHE files.

## CTD processing

##### Conversion of Full Files from Raw Data

- Configuration did not change during the cruise. 2024-008start.xmlcon was used throughout.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, depth, conductivity & temperature only in the full cast files (*.CNV).

- Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

- The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

ALIGNCTD was run on all casts to advance the oxygen voltage by +2.5s, a setting which has worked well in the past for this type of sensor.

##### CELLTM

The noise in the upcast data makes tests for the best parameters for this routine very difficult to interpret. In the past when upcast data were not so noisy, the default setting of (α = 0.0245, β=9.5) was generally found to be the best choice—these values were used here.

##### DERIVE

Program DERIVE was run on all casts to calculate primary and secondary salinity, dissolved oxygen concentration (ml/l) and descent rate.

IOS Shell

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert the final SeaBird CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number. The ECO Fluorometer slightly below zero values were replace with PAD values.

##### Checking Headers

- The cross-reference check and header check were run. No problems were noted.

- Surface check was run however it didn’t work

- Cruise tracks were plotted and added to the end of this report.

The following header changes were made:

- Cast 0003 – Haro59: the station name is written on the Tully: 2024-008 line. The Depth is written as “deck”.  Should be 240 m.

- Cast 0014 – P4: has wrong depth in header, should be 1324 not 811.

- Cast 0056 – P26: has the wrong depth in header, should be 4211, not 4321.

- Cast 0063 – P21: has the wrong depth in header, should be ~3950, not 3873.

- Cast 0067 – P17: has the wrong depth in header, should be ~3610, not 2050.

- Cast 0069 – P16: has the wrong depth in header, should be ~3630, not 1971.

- Cast 0071 – P16: has the wrong depth in header, should be ~3630, not 1971.

- Cast 0073 – P16: has the wrong depth in header, should be ~3630, not 3600.

- Cast 0145 – Kn09: station name written as Kn08, should be Kn09.

The altimeter and water depth readings from the headers were exported to a spreadsheet. A check value was calculated by subtracting water depth from maximum depth sampled plus altimetry header.

Note: it was noted that the sounder seemed to be ~30db too shallow compared to historical depths.

The following changes were made:

- Cast 0053 - listed a max sample depth of 4239m and an altimeter value of 9.9m, but the water depth was 4211m. Since there was a listed concern about the depth sounder accuracy of around 30 db, values were left as is.

- Cast 0060 – listed a max sample depth of 11m and an altimeter value of 0.8m, but the water depth was 4021m. Altimeter value was deemed nonsense and removed from header.

- Cast 0053 - listed a max sample depth of 1409m and an altimeter value of 9.16m, but the water depth was 1352m. In the daily science log, the water depth when the rosette was at the bottom was 1395m, since there was a listed concern about the depth sounder accuracy of around 30 db, using the deeper value makes sense. Cast file was changed.

Note:

- Primary and secondary temperature and salinity were closely aligned. For further processing the primary sensors were chosen.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel for all casts using the usual advance of +24 records. Plots show that the fluorescence offset is reasonably close to the temperature offset after this step.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it did. No further alignment was needed for the DO concentration channel.

Conductivity

- T-S plots were studied looking for casts that would be good candidates to test conductivity shifts—no good candidates were found. No conductivity shift was applied.

##### DELETE

DELETE was run on the shifted files. The following DELETE parameters were used:

- Surface Record Removal: Last Press Min

- Maximum Surface Pressure (relative): 10.00

- Surface Pressure Tolerance: 1.0                  Pressure filtered over 9 points

- Swells deleted. Warning message if pressure difference of 2.00

- Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

- Drop rate applies in the range:  10db to 10db less than the maximum pressure

- Sample interval = 042 seconds. (taken from header)

##### Other Comparisons

Experience with these sensors since last factory service –

- For CTD2 (0443) --This CTD was last used for 2024-020, and it functioned fine.

##### DETAILED EDITING

- The *.DEL files were zipped in and submitted to the CTD-QC-File Processor—however the software didn’t work. Continued on without this input.

- Most casts had shed wakes and spikes which were removed along with all data captured when the descent rate was below 0.2 m/s. Some of the deep casts had quite large shed wakes.

- After editing, T-S plots were examined for all casts and no further editing was found necessary for temperature and salinity; a few unstable features remain but may be real.

##### Corrections to Pressure, Salinity and Dissolved Oxygen Concentration

File 2024-008-recal-CTD.ccf was prepared apply corrections as follows:

- Salinity:T0:C0 was recalculated as: CTD Salinity Corrected = CTD Salinity - 0.0003

- dissolved oxygen was recalculated as: CTD DO Corrected = CTD DO - 0.0191

COMPARE was rerun after recalibration. It was determined a second oxygen correction was required.

File 2024-008-recal-CTD2.ccf was prepared apply corrections as follows:

- dissolved oxygen was recalculated as: CTD DO Corrected = CTD DO - 0.0282

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

- Bin channel = pressure;	Averaging interval = 1.000;  Minimum bin value =   .000

- Average value will be used. Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined and no significant differences were found from climatology.

##### Final CTD File Steps (REMOVE and HEADEDIT)

For all casts REMOVE was run to remove the following channels: Scan_Number, Temperature:Secondary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T1:C1, and Flag.

- PAR from casts 33 and 53

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add comments about processing.

The Standards Check routine and Header Check were run; no problems were found.

Profile and T-S plots were examined and look ok.

##### Thermosalinograph Data

An IOS TSG45 was used for this cruise and separate flow to a fluorometer in the lab.

- TSG data came in five text files and contained extra fields with data from other instruments. These files were moved to excel and simplified to only include the relevant TSG data.

- Two-line Headers were added with variable names and units. Data were rearranged to separate date and time, convert fluorescence in volts to ug/L (using the Wetlabs chlorophyll characterization sheet), add a pressure column with all values set to 4. All NAN values were replaced with -99. A “File Name” column was then added with format “2024-002-2024MMDD”.  This will be used as a file break column to create 1 file per day. These files were saved in CSV format.

- The .csv was converted into IOSShell format with a file per day.

- CLEAN was run to reset the number of records, min and max values, set the start and end times, and latitude and longitude limits.

- ADD TIME CHANNEL was used to add Julian dates (Offset from Time Zero – i.e. Day of Year). A record number was also added to enable averaging (for use in comparison to CTD files).

  - Time zero was set to 31 December 2023 0:00:00. (Note that this step leads to problems plotting until REORDER is run.)

- DERIVED QUANTITIES was run to derive salinity using the lab temperature. (Sigma-T was derived later.)

- REORDER was run to move the Julian date to after the Time/Date channels and to put salinity and fluorescence after the lab temperature. Also the record # was moved to the end. Pressure was dropped.

a.) Plots

- A track plot was produced and looked fine; it was added to the end of this report.

- Time-series plots were produced and there are some salinity spikes that can mostly be removed using a graphical editor.

- A plot of all differences (Lab Temp – Intake Temp) through the whole record shows variability, but in quieter sections heating in the loop is in the 0.4 to 0.5C° range which looks reasonable.

- The main flow rate was zero for most of the cruise, however reasonable flow rates were recorded (~1.3 m/s) for the fluorescence sensor. However TSG data appeared good throughout and matched well with the CTD data.

- There were only a few good fluorescence values from the TSG, not enough for comparisons with the CTD.

b.) Comparison of TSG data to CTD data

- The CTD files were thinned to reduce the files to a single point from the downcast at or within 0.5db of 4.5db. These were exported to a spreadsheet.

- The TSG files were averaged over 5 records (~30 seconds). Records were picked out for the times that match the CTD casts and added to the spreadsheet 2024-008-CTDvsTSG.xlsx. Columns were added to calculate the differences between TSG and CTD latitude and longitude. This is used to ensure the matches are good.

- A comparison was made of positions for the CTD and TSG data to check for good matches. The differences in positions are expected to be small despite the averaging because the ship was stopped at these times.

The initial comparison between TSG and CTD data using all casts was:

|  | Tintake-Tctd | Tlab-Tctd | TSG Sal-CTD Sal | TSGfl/TSGctd |
| --- | --- | --- | --- | --- |
| average | -0.20027 | 0.253301 | 0.180215 | n/a |
| stdev | 0.591178 | 0.279603 | 2.64362 | n/a |
| median | 0.00125 | 0.1855 | -0.306 | n/a |
| min | -2.5284 | -0.4708 | -9.6147 | n/a |
| max | 0.6825 | 1.745 | 10.7795 | n/a |

The median and average differences are reasonably close together. There were a few spikes in the data which will be removed manually.

Heating in the Loop

The average difference between the intake temperature and lab temperature was 0.4 deg C with a standard deviation of 0.6 deg C.

Comparisons of Loop samples and TSG data

There were 24 overlapping loop salinity samples.

The TSG salinity is lower than loop samples by a median of 0.3098 PSU with a minimum of 0.2898 PSU and a maximum of 0.3776 PSU.

f.) Editing

TSG data were edited using CTDEDIT.

- Isolated spikes were removed throughout.

- In file 20240915, the flow rate dipped to zero for a short period—other data was removed for this time

- In file 20240925, a few hours in the flow rate dropped to zero and remained here for the rest of the cruise. Data appears to be fine and it was assumed the flow metre failed.

g.) CALIBRATE, REMOVE and CLEAN

- Based on the small salinity difference between the bottle and TSG CALIBRATE was deemed unnecessary.

- Sigma-T was derived earlier. First, Pressure needed to be added to enable the calculation. ADD CHANNEL was run to add pressure = 4.5db. Then DERIVE was run to derive sigma-T.

- REMOVE was run to remove channels Temperature:Difference, Record # and pressure. The flow rate channels were left in the files.

h) Preparing Final Files

- HEADER EDIT was used to change the DATA DESCRIPTION to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats and to add comments.

- The TSG sensor history was updated.

- As a final check plots were made of the cruise track and time-series and all look fine.

- A cross-reference list was prepared.

##### Loop File

- Derived Quantities was run to derive Sigma-T and add it to the CHE files. These files were clipped to 7bd then exported to a spreadsheet.

- Loop samples were added into the spreadsheet. There were no DMS or DMSP samples corresponding to the loops, these parameters were dropped.

- The sampling method column indicates ROS or USW for loops.

- 6-line headers were added and the usual headers removed.

- The columns were rearranged into the usual order of the 6linehdr files.

- The file was saved as 2024-008-surface-6linehdr.csv

- The file break column was filled with value 1 so all data to ensure only a single file is created in conversion.

- CONVERT was run to produce an IOS Header file.

- CLEAN was run to get start and stop times and to add flag 0 to empty flag cells.

- Edit Headers was used to add header comments about the cruise, CTD data and sample analysis.

- The track plot looks reasonable.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |

| Calibration Information - 0443 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2106 | 16Feb2023 | Factory |  |  |
| Conductivity | 2754 | 18Jan23 | Factory |  |  |
| Secondary Temp. | 5130 | 23Feb2023 | Factory |  |  |
| Secondary Cond. | 1766 | 18Jan2023 | Factory |  |  |
| Transmissometer | 1185DR | 11Jan2023 | Factory |  |  |
| Transmissometer | 1883DG | 11Jan2023 | Factory |  |  |
| SBE 43 DO sensor | 4372 | 25Jan2023 | Factory |  |  |
| PAR sensor | 70613 | 24Feb2021 | Factory |  |  |
| SeaPoint Fluor. | 4186 | May 2023 | Factory |  |  |
| WetLabs Fluor. | 2216 | 27Nov2018 | Factory |  |  |
| Pressure Sensor | 0443 | 22Feb2023 | Factory |  |  |
| Valeport Altimeter | 6341 | 23Nov2021 | Factory |  |  |

# TSG Make/Model/Serial#:	SEABIRD/45/0620

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature-SBE45 | 45-0620 | 12Jan22 | Factory |  |  |
| Conductivity-SBE45 | 45-0789 | 2Feb24 | Factory |  |  |
| Wetlabs WETStar Fluor. For depths deeper than, and including, 125 dbar, we would wait 30 seconds before closing a bottle. For depths shallower than, and including, 100 dbar, we would wait 60 seconds before closing a bottle. | 1656 | 12Mar2021 | Factory |  |  |
| Temperature-SBE38 | 603 | 4Jan 2024 | Factory |  |  |