## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2024-009

Agency: OSD

Location: Strait of Georgia

Project: NRCAN/ONC Maintenance

Chief Scientist: Jackson J.

Platform: John P. Tully

Date: 10 July 2024 – 17 July 2024

Processed by: Germaine Gatien

Date of Processing:  12 December 2024 –  12 December 2024

Number of HEX files: 3		Number of CTD files processed: 3

# INSTRUMENT SUMMARY

A SeaBird 19+ CTD (#5132) was used with an SBE 43 DO sensor (#4239 and an altimeter attached.

Seasave version 7.26.7.121 was used for acquisition.

# SUMMARY OF QUALITY AND CONCERNS

These data were collected opportunistically during an NRCAN/ONC cruise. There was no calibration sampling and we have no history of use for this CTD.

Pressure calibration looks excellent.

There is no record of the deployment strategy but the downcast data suggest there was no 10m soak before the full casts were run. There are suspicious sub-surface maxima at about 5m in dissolved oxygen concentration that may be due to bubble release as pressure increased. This is most notable at station 41.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

- There were scans information about the 3 casts.

- The cruise summary sheet was completed.

- There was no calibration sampling for the salinity or dissolved oxygen sensors.

- These data were collected opportunistically, and no calibration history was available.

The configuration file was the same for all three files; a copy was saved as 2024-009-ctd..xmlcon.

##### Conversion of Full Files from Raw Data

All files were converted using file 2024-009-ctd.xmlcon.

The Tau function was selected but not the hysteresis function since there was no deep sampling.

Plots were examined and all expected channels are present.

##### WILDEDIT

Not run as no noticeable spiking.

##### FILTER

The resolution of this instrument appears to be good, so the pressure does not obviously need filtering. However, tests on previous cruises showed some improvement in salinity and dissolved oxygen if a low-pass filter, size 1s, was run on pressure, so that was done for all casts.

The temperature and conductivity were examined and the usual approach of applying a cosine filter size 8 in routine WFILTER did a good job of removing small reversals.

This step had little effect on Dissolved oxygen data.

##### ALIGN DO

ALIGNCTD was run on all casts using +2.5s. That setting proved to be too high, so it was run again using +2s and +1.5s.

##### CELLTM

CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate salinity, dissolved oxygen concentration and descent rate of the CTD.

Plots were checked to see how the various steps worked.

- The temperature and conductivity channels are smoother.

- The DO alignment looks good.

- Salinity looks much smoother and T-S plots show most unstable features were removed.

- CELLTM worked as expected, lowering salinity where temperature is decreasing (cell warms water raising conductivity but temperature sensor reading is correct, so salinity is overestimated). Conversely, it raises salinity where temperature is increasing (conductivity too low). This brings downcast and upcast traces closer together, so T-S downcast and upcast plots are in better agreement after this step. Downcast and Upcast T-S curves also moved closer together where temperature is decreasing with depth.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

- The cross-reference check was run and details were checked against scans of log sheets for each cast. Times are close to the log entries except that logs were in local time and headers in UTC. Positions are close to those reported from other cruises at those sites. The station names were simplified by dropping the initial “SoG_CTD” leaving just the usual names for those sites, stations 41, 42, 43.

- The interval within the files is the expected 0.25s but in the header it is 0.5s. No explanation could be found for that, so the header was changed to 0.25s. The filters were applied earlier with the assumption that the interval was 0.25 and appear to have worked well, and nowhere in the SeaBird headers is there any indication that 0.5s is correct.

- The header check indicates that there are some negative pressures, but the only incidence was a single spike in the first cast during the soak period.  Conductivity profiles show values dropping precipitously when pressure was about

- Judging by conductivity values, the CTD was in water at 0.3db of downcast, but that may just be when the pumps came on. The upcast shows the CTD was at or passing through the surface at about 0.1db. These values indicate that the pressure sensor was working within specifications.

- Cruise tracks were plotted and no problems were found and added to the end of this report.

##### Shift

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier.

Conductivity

Tests were run to assess what setting is best to align conductivity with temperature (as judged by the effect on salinity as seen in T-S space). The best setting was -1.7 records.

SHIFT was run applying a shift of -1.7 records to conductivity. Salinity was recalculated.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure was not filtered

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 5 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

The DEL files were copied to *.edt so there will be a complete set even if no editing is required for any casts.

##### Other Comparisons

Experience with these sensors since last factory service –  None

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. All data fell within the climatology.

Post-Cruise Calibration -None available.

##### DETAILED EDITING

CTDEDIT was used to edit all casts by removing data form the top of casts where the pump was obviously not operating and to remove data from the bottom of casts that were corrupted by shed wake. Light editing of salinity was applied to cast #2 only, near the bottom.

After editing T-S plots were examined for all casts and no further editing was found necessary.

##### Corrections to Pressure, Salinity and Dissolved Oxygen Concentration

There is no indication of a problem with pressure calibration and there is no information available upon which to base recalibration of salinity or dissolved oxygen. No recalibration was applied.

##### Fluorescence Processing

No fluorometer present.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined. No problems were found.

Profile plots were examined. No problems were found

##### Final CTD File Steps (REMOVE and HEADEDIT)

For all casts REMOVE was run to remove the following channels:

Scan_Number, Oxygen:Voltage:SBE, Descent_Rate and Flag.

Dissolved Oxygen was derived in mass units and that was used to calculate DO saturation.

Values ranged from 80% to 140%, increasing from east to west.

There is a peak in DO at station 41 at 5m. This looks rather like it could be due to bubble release as pressure increased. Without a 10db soak, there can be a problem with flow and bubble clearance. The conductivity and salinity do look ok, though there may be some effect on those as well.

There was a long soak at the surface for that cast, but that may not have been enough.

The upcast does not show a subsurface peak like the downcast.

The cast at station 42 also has a suspicious peak but it is smaller.

At station 43 there is no obvious peak and the upcast looks similar to the downcast.

So there is doubt about the oxygen values above 10db, but the evidence is not clear enough to remove the data.

REORDER was used to get the 2 dissolved oxygen channels together.

HEADER EDIT was used to fix formats and channel names and to add comments about processing.

The Standards Check routine was run and no problems were found.

The Header Check was run; no problems were found.

Profile and T-S plots were examined and no problems were found other than the surface DO issue noted above.

CRUISE SUMMARY – CTD

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | SeaBird | 19+ | 5132 | No | Yes |  |  |
| Calibration Information - 585 |  |  |  |  |  |  |  |
| Sensor |  |  | Pre-Cruise |  | Post Cruise |  |  |
| Name |  | S/N | Date | Location | Date | Location |  |
| Temperature |  | 5132 | 07-Feb-23 | Factory |  |  |  |
| Conductivity |  | 5132 | 07-Feb-23 | Factory |  |  |  |
| SBE 43 DO sensor |  | 4239 | 05-Jan-22 | Factory |  |  |  |
| Altimeter |  | ? | ? | Factory |  |  |  |
| Pressure Sensor |  | 5132 | 23-Jan-23 | Factory |  |  |  |