
REVISION NOTICE TABLE

	DATE
	DESCRIPTION OF REVISION

	
	

	
	


PROCESSING NOTES
Cruise: 2024-010




Agency: Ocean Sciences Division (OSD)
Location: Juan de Fuca Strait & Strait of Georgia
Project:  Salish Sea Biophysical Survey
Party Chief: Bluteau C.
Platform: Vector
Date: 2 October – 8 October 2024 
Processed by: Germaine Gatien
Date of Processing: 12 November 2025 – 4 December 2025
Number of original HEX files: 79
Number of CTD files:  79

Number of CTD casts processed:  79
Number of bottle files: 27

Number of bottle casts processed: 27
TSG files: 1



Number of TSG files processed: 1
INSTRUMENT SUMMARY
SBE911+ CTD 1453 was mounted in a rosette and attached were Wetlabs CSTAR transmissometer 1201DR, SBE 43 DO sensor 1438, SeaPoint fluorometer 3640, pH sensor 0691, PAR sensor 4565, SPAR sensor 20518 and altimeter 73171.  

Seasave version 7.26.7.121 was used for acquisition. 

The salinometer used at IOS was a Guildline model 8400B Salinometer Serial Number 73274.

An IOS rosette with 24 10L bottles was used. 

A thermosalinograph (SeaBird 45 S/N 3411) was used. 

SUMMARY OF QUALITY AND CONCERNS
The Science Log was in digital format. The header page had complete information about participants and equipment. There were good notes in the comments sections of both the CTD and sampling logs about problems encountered including a serious problem in the TSG conductivity. There were 3 casts during which the buffer was left on the pH sensor that were not noted in the log.
The comparison of CTD salinity with bottles during this cruise and 2024-035 during which the same CTD was used, was based on few samples and produced inconsistent results, but neither suggest a large error. It is likely the salinity is reading a little high but within ±0.003psu. This cruise had good dissolved oxygen sampling.
All temperature data fell within the local climatology. The casts around Texada Island and one at station 56 had high salinity around 30db to the bottom. The deep water from stations 44 to 48 in the southern Gulf of Georgia had high salinity in the deeper part of the casts, starting at between 120 and 180db and continuing to the bottom. All salinity in Juan de Fuca Strait and north of Texada Island were within the climatology. These variations are not considered evidence of problems in CTD data, but rather an illustration of the high variability in this region and the limitations of the climatology. 
There were many problems with spikes in the fluorometer signal during the first part of the cruise; most occurred during upcasts, but not when stopped for bottles. The only spikes noted in downcast data occurred during cast #35; those bad data were padded. The fluorometer was replaced after event #36. 

The SBE DO sensor has a fairly long response time so data accuracy is not as high when it is in motion as it is during stops for bottles. This will be especially true when vertical DO gradients are large. To get an estimate of the accuracy of the SBE DO data during downcasts (after recalibration) a rough comparison was made between downcast SBE DO and upcast titrated samples. Some of the difference will be due to problems with flushing of Niskin bottles and/or analysis and sampling errors, so the following statement likely underestimates SBE DO accuracy. 
Downcast (CTD files) Oxygen:Dissolved:SBE data for this cruise are considered, very roughly, to be:

      ±0.15 mL/L from 0-50db

      ±0.10 mL/L from 50db to 100db 
      ±0.05 mL/L from 100db to 200db 
      ±0.02 mL/L below 200db.

PROCESSING SUMMARY

1. Seasave
This step was completed at sea; the raw data files have extension HEX.

2. Preliminary Steps

· The digital daily log and rosette log sheets were obtained and checked for comments. 
· Salinity, Nutrients, Extracted CHL and Dissolved Oxygen data were obtained in QF spreadsheets. 
· The history of the pressure sensor, conductivity and dissolved oxygen sensors were checked. Most sensors were recalibrated in early 2024 and have only been used since then during 2024-035.
· The calibration control files were obtained and were the same as for 2024-035. The only change during the cruise was the replacement of the fluorometer after event #36. 
3. BOTTLE FILE PREPARATION 
The ROS files were created using 2024-010-ctd.xmlcon. The hysteresis correction was not selected since there were no very deep casts; the Tau correction was selected.

The ROS files were converted to IOS format

The files were put through CLEAN to create BOT files. 

Temperature and salinity were plotted for all BOT files to check for outliers, none were found. 

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. Sample numbers were added based on the rosette logs.
The ADDSAMP file was used to add sample numbers to the BOT files – output *.SAM.
The SAM files were bin-averaged on bottle # and called SAMAVG.  

A new method was used to create MRGCLN1 files – python program “BottleData_Processing.py”.
To run this it is necessary to provide CSV files for each analysis type and a metadata form, and the raw HEX, BL and XMLCON files.

First, the analysis data were used to create individual CSV files. The information from the top of each of the QF files was extracted and placed in file 2024-010-bot-hdr.txt.

EXTRACTED CHLOROPHYLL 

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2024-010_CHL QF*.xlsx. The file included comments, flags and a precision study. A simplified version of the spreadsheet was prepared and saved as 2024-010chl.csv. 

The csv file was then converted to individual CHL files. 
SALINITY  
Salinity analysis was obtained in file QF2024-010_SAL*.xlsx. The analyses were carried out in a temperature-controlled lab within 2-8 days of collection. The files were simplified and saved as 2024-010SAL.csv. 
DISSOLVED OXGYEN  

Dissolved oxygen data were provided in spreadsheet QF2024-010_OXY_11Mar2024.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and saved as 2024-010OXY.csv. 
NUTRIENTS 

The nutrient data were obtained in spreadsheet QF2023_010_NUTS.xlsx which includes flags, comments and a precision study. The spreadsheet page with the final data was simplified and saved as 2024-010NUTS.csv. 
At this point python program BottleData_Processing.py was used to prepare bottle files.
Input for the program are:
HEX, BL, XMLCON files 

CSV files for each analysis type available
PSA file – create in Seasoft as if doing the conversion but don’t run

2024-010_METADATA.csv – form updated for each cruise.
The output files are named CHE. (They could be called CHEPRELIM but are not really CHE files.)

*They were put through Edit Header to fix the bottle number channel – output MRGEDIT.

*They were then put through CLEAN to fix the channel limits in the header – output MRGCLN1

Those files were then put through SORT to sort on sample # with output MRCLN1s; this is not always necessary but if done as standard step avoids errors.

*These steps could become unnecessary if there is a fix to the python routine.

Next, the SAM files were prepared.

The hex files were converted to ROS files.

Those were converted to IOS Headers (*.ios).
They were put through CLEAN to add event #s and were saved as *.BOT

Temperature and salinity were plotted for all BOT files to check for outliers. File 84 had noisy primary salinity at 150m; it was cleaned lightly and saved as *.ED1. ED1 was copied to *.BOT.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. Sample numbers were added based on the rosette logs.

The ADDSAMP file was then sorted on event number & then sample number.

It was used to add sample numbers to the BOT files – output *.SAM.
The SAM files were put through Edit Headers to change:

Bottle_Number ==> Bottle:Firing_Sequence
Bottle:Position ==> Bottle_Number
The order of doing those changes matters. The output files were named SAMEDIT.

The SAMEDIT files were bin-averaged on bottle # and called SAMAVG 
The SAMAVG files were put through Header Edit to change channel name Bottle_Number to Bottle and Bottle:Position to Bottle:Firing_Sequence. Output was called SAMAVGhdr. 

Files SAMAVG were merged with MRGCLN1s to create MRG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. Some problems were found and the program was adjusted and rerun until issues were all resolved.

A header check was run and no problems were found.
4. Compare  

Salinity  
There were 9 salinity samples. 2 bottles were removed from the comparison because the standard deviation in the CTD salinity was >0.001psu. The primary salinity was high by an average of 0.0048psu (std dev 0.0007psu). The secondary salinity was low by 0.00002psu (std dev 0.0016). 
Using just the 2 samples from near the mouth of Juan de Fuca where flushing of Niskins is likely to be better, the primary was high by 0.0052psu and the secondary by 0.0017psu. 

For the primary the results are quite consistent, so flushing may not be a big issue.

For the secondary the results are different, but it is noted that in the full comparison the standard deviation was quite high.

The variability with pressure and time is higher in the secondary than in the primary, but this observation is based on too few data to be significant.

For more detail see file 2024-010-sal-comp1.xls.
Dissolved Oxygen

COMPARE was run between measured and bottle dissolved oxygen. 
A few large outliers were excluded based on large differences:

Event 1 – Sample #11 – 5m very high vertical gradient. Values likely both ok but slightly displaced vertically. 
Event 11 – Sample #49 – 10m CTD data noisy; shed wakes.
Event 22 - Samples #155 – 10m CTD high variability during stop.
No flags are appropriate.
Other outliers were identified based on residuals. The offset for the trendline was set to match the average of 2 near-zero titrated samples from Saanich Inlet. The result is:
CTD DO Corrected = CTD DO *1.0324 + 0.007
Fits using the same data show no indication of time variability. 

The fit against pressure indicates that there is pressure-dependence above 40m. To check if this is a significant issue, all data above 40m were excluded and the fit was very similar with a slope of 1.0317.

Plots were made of Oxygen:Dissolved and Oxygen:Dissolved:CTD vs Salinity:T0:C0 did not indicate that further quality flags should be attached to samples.

Fluorescence

COMPARE was run with pressure as the reference channel.
There was a malfunction in the fluorescence signal during the cruise, so the fluorometer was switched after cast #36. (Fluorometer #3640 also malfunctioned a few times during cruise 2024-035.)

Uranine tests in February 2025 found no problem with either fluorometer. 

The dark values were found by examining plots below 300db.

Dark Values: 
SeaPoint 3640   +0.011ug/L

SeaPoint 3641   +0.012-0.017ug/L except in areas of deep mixing where higher values were found, as expected.

Comparisons were done separately for the two fluorometers.

Event #36 stands out as very bad, so this was skipped in the chart below:

[image: image1.png]Fluorescence

7.00

6.00

5.00

4.00

3.00

2.00

1.00

0.00

2024-010 Fluor #3640 vs CHL (Excluding cast #36)

Extracted CHL




[image: image2.png]Fluorescence

5.00
4.50
4.00
3.50
3.00
2.50
2.00
1.50
1.00
0.50
0.00

2024-010 Fluor #3641 vs CHL (events 39-93)

o a2

0.00

o®

. 0
oo.... 1
$

1.00 2.00 3.00 4.00 5.00 6.00
Extracted CHL

7.00




[image: image3.png]Fluorometer/CHL

2024-010 Fluor #3640 FL/CHL vs CHL (Events 1-34)

1.60

1.40
o0 0

B X
100 ‘oe
0.80
0.60
0.40
0.20

0.00

0.00 5.00 10.00 15.00
Extracted CHL

20.00

25.00




[image: image4.png]Fluor/CHL

2.00
1.80
1.60
1.40
1.20
1.00
0.80
0.60
0.40
0.20
0.00

0.00

2024-010 - Fluor #3641 FL/CHL vs CHL

“’:og'.-. X

1.00 2.00 3.00 4.00 5.00 6.00
Extracted CHL

7.00




There are some unusual fits for #3641 where fluorescence/CHL is higher than usually seen, but examination of CTD plots suggests these from an area with high fluorescence gradients. We expect samples to come from a little lower in the water column than the CTD data, so the high ratio is likely reflective of incomplete flushing of Niskin bottles. 

Most of the data reflects the usual observation that fluorescence reads high at low CHL and low at high CHL. Most of the fluorescence data from both sensors are lower relative to CHL than usually seen but this does vary a lot, possibly related to variable environmental conditions.
Overall the fluorometers appear to both have performed reasonably well while stopped.

For more details see document 2024-010-fl-chl-comp1.xls.

5. Conversion of Full CTD Files from Raw Data

All files were converted using files 2024-010-0001.xmlcon. The hysteresis correction was not selected since there were no deep casts; the Tau correction was selected. A few casts were examined and all expected channels are present.  
6. WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the 
full cast files (*.CNV).  

Parameters used were: 
Pass 1    Std Dev = 2 
Pass 2    Std Dev = 5 
Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

7. ALIGN DO

ALIGNCTD was run on all casts using +2.5s.

8. CELLTM

CELLTM was run using the default settings (α = 0.0245, β=9.5) for both the primary and secondary conductivity. 
9. DERIVE 

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.
At this point tests were done to ensure steps 6, 7, 8 and 9 worked appropriately.

· A few files were found with large and small spikes in conductivity profiles before this step; they were all removed by WILDEDIT. 

· The alignment of dissolved oxygen with temperature looks good after the ALIGNCTD step.  

· CELLTM worked as expected, lowering salinity where temperature is decreasing (cell warms water raising conductivity but temperature sensor reading is correct, so salinity is overestimated). Conversely, it raises salinity where temperature is increasing (conductivity too low). This brings downcast and upcast traces closer together, so T-S downcast and upcast plots are in better agreement after this step.
DERIVE was run a second time on a few of the deeper casts to find the differences between the pairs of temperature, conductivity and salinity channels. 
The sensors were used during cruise 2024-035 so a few differences from that cruise are included for comparison (shaded).
	Cast #
	Press
	T1-T0 
	C1-C0
	S1-S0
	Descent Rate

	2024-035-0083
	340
	-0.0002
	-0.00043
	-0.0042
	Mod, V. Steady

	2024-035-0083
	420
	-0.0002
	-0.00041
	-0.0042
	Mod, V. Steady

	2024-010-0026
	340
	-0.0001
	-0.00061
	-0.0060
	High, V. Steady

	2024-010-0052
	340
	 0.0000
	-0.00053
	-0.0054
	High, Moderate

	2024-001-0057
	300
	-0.0002
	-0.00046
	-0.0044
	High, F. Steady


Temperature differences are small with insignificant variation.
Conductivity differences vary more but don’t suggest temporal drift.
Salinity differences are moderate and vary in the same way as conductivity, as expected.
The salinity differences are reasonably close to the range of differences found in the comparison of the 2 CTD salinity channels with bottles (about 0.0035psu-0.0048psu). Some differences are expected since the CTDs were in motion for one and stopped for the other, and sampling was limited.
10. Conversion to IOS Header Format

The IOSSHELL routine was used to convert SeaBird 911+ CNV files to IOS Headers. 
CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number. There are no negative fluorescence values.
11. Checking Headers

A header check and cross-reference list was run. No problems were found.
Surface check – mean surface pressure was 1.8 db which is within the expected range from the Vector. The minimum pressure recorded was -0.05db. That was at the end of a cast with pumps off, but transmissivity suggests that the CTD was in water. Pressure could be reading a little low but there is no way to establish a correction. 
The track plots were added to the end of this report, one with station names and the other with cast number.
A check was made of the water depths entered in the file headers by adding the maximum depth sampled plus the altimetry reading when the CTD was at the bottom of the cast.
For 8 casts (events #7, 11, 30, 31, 32, 65,72, 83)  the check values were larger than ± 5m.
The log entries for bottom depth were checked and in 7 cases the entries in the log lead to small differences, <4db.  So those entries for bottom depth were entered in the file headers.

Only 2 of the affected casts had rosette sampling - #7 and 11.

For event 72 the altimetry signal is very clear. It is likely that there was shoaling during the cast but a bottom sounder reading was not recorded. No change was made to the header.
Changes to water depth were made to affected SAMAVG and CLN files.

12. Shift
Fluorescence

SHIFT was run on the fluorescence channel with an advance of 24 records. Plots showed this setting was appropriate for early casts with the first fluorometer and all casts using the second fluorometer. The first fluorometer did malfunction, most obviously during upcasts. Two casts checked showed good alignment in the downcasts, so it is likely that the downcast data will be salvageable. 
Dissolved Oxygen 

The Dissolved Oxygen voltage channel was aligned earlier using Seabird software. No further alignment is needed for the DO concentration channel.
Conductivity
SHIFT was run on all casts using -0.8 records for the primary conductivity and -0.5 records for the secondary based on test using a variety of settings. Salinity was recalculated for both channels.

pH
SHIFT with an advance of +24 scans was run on the pH channel based on tests. This brings the downcast profiles into better correspondence with temperature. Most of the offset in the down/up pH profiles is due to the upcast.
13. DELETE

The following DELETE parameters were used: 

Surface Record Removal: Last Press Min
Maximum Surface Pressure (relative): 10.00
Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure 
Sample interval = 0.042 seconds. (taken from header)
There were no warnings in the log.

T-S plots were examined and show some editing needed. For cast #97 the primary data look bad but the secondary are probably ok.

Profiles look reasonable with PAR generally close to SPAR at the surface.
14. Other Comparisons

Previous experience with these sensors – 
2024-035: This was the first cruise since most of the sensors were serviced in early 2024. The pressure sensor was used on many previous cruises and the last deck pressure measurement was ~0.2db. There were only 15 salinity samples and most were shallow. Based on 5 samples below 40db the primary salinity was high by an average of 00025psu and the secondary low by 0.0019psu.The differences were slightly lower using only 2 deep samples with “quiet” CTD data at +0.0020psu and -0.0015psu, respectively.

Historic ranges – All temperature data fell within the local climatology but much of the salinity data did not. The casts around Texada Island and one at station 56 had high salinity around 30db to the bottom. The casts north of Texada were within the climatology. The deep water from stations 44 to 48 in the southern Gulf of Georgia had high salinity in the deeper part of the casts, starting at between 120 and 180db and continuing to the bottom. All salinity data in Juan de Fuca Strait were within the climatology. Given that this is an area subject to occasional flux of deep water and the high salinity values are restricted to particular areas, this looks to be a real phenomenon, not evidence of calibration problems with the CTD.
Post-Cruise Calibration – There were no post-cruise calibrations available. 

15. DETAILED EDITING
The secondary T-S pair were chosen for editing and eventual archiving since the secondary had slightly less noise in T-S space than the primary and the secondary salinity was closer to bottle salinity.
All DEL files were uploaded into the CTD-QC File Processor to obtain AI generated data quality flags.
CTDEDIT was used to remove records that appear to be corrupted by shed wakes and to clean salinity where small spikes appear to be due to small misalignment or instrumental noise. 

All casts needed some editing, mostly at the top and/or bottom of casts.

The edited files were copied to *.EDT. They were zipped and uploaded to the CTD_QC file Processor.
After editing T-S plots were examined for all casts; there remain many unstable features but they are in areas when they may well reflect real conditions. A little touch-up editing was applied to 1 file. 
16. Recalibration
· The minimum CTD Salinity was ~27.4psu, above the <25psu threshold that requires correction to silicate samples. 
· There is no evidence that pressure needs recalibration.

· Salinity was not recalibrated due to limited sampling and scatter in fits. For 2024-035 salinity, the fit makes it clear there was poor flushing and there were only 2 samples below 200m. Using only those 2 samples the primary salinity is found to be high by 0.0025 and the secondary low by 0.0015psu. For 2024-010 the primary salinity was found to be high by 0.0048psu and the secondary very close to the bottle. However, based on 2 samples near the mouth of Juan de Fuca where flushing is usually better the primary is high by 0.0052psu and the secondary high by 0.0017psu.
The secondary salinity is likely within 0.003psu.
· Recalibration of SBE Dissolved Oxygen was based on the bottle comparison; the results of the previous use were not useful due to few samples and a very small range of DO values.

File 2024-010-rcal1.ccf was prepared to apply the following correction to channel Oxygen:Dissolved:SBE in the MRGCLN2 files and the SAM files:
CTD DO Corrected = CTD DO *1.0324 + 0.007

COMPARE was rerun on dissolved oxygen after recalibration of the SAM files were recalibrated.
The fit is much improved with DO reading high by an average of 0.006mL using the points that were included in the original comparison.

See 2024-010-dox-comp2.xls for details.
CALIBRATE was then run on the EDT files.

17. Final Comparison of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and delayed response and noise in CTD data.

Downcast recalibrated files (bin-averaged to 0.5m-bins) were thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles. Outliers were removed based on residuals in a fit against DO. Plots were then made against pressure using the same data.  

The SBE DO sensor has a fairly long response time so data accuracy is not as high when it is in motion as it is during stops for bottles. This will be especially true when vertical DO gradients are large. To get an estimate of the accuracy of the SBE DO data during downcasts (after recalibration) a rough comparison was made between downcast SBE DO and upcast titrated samples. Some of the difference will be due to problems with flushing of Niskin bottles and/or analysis errors, so the following statement likely underestimates SBE DO accuracy. 
Downcast (CTD files) Oxygen:Dissolved:SBE data for this cruise are considered, very roughly, to be:

      ±0.15 mL/L from 0-50db

      ±0.10 mL/L from 50db to 100db 
      ±0.05 mL/L from 100db to 200db 
      ±0.02 mL/L below 200db.

For details see files 2024-010-dox-comp3.xls

18. Fluorescence Processing
A median filter with a fixed width of 11 was applied to the fluorescence channel.
19. BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure
Averaging interval = 1.000
Minimum bin value =   .000

Average value will be used.
Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined. There are some small unstable features that are likely real due to active mixing and/or ship drift during the cast. 
Profile plots were examined to see if there any problems. 
· The cap was obviously left on the pH sensor during casts 24, 53, 90.  That channel will be removed for those casts. (There was no rosette sampling during those casts.)

· Fluorescence has a sudden shift to high values between 230db and 240db during cast #35. No other variable shows anything unusual at this level and the sudden start/stop indicates this was not a real feature. Those values were padded and the channel limit was corrected in the header section.

No other problems were noted.
20. Final CTD File Steps (REMOVE and HEADEDIT)
· For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Primary, Conductivity:Primary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T0:C0, Flag and Prediction_Flag.

For events 24, 53 and 90 channel pH:SBE was removed.

· A second SBE DO channel (with umol/kg units) was derived and REORDER was run to place the 2 DO channels together. 
· DO saturation was calculated and values ranged from 55% to 110% at the top of casts, with the exception of Saanich Inlet where it was about 160%. The lowest values were in the Gulf Island and Juan de Fuca areas, which is normal. As usual in inshore cruises, this test offers little guidance as to the calibration of the oxygen sensor. 
· HEADER EDIT was used to fix formats and channel names and to add comments to the headers. 
· The Standards Check routine was run and no problems were found. 
· The Header Check was run; no problems were found. 
· Profile and T-S plots were examined and no problems were found. 
· The sensor history was updated. 

21. Final Bottle Files
· The MRGCOR1 files were put through SORT to order on increasing pressure. 
· For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Primary, Conductivity:Primary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T0:C0 and Flag.
      
There was no rosette sampling during the casts that had the cap left on the pH:SBE sensor.
· A second SBE DO channel with mass units was added for both the CTD DO and REORDER was run to get the pair of DO channels together.
· HEADER EDIT was run to ensure formats and units are correct; the geographic area entry was corrected.
· Standards check was run and a few problems were found and resolved.
· Bottle data were exported to a spreadsheet and no problems were found.

· A header check were run. No problems were found. 
· A cross-reference listing was produced for the CHE files.
22. Thermosalinograph

a.) Initial steps

There was 1 thermosalinograph hex file. 

The configuration file created in acquisition was used to convert the hex file to CNV format.

There was no external thermistor or fluorometer in use.
The file was converted to IOS Header format. This includes Start Time.
As mentioned in the log the conductivity values are very low, so salinity and conductivity will not be archived.
CLEAN was run to add Start and End times and Longitude and Latitude minima and maxima to the headers. 
ADD TIME CHANNEL was used to add Time and Date channels. (Julian Time was already included.)
The track plot was run and looks fine. It was added to the end of this report.

A time-series plot was produced. An initial examination shows temperature data that look ok and salinity data that are clearly bad.
b.) Time checks
Checks were made that the time in TSG files are accurate. 
· The CTD files were thinned to reduce the files to a single point from the downcast within 0.5db of 2db. These were exported to a spreadsheet which was saved as “2024-010-CTD-TSG-comp.xlsx”.There were 65 CTD casts that overlapped with TSG records, and had data within 0.5m of 2m.  

· The TSG ATC file was opened in EXCEL and reduced to times of CTD data in the spreadsheet. They were added to 2024-010-CTD-TSG-comp.xls.

· To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. The median differences were <0.0000º for both latitude and longitude, but there were some large differences during events 31, 32, 37, 41, 83, 84 and 89; those are locations where some drift during the cast is likely. Overall, the matches look ok, so the TSG can be compared with the CTD.

c.) Calibration studies

Next the differences between the TSG and CTD temperature at 2m were examined. 

	
	Lat Diff
	Long Diff
	Tlab-Tctd

	Min
	-0.0016
	-0.0021
	-0.7654

	Max
	0.0016
	0.0036
	-0.0086

	Std Dev
	0.0004
	0.0007
	0.1119

	Average
	0.0000
	0.0001
	-0.3022

	Median
	0.0000
	0.0000
	-0.2722


The following plot shows the temperature differences and the CTD temperature at 2m versus event numbers. We might expect more heating when the intake temperature is lower, but that does not seem at all obvious. There is too much noise in the differences to pick out a useful pattern. However, examination of surface temperature profiles in CTD files shows that late in the cruise, in Juan de Fuca Strait, the upper 10m are very well mixed, so small differences in the depth from which the TSG draws water and the data included in the CTD files will have little effect on the differences in temperature. The averaging in the CTD files may lead to the water coming from a little deeper in the water column than the TSG intake, though this will be highly variable.
The average seen in Juan de Fuca is likely the best estimate of heating in the loop. Using events 82 to 92, the average difference is 0.30Cº. Subtracting that amount from the temperature in the lab will enable an estimate of seawater temperature.

[image: image5.png]TSG Temp - CTD Temp

(TSG Temp-CTD Temp)and CTD Temp vs Event

56660
S

5665
o> Yoo

Event#

@®Tlab-Tctd @ Temperature:CTD

CTD Temp at 2m




A plot of heating in the loop versus intake temperature shows a lot of variability in the area with higher temperatures. That is where the largest near-surface gradients are found including some with temperature inversions. 
[image: image6.png]TSG Temp - CTD Temp

0.0000
-0.1000
-0.2000
-0.3000
-0.4000
-0.5000
-0.6000
-0.7000
-0.8000
-0.9000

(TSG Temp-CTD Temp) vs CTD Temp

9 10 " 14
.
o &
%%, %o
% K .
. .
°
°

CTD Temperature




See 2024-010-CTD-TSG-comp.xlsx for details.

d.) Calibration history

The TSG has been used during 2 previous cruises since it was last serviced.

2024-001: T high by median 0.26C; Sal low by median 0.009psu.

2024-004: Salinity differences too variable for recal but likely within 0.006psu for early part of cruise; some bad data removed later. A proxy for intake temperature was made by subtracting 0.20Cº.
e.) Conclusions

· The TSG clock worked well. 
· The conductivity & salinity data are bad. 
· An estimate for seawater temperature can be made by subtracting 0.30Cº from the temperature in the lab. This temperature will be called Temperature:Primary and will be a proxy for intake temperature.
f.) Addition of intake temperature proxy and Calibration

Add Channel was run to add channel Temperature:Lab, setting it equal to Temperature:Primary and channel Pressure set to 2.0db to enable derivation of sigma-T.

CALIBRATE was run to subtract 0.30Cº from Temperature:Primary.

Derived Quantities was not run to calculate sigma-T since salinity data are bad.
REMOVE was used to remove the following channels: Scan Number, Flag and Pressure.

g.) Preparing Final Files 

HEADER EDIT was used to change the DATA DESCRIPTION to THERMOSALINOGRAPH and to change channel names to standard names and formats. Comments about the data were added.

The TSG sensor history was updated. 
NOTE: There is a file called 2024010_Surface_PAR.xlsx. These data did not come from the CTD. There is no accompanying information about how to calculate PAR from the data in the file. It was not used in processing.
Particulars
10. Brought back to surface after soak because pH buffer not removed. Turned pump on and off.

15. Pump off for a few minutes because possible SAR call. 

15. Large fluorescence spike ~130m on upcast. Stayed consistently high to surface.

17. Fluorometer dis-attached from pump from 75m to 12m on upcast. 

      Depth sounder showing inconsistent depths.

24. pH buffer left on.

31. Fluorometer giving high values on way up. Resolved itself.

34. Scott cleaned fluorometer cable to prevent spiking
35. Lots of swell. Still problems with fluorometer spiking badly.

36. Swapped fluorometers after this event with sensor 3641.
53. pH buller left on.

83. Stn Name wrong in header. Should be 70, not 83.
91. pH buffer left on.

Thermosalinograph: Conductivity was very low to start and then got worse and worse. Needs service.
CRUISE SUMMARY – CTD
	CTD#
	Make
	Model
	Serial#
	Used with Rosette?
	CTD Calibration Sheet Competed?

	1
	SEABIRD
	911+
	1453
	Yes
	Yes

	Calibration Information - 1453

	Sensor
	Pre-Cruise
	Post Cruise

	Name
	S/N
	Date
	Location
	Date
	Location

	Temperature


	2023
	8Feb2024
	Factory
	
	

	Conductivity
	3500
	23Feb2024
	Factory


	
	

	Secondary Temp.


	6565
	17Feb2024
	Factory


	
	

	Secondary Cond.


	5043
	23Feb2024
	Factory


	
	

	Transmissometer


	1201DR
	11 Jan2024


	IOS
	
	

	PAR
	4615
	24Feb2024
	Factory
	
	

	SPAR
	20518
	16Feb2024
	Factory
	
	

	SBE 43 DO sensor
	1438
	27Feb2024
	Factory
	
	

	SeaPoint Fluor.
	3640
	?
	
	
	

	SeaPoint Fluor.
	3641
	?
	
	
	

	pH
	0691
	9Mar2023
	Factory
	
	

	Pressure Sensor
	1453
	21July2022
	Factory
	
	

	Altimeter
	73171
	30Jan2024


	Factory
	
	


	Calibration Information – TSG 3411

	Sensor
	Pre-Cruise
	Post Cruise

	Name
	S/N
	Date
	Location
	Date
	Location

	Temperature
	3411
	25Jan2024
	Factory
	
	

	Conductivity
	3411
	25Jan2024
	Factory
	
	


[image: image7.png]2024-010 Event #s

125,50 125,00 1245 124,00 172,50 122,00 122,50
50.50 . . . . . 50.50
50.00 - 50.00
'g 43,50 - 49,50
=
E
=
<
3
=
=1
S @0 | .00
8.5 | .50
8.0 : : : : : 8.0
125.50 125.00 124,50 124.00 123.50 123.00 122,50

West Longitude




[image: image8.png]2024-010 Stn Names

125,50 125,00 1245 124,00 172,50 122,00 122,50
50.50 . . . . . 50.50
50.00 - 50.00
'g 43,50 - 49,50
=
E
=
<
3
=
=1
S w0 | .00
8.5 | .50
8.0 : : : : : 8.0
125.50 125.00 124,50 124.00 123.50 123.00 122,50

West Longitude




[image: image9.png]North Latitude

2024-010 TSG

124.00
I

123,50
I

122.50

50.50

50.00~

43,50

49,00

48,50

- 50.00

I~ 48.50

49,00

- 4B.50

124,00
- West Longitude




PAGE  
13

