
REVISION NOTICE TABLE

	DATE
	DESCRIPTION OF REVISION

	
	


PROCESSING NOTES
Cruise: 2024-013




Agency: OSD

Location: Strait of Georgia, Juan de Fuca Strait
Project: Acoustic Moorings
Chief Scientist: O’Neill C.
Platform: Vector
Date: 3 July 2024 – 9 July 2024 
Processed by: Germaine Gatien
Date of Processing:  2 December 2024 – 11 December 2024
Number of HEX files: 80

Number of CTD files processed: 80
INSTRUMENT SUMMARY
Two CTDs were used during this cruise. There was no rosette.
For Events 1 – 59 CTD #0585 was used with an SBE 43 DO sensor on the primary pump (#1176), SeaPoint Fluorometer (#3950) with 3X gain on the secondary pump and an altimeter (#73171) attached.
For Events 60 – 101 CTD #1453 was used with a Wetlabs CSTAR transmissometer (1201DR), a SBE 43 DO sensor on the primary pump (#1438), SeaPoint Fluorometer (#3640) with 3X gain on the secondary pump and an altimeter (#73171) attached. (Fluorometer not present for cast #60 only.)
(Note that a Transmissometer and PAR and SPAR sensors were listed in the configuration file but were not actually present. For cast #60 only no altimetry or fluorescence data were acquired.)

Seasave version 7.26.7.121 was used for acquisition.
SUMMARY OF QUALITY AND CONCERNS
The digital Daily Science Log Book was in good order with useful comments about the CTD. 
There was only one previous cruise that used CTD #0585 since it was last serviced and the sampling data were not available at the time this cruise was processed; that sampling was also primarily very shallow so not likely to provide useful information for calibration purposes. The sensor pairs were in excellent agreement so both are likely well calibrated. The primary salinity looks slightly smoother, so the primary channels were chosen for editing and archiving.
The pressure sensor on CTD #1453 has performed well in previous uses with no need for recalibration, but there was a deck measurement of 1.2db. There was no evidence in the data that the pressure was out of line, so with only a single deck measurement to go by, no recalibration was applied. The other sensors on that CTD have not been used since the last factory service. The temperature channel differences are a little higher than expected so soon after factory service. The salinity differences were about 0.004psu. There is no way to determine which is more accurate, though the secondary salinity generally looked smoother so the secondary channels were selected for editing and archiving.
During cast #59 all data were thought to look poor, so the CTD was switched. However, the data look fine, so the problem was more likely with the display.
The configuration file for CTD #1453 included information about a transmissometer, PAR and SPAR, but there is no entry for such instruments in the log and no data were found in those channels, so it is presumed there was no plan to use them. Altimetry and fluorescence data were not collected during cast #60 due to some set-up issues, but both were collected during all other casts. 
PROCESSING SUMMARY

1. Seasave
This step was completed at sea; the raw data files have extension HEX.

2. Preliminary Steps
· The digital Log Book was obtained.
· The cruise summary sheet was completed. 

· There was no calibration sampling for the salinity or dissolved oxygen sensors.
· Two CTDs were used for this cruise. 
· There was only one previous cruise that used CTD #0585 since it was last serviced and the sampling data were not available at the time this cruise was processed; that sampling was also primarily very shallow so not likely to provide useful information for calibration purposes. 
· The pressure sensor on CTD #1453 has performed well in previous uses with no need for recalibration. The other sensors on that CTD have not been used since the last factory service. 
The configuration files used at sea were checked. 

There was an error in the offset for transmissometer #1201 on CTD #1453; this was corrected. 

The configuration files were saved as 2024-013-ctd1.xmlcon (files 1-59) and 2024-013-ctd2.xmlcon (files 60-108). 

For cast #60 there was no altimeter and the fluorometer signal was bad.

For cast #61 the altimeter was installed and the fluorometer cable was switched.

3. Conversion of Full Files from Raw Data

All files were converted using files 2024-013-ctd1.xmlcon and 2024-013-ctd2.xmlcon.
Altimetry was converted for cast #60 so that all files have the same channels, but it was not actually mounted. The header entry will be removed later as it is based only on noise.
The Tau function was selected but not the hysteresis function since there was no deep sampling. 

A few casts were examined and all expected channels are present.
4. WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, depth, conductivity & temperature only in the full cast files (*.CNV).  

Parameters used were: 
Pass 1    Std Dev = 2 
Pass 2    Std Dev = 5 
Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

5. ALIGN DO

ALIGNCTD was run on all casts using +2.5s.

6. CELLTM

CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

7. DERIVE and Channel Comparisons
Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration. 
Plots were checked to see how the various steps worked.

· The temperature and conductivity channels are fairly close during downcasts but noisier and further apart during upcasts. There are large differences between downcasts and upcasts for many casts, which is likely due to the areas being sampled having high variability.  
· WILDEDIT did a good job getting rid of single-point spikes.

· The DO alignment looks reasonably good. In some spots it looks like the advance was too large and in others too small, but generally downcast features match those in temperature quite well, especially in some of the deeper casts.
· CELLTM worked as expected, lowering salinity where temperature is decreasing (cell warms water raising conductivity but temperature sensor reading is correct, so salinity is overestimated). Conversely, it raises salinity where temperature is increasing (conductivity too low). This brings downcast and upcast traces closer together, so T-S downcast and upcast plots are in better agreement after this step. Downcast and Upcast T-S curves also moved closer together where temperature is decreasing with depth.
Differences between channel pairs are hard to assess due to shallow casts and active mixing, but estimates were made in a few casts for each CTD. 
	Cast #
	Press
	T1-T0 
	C1-C0
	S1-S0
	Descent Rate

	2024-013-0008
	220
	-0.0004
	~0
	~0
	High, V. Steady

	2024-013-0022
	240
	-0.0006
	+0.0002
	+0.0008
	Mod, X. Noisy

	2024-013-0068
	320
	-0.0012
	-0.0004
	-0.0036
	High, V. Steady

	2024-013-0092
	300
	-0.0017
	-0.0005
	-0.0037
	

	2024-013-0107
	305
	-0.0013
	-0.0005
	-0.0042
	High, V. Steady


The differences for CTD #0585 are very small. Cast 8 was in Haro Strait which was very well mixed, and cast 22 came from the offshore area where fairly reliable differences may be found in shallow water. 
For CTD #1438 the differences in temperature are somewhat higher than usually seen which is partially offset by conductivity differences, so that salinity values differ by ~0.004psu. These differences are a little higher than we would expect from recently calibrated sensors, though such judgments are less reliable in shallow water.
8. Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers. 
CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.
9. Checking Headers

· The cross-reference check and header check were run. The station names are long so don’t show up fully.
· Surface check was run and found an average  of 2.5 for CTD1 and 2.2db which are reasonable. depths for  starting acquisition. 
· All files were put through REVERSE and surface check was rerun looking for casts stopping with pressure <1db but none were found for either CTD. 
· Header check was run for each CTD to find minimum pressures recorded. 
· They were 0.76db for CTD #1 near the end of event #21. The conductivity and salinity do not suggest the CTD was right at the surface. 
· For CTD #2 the minimum pressure was 1.58db during event #83. The CTD was clearly in water at that point. The upcasts generally stopped at about 2db for both CTDs so there is no evidence to suggest the pressure is reading low other than the one deck measurement. Such measurements often vary, possibly with temperature. The history of the sensor shows no pressure drift including a deck pressure of -0.2db in October 2023. No recalibration is justified. 
· Cruise tracks were plotted and no problems were found and added to the end of this report. 
10. Shift
Fluorescence

SHIFT was run on the SeaPoint fluorescence channel in all casts using the usual advance of +24 records. Plots show that the vertical offset between downcast and upcast for fluorescence is reasonably close to that for temperature after this step, although some CTD #2 casts had bad upcast fluorescence.
Dissolved Oxygen 

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked good, and it looked very good for CTD #1 and generally good for CTD #2, though as noted earlier, there are some cases where the correction looks too large or too small. No further alignment was applied to the DO concentration channel.
Conductivity
Tests were run on a few casts to assess what settings are best to align conductivity with temperature (as judged by the effect on salinity as seen in T-S space). The best settings were -0.5 record for the primary and -0.9 records for the secondary for CTD #0585 and -0.5 for both primary and secondary for CTD #1438. 

SHIFT was run twice for each CTD. Salinity was recalculated.

11. DELETE

The following DELETE parameters were used: 

Surface Record Removal: Last Press Min
Maximum Surface Pressure (relative): 10.00
Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure 
Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.
The DEL files were copied to *.edt so there will be a complete set even if no editing is required for any casts.
12. Other Comparisons

Experience with these sensors since last factory service –  
#0585 – All sensors for one other cruise that has not yet been processed; all sampling from that cruise is shallow and not likely to provide reliable recalibration information.
#1438 – Pressure sensor has had many uses with good results, no recalibration found necessary. T, C and DO sensors have no history of use since last recalibration.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. All salinity fell within the climatology except for low salinity around 100db for 2 of the Juan de Fuca casts and high salinity below 100db at station 56. As usual for that site, the salinity data fell within the climatology for the Gulf Islands. Near-surface temperatures were frequently above the maximum of the climatology in the Strait of Georgia and Gulf Islands; this is likely not a problem with calibration but reflection of real change not represented in the climatology. .
Post-Cruise Calibration – There were no post-cruise calibrations available. 

13. DETAILED EDITING
The primary channels were chosen for recalibration for CTD #0585 and the secondary for CTD #1438.
The differences were small.
The DEL files were submitted to the CTD_QC File Processor. The DELPRED files produced will be used for editing.
Most casts required light editing at the top and/or near the bottom where the instrument descent rates were very low. 
Editing was kept light in the inshore areas where descent rates were quite steady and local mixing leads to complex gradients.

Editing of casts in the offshore area was difficult since there were many shed wakes due to noisy descent rates combined with active mixing. Data were removed where they were obviously due to shed wakes but in many cases where it was not clear, data were left unedited.
Casts that did not require editing were: 8, 9, 10, 28, 51, 71, 91, 108.
Notes about editing applied were added to the files.
The edited files were copied to *.EDT.
After editing T-S plots were examined for all casts. Unstable features remain but these look likely to be real features; no further editing was applied.
The EDT Files were submitted to the CTD_QC File Processor.
14. Corrections to Pressure, Salinity and Dissolved Oxygen Concentration

There is no indication of a problem with pressure calibration and there is no information available upon which to base recalibration of salinity. No recalibration was applied.
15. Fluorescence Processing
A median filter, size 11, was applied to the fluorescence channel in the COR1 files. (Output:*.FIL)
16. BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure
Averaging interval = 1.000
Minimum bin value =   .000

Average value will be used.
Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined. No problems were found.
Profile plots were examined. No problems were found
17. Final CTD File Steps (REMOVE and HEADEDIT)
For casts 1-59 REMOVE was run to remove the following channels:

Scan_Number, Temperature:Secondary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, 
For casts 60-108 REMOVE was run to remove the following channels:

Scan_Number, Temperature:Primary, Conductivity:Primary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T0:C0 and Flag.
For cast #60 only, Fluorescence was removed.
The altimetry header was removed from event #60 as the sensor was not actually mounted.
Dissolved Oxygen was derived in mass units and that was used to calculate DO saturation. 
Plots of near-surface saturation show a range of 55% to 125%, with the exception of cast #41 (station Nitinat North), a cast in shallow water near, which was 190%.. The SBE DO value at that site was >11.8mL/L at the surface.
The low values are all from active mixing regions including Haro Strait and Juan de Fuca Strait.

The values from the southern Strait of Georgia are in the range 105-125%.

A check value was calculated by adding maximum depth sampled plus altimetry at bottom of the cast minus the water depth entry in the header. Where that difference was greater than 5m, the altimetry value and depths in headers were checked. There were 7 such entries, with only 2 having differences of more than 10m. First the altimetry was checked and looks fine except for cast #60 where it was not actually mounted but there was a signal. Casts investigate were:
· For events 58, 65, 74, 75 the log entries for water depth differed from the header entries. When the log depth was used the check value was within 5m for all 4 casts. The water depth in the headers were changed to match the log entries.
· Event 54  has the same water depth header as in the log and altimetry looks fine. This cast was in Haro Strait where shoaling is likely. Since the difference is <10m no change was made to the header.

· Event 60 is likely fine – the altimetry is bad and was removed from the header.  No change was made.

· Event 94. There is no cast 94 in log – this should be cast #96. Minimum altimetry looks good at 21m. Suggests depth was 258. Log entry for #96 is 248. Using that water depth the check value is 10m. Changed water depth entry to 258.
Having found the error in file names for 94/96, all file names were checked against the log and one other error was found:
· Event 95. This should be #99. The water depth is correct for that site. The check value was good. No change needed except to event #. 
REORDER was used to get the 2 dissolved oxygen channels together.
HEADER EDIT was used to fix formats and channel names and to add comments about processing. 
The Standards Check routine was run and no problems were found. 
The Header Check was run; no problems were found. 
Profile and T-S plots were examined. No problems were found.
The sensor history was updated. 
Particulars

32. Stop on downcast at 50m for ship to reposition onto station due to drift.
43. Pumps off during upcast.

44. Pumps ok.

59. Data in all channels reported to look bad, but seems ok.
60. Switch to secondary CTD 1453. No altimeter or fluoro readings - altimeter not on package, fluorometer not plugged in correctly. Conductivity/Temp readings OK but ~90m from bottom altimeter didn't kick.
61. Added altimeter from primary CTD to secondary (SN# 73171) into AUX2, split AUX1 so both DO and fluoro sensors going in. Swapped fluoro cable to one from other package - same gain. 
61. Moved SCStar_1 SSE to avoid traffic and test cast new CTD.
62. Fluoro still not going to 0. But profiles for other things are much improved and altimeter working great.
63. Deck pressure = 1.2
64. Switched fluorometer cable out to try to help with noise. Fluorometer is still super noisy but a bit better with different cable. Going to leave it as-is.
65. Bottom depth variable while on station - lots of topography/drift
70. Super large fluoro spikes. Very unhappy fluorometer.
74. Used 071.xmlcon file - this calls on the correct sensor/config data but just not from the con file folder.

75. Used 071.xmlcon file

78. Used 071.xmlcon file

79. Used 071.xmlcon file
80. Back to confile 9p-1453 4 june 2024
81. Fluorometer seems to have settled down as much as a fluorometer can.

89. Fluoro at ~24 on upcast in upper 10m that wasn't seen on downcast.
91. Slightly NE of station coords - off by ~1 cable
94 & 95. See 96 and 99. Mooring casts.
96. No CTD data found. Log says “Altimeter read 20m, but depth sounder read 248m (floor depth). Went with sounder depth for 10m off the sea floor.”  CTD file #94 found which in the log is a Mooring cast. So this file was renamed as #96 at the end of processing.
99. No CTD data found. File #95 was renamed as #99 at end of processing.
CRUISE SUMMARY – CTD

	CTD#
	Make
	Model
	Serial#
	Used with Rosette?
	CTD Calibration Sheet Competed?

	1
	SEABIRD
	911+
	585
	No
	Yes

	2
	SEABIR
	911+
	1453
	No
	Yes

	Calibration Information - 585

	Sensor
	Pre-Cruise
	Post Cruise

	Name
	S/N
	Date
	Location
	Date
	Location

	Temperature


	2449
	22 Feb 2023
	Factory
	
	

	Conductivity
	1764
	17 Feb 2023
	Factory


	
	

	Secondary Temp.


	4484
	22 Feb 2023
	Factory


	
	

	Secondary Cond.


	2128
	17 Feb 2023
	Factory


	
	

	SBE 43 DO sensor
	1176
	01 Mar 2023
	Factory
	
	

	Altimeter
	73171
	30 Jan 2024
	Factory
	
	

	Pressure Sensor
	1222
	20 Feb 2023
	Factory
	
	

	SeaPoint Fluor.
	3950
	May 2023
	
	
	


Calibration Information -1453
	Sensor
	Pre-Cruise
	Post Cruise

	Name
	S/N
	Date
	Location
	Date
	Location

	Temperature


	2023
	8 Feb 2024
	Factory
	
	

	Conductivity
	3500
	23 Feb 2024
	Factory


	
	

	Secondary Temp.


	6565
	17 Feb 2024
	Factory


	
	

	Secondary Cond.


	5043
	23 Feb 2024
	Factory


	
	

	Transmissometer
	1201DR
	11 Jan 2024
	IOS
	
	

	SBE 43 DO sensor
	1438
	27 Feb 2024
	Factory
	
	

	SeaPoint Fluor
	3640
	14 Mar 2024
	
	
	

	Altimeter
	73171
	30 Jan 2024
	Factory
	
	

	PAR *
	4615
	24 Feb2021
	Factory
	
	

	SPAR *
	20518
	24 Feb2021
	Factory
	
	

	Pressure Sensor
	1453
	17 Mar 2022
	Factory
	
	


* Listed in configuration file but not in use during this cruise.

[image: image1.png]2024-013 Event #s

125,00 124.50 124.00 123.50 123.00
49.40 L L ) 48.40
43,20 - 48,20
43,00 45,00
5]
=1
-]
=1
=
<
— 48.80+ 5 48,80
o
=
=}
Z
48,60 - 4B.60
48,40 - 4B.40
48.20 T T T 48.20
125,08 124,58 124,00 123,50 123.00

‘West Longitude




[image: image2.png]125,00

124.50
I

2024-013 Stn Names
12400

123.00

49.40

43,20

43,00

48,80

North Latitude

48,60

48,40

SofSstar 8
SoBstar 1 sofsstar 7

SoBSstor_2

S0 St 5

48.40

45,00

48.20
125,08

T
124,58

124,00
West Longitude

48.20

123.00




PAGE  
1

