## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2024-019				Agency: OSD

Location: North-East Pacific			Project: La Perouse

Chief Scientist: Sastri A.				Platform: John P. Tully

Date: 14 - 27 May 2024

Processed by: Jeannette Bedard 		Date of Processing: August-November 2025

Number of original HEX files: 94			Number of processed CTD files: 94

Number of rosette casts: 58			Number of processed CHE files: 48

Number of original TSG txt files: 4		Number of processed TOB files: (1 per day) 13

# INSTRUMENT SUMMARY

SeaBird 911+ CTD #0443 was mounted in a rosette and attached were 2 Wetlabs CSTAR transmissometer (1185DR & #1883DG), a SBE 43 DO sensor on the primary pump (#4372), SeaPoint Fluorometer on the secondary pump (#4186), a WetLabs fluorometer (#2216), a Biospherical QSP-400 PAR sensor (#70613) and a Valeport altimeter (#76341).

A thermosalinograph (SeaBird 45 S/N 0620) was mounted with a Wetlabs WETStar fluorometer (#1656), intake thermometer (SeaBird 38 #603) and flow meter; sampling interval was 5s.

Seasave version 7.26.7.121 was used for acquisition. The data logging computer was TULLY. The deck unit was a Seabird model 11+ #508. An IOS rosette with 24 10L bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log Book was in excellent order with comments about problems encountered and a detailed list of equipment. The sampling logs were in good order.

The standard deployment procedure for this cruise as follows:

The rosette was brought to the surface. Pumps were turned ON. The rosette was brought down to 10m and kept there for 30 seconds. Once back at the surface, the data started to be archived, with the rosette at the surface for 30 seconds longer. Then the cast would start.

There were 2 WetLabs CStar transmissometers in use during this cruise:

- Channel Transmissometer refers to sensor #1185DR (650nm - red)

- Channel Transmissometer:Green refers to sensor #1883DG (530nm - green)

For comparison with other Institute of Ocean Sciences cruises, note that the transmissometer wavelength is 650nm unless otherwise stated.

The SBE DO sensor has a fairly long response time so data accuracy is not as high when it is in motion as it is during stops for bottles. This will be especially true when vertical DO gradients are large. To get an estimate of the accuracy of the SBE DO data during downcasts (after recalibration) a rough comparison was made between downcast SBE DO and upcast titrated samples. Some of the difference will be due to problems with flushing of Niskin bottles and/or analysis errors and small mismatches in depth in the presence of large DO gradients, so the following statement likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data for this cruise are considered, very roughly, to be:

<0.08mL/L for pressure <200db

<0.04mL/L for pressure between 200-300db

<0.02mL/L for pressure >300db

Despite the flow meter for the TSG not reporting a flow, the data appeared to be fine (suggesting the flow meter was faulty). There was good agreement between the different data types.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

- The Log Book and rosette log sheets were obtained.

- Nutrients, extracted chlorophyll, dissolved oxygen, salinity and NH4 data were obtained in QF spreadsheet format from the analysts.

- The cruise summary sheet was completed.

The configuration files changed during the cruise. Before cast 5, the Seapoint fluorometer was dummied off. Files used in conversion were:

- 2024-019-v1.xmlcon for events #1-4 (CTD #0443)

- 2024-019-v2.xmlcon for events #5-218. (CTD #0443 without Seapoint fluorometer)

##### BOTTLE FILE PREPARATION

- ROS files were created using 2024-019-ctd*.xmlcon. The hysteresis correction and tau corrections were selected.

- The IOS files were put through CLEAN to create BOT files.

- Temperature and salinity were plotted for all BOT files to check for significant outliers.

- The BOT files were bin-averaged on bottle number.

- The output was used to create file ADDSAMP.csv.

- The file was then sorted on event number and Bottle Position order. Then sample numbers were added based on the rosette logs.

  - Cast 40 – bottles tripped out of order

- The ADDSAMP file was then reordered on event # & sample #.

- The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

- The ADDSAMP file was used to add sample numbers to the BOT files – output *.SAM.

- The SAM files were bin-averaged on bottle # and called SAMAVG.

- Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create the header comments file which was updated during processing.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2024-019_OXY*.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and saved as 2024-019_OXY.csv. That file was converted into individual *.OXY files.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2024-019_CHL QF*.xlsx. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared and saved as 2024-019_CHL.csv. The csv file was then converted to individual CHL files.

SALINITY

Salinity analysis was obtained in file QF2024-019_SAL.xlsx which included a precision study. The analyses were carried out in a temperature-controlled lab 11 to 25 days after collection. The files were simplified and saved as 2024-019_SAL.csv. That file was then converted to individual SAL files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2024-019_NUTS*.xlsx. This includes a precision study. The file was simplified, saved as 2024-019_NUTS.csv. The file was converted to individual NUTS files.

NH4

The NH4 data were obtained in spreadsheet QF2024-019_NH4*.xlsx. This includes a precision study. The file was simplified, saved as 2024-019_NH4.csv. The file was converted to individual NUTS files.

The SAL, CHL, OXY, NUTS, NH4 files were merged with CST files in 5 steps.

The files were then put through CLEAN to reduce the headers to File and Comment sections only.

These files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number and saved as *. MRGCLN1s.

The MRGCLN1s files were then merged with SAMAVG files using merge channel Bottle_Number.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions.

CTD salinity was checked and the lowest value was ~29 PSU well above the 25 PSU threshold for needing a silicate correction. No correction to the silicate was applied.

##### Compare

Salinity

Compare was run with pressure as reference channel. Only a small salinity correction was needed:

- CTD Sal Corrected = CTD Sal - 0.001

For full details for the COMPARE run see file 2024-019-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel. Only a small dissolved oxygen correction was needed:

- CTD DO Corrected = CTD DO  - 0.0037

For full details for the COMPARE run see file 2024-019-oxy-comp1.xls. A second comparison was done and a second correction applied:

- CTD DO Corrected = CTD DO  + 0.0019

For full details for the COMPARE run see file 2024-019-oxy-comp2.xls.

##### Conversion of Full Files from Raw Data

The configuration files changed during the cruise. Before cast 5, the Seapoint fluorometer was dummied off. Files used in conversion were:

- 2024-019-v1.xmlcon for events #1-4 (CTD #0443)

- 2024-019-v2.xmlcon for events #5-218. (CTD #0443 without Seapoint fluorometer)

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, depth, conductivity & temperature only in the full cast files (*.CNV).

- Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

- The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

ALIGNCTD was run on all casts to advance the oxygen voltage by +2.5s, a setting which has worked well in the past for this type of sensor.

##### CELLTM

The noise in the upcast data makes tests for the best parameters for this routine very difficult to interpret. In the past when upcast data were not so noisy, the default setting of (α = 0.0245, β=9.5) was generally found to be the best choice—these values were used here.

##### DERIVE

Program DERIVE was run on all casts to calculate primary and secondary salinity, dissolved oxygen concentration (ml/l) and descent rate.

IOS Shell

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert the final SeaBird CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

- The cross-reference check and header check were run. No problems were noted.

- Surface check was run for each CTD. Average depth of the surface mixed later was 3.47 db. With a minimum of 3.92 db and a maximum of 5.03 db.

- Cruise tracks were plotted and added to the end of this report.

The altimeter and water depth readings from the headers were exported to a spreadsheet. A check value was calculated by subtracting water depth from maximum depth sampled plus altimetry header. Five casts had values much greater that +/- 5m and in all cases it was due to the bathymetry changing while the cast was taking place. The following corrections were made:

- Cast 52 – File listed max depth as 1170 m, but the log listed it as 1180 m – cast file was corrected

- Cast 57 – File listed max depth as 1775 m, but the log listed it as 1808 m – cast file was corrected

- Cast 62 – File listed max depth as 1443 m, but the log listed it as 1465 m – cast file was corrected

- Cast 139 – File listed max depth as 1046 m, but the log listed it as 1093 m – cast file was corrected

- Cast 160 – File listed max depth as 1385 m, but the log listed it as 1425 m – cast file was corrected

Note:

- The primary and secondary temperature and salinity were closely aligned in all casts. Upon manual inspection if was found that the secondary sensors had more spikes than the primary. For further processing the primary sensors were chosen.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel using the usual advance of +24 records.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it did. No further alignment was needed for the DO concentration channel.

Conductivity

Tests were run on a selection of casts to find the alignment shift best for the conductivity sensors as judged by noise in T-S space. A shift of -0.7 was chosen. After the shift was applied, salinity was recalculated.

##### DELETE

DELETE was run on the shifted files. The following DELETE parameters were used:

- Surface Record Removal: Last Press Min

- Maximum Surface Pressure (relative): 10.00

- Surface Pressure Tolerance: 1.0                  Pressure filtered over 9 points

- Swells deleted. Warning message if pressure difference of 2.00

- Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

- Drop rate applies in the range:  10db to 10db less than the maximum pressure

- Sample interval = 042 seconds. (taken from header)

##### DETAILED EDITING

- The *.DEL files were zipped in and submitted to the CTD-QC-File Processor. DELPRED files were returned.

- Most casts had small shed wakes and spikes which were removed along with all data captured when the descent rate was below 0.2 m/s.

- After editing, T-S plots were examined for all casts and no further editing was found necessary for temperature and salinity; a few unstable features remain but may be real.

##### Corrections to Pressure, Salinity and Dissolved Oxygen Concentration

Same as for bottle data

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

- Bin channel = pressure;	Averaging interval = 1.000;  Minimum bin value =   .000

- Average value will be used. Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined and no significant differences were found from climatology.

##### Final CTD File Steps (REMOVE and HEADEDIT)

For all casts REMOVE was run to remove the following channels: Scan_Number, Temperature:Secondary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T1:C1, Flag and Prediction_Flag.

- Seapoint Fluorometer removed for all casts after cast 1

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add comments about processing.

The Standards Check routine and Header Check were run; no problems were found.

Profile and T-S plots were examined and look ok.

##### Final Bottle Files

- SORT was run to arrange casts in pressure order.

- For all casts REMOVE was run to remove the following channels:

  - Scan_Number, Temperature:Secondary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T1:C1,and Flag

- A second SBE DO channel with mass units was added for both the CTD DO and titrated DO and REORDER was run to get the pairs of DO channels together.

- EDIT HEADERS was run to fix formats and channel names and to add comments about analyses and CTD processing.

- Standards check was run. No problems were found.

- The track plot looks fine.

- Plots of each file were examined and no problems were found.

- A cross-reference listing and header check were produced for the CHE files.

##### Thermosalinograph Data

An IOS TSG45 was used for this cruise and separate flow to a fluorometer in the lab.

- TSG data came in four text files and contained extra fields with data from other instruments.

- A python script was used to simplify the data to only the relevant .Two-line Headers were added with variable names and units. Data were rearranged to separate date and time, convert fluorescence in volts to ug/L (using the Wetlabs chlorophyll characterization sheet), add a pressure column with all values set to 4.5. All NAN values were replaced with -99. A “File Name” column was then added with format “2024-002-2024MMDD”.  This will be used as a file break column to create 1 file per day. These files were saved in CSV format.

- The .csv was converted into IOSShell format with a file per day.

- CLEAN was run to reset the number of records, min and max values, set the start and end times, and latitude and longitude limits.

- ADD TIME CHANNEL was used to add Julian dates (Offset from Time Zero – i.e. Day of Year). A record number was also added to enable averaging (for use in comparison to CTD files).

- Time zero was set to 31 December 2023 0:00:00. (Note that this step leads to problems plotting until REORDER is run.)

- DERIVED QUANTITIES was run to derive salinity using the lab temperature. (Sigma-T was derived later.)

- REORDER was run to move the Julian date to after the Time/Date channels and to put salinity and fluorescence after the lab temperature. Also the record # was moved to the end. Pressure was dropped.

a.) Plots

- A track plot was produced and looked fine; it was added to the end of this report.

- Time-series plots were produced and there are some large distinct spikes in salinity that can be removed using a graphical editor.

- A plot of all differences (Lab Temp – Intake Temp) through the whole record shows little variability with an average of 0.34°C which looks reasonable.

- The main flow rate was zero due to an ongoing issue with the flow metre, however reasonable flow rates were recorded (1.4-1.6 m/s) for the fluorescence sensor. However TSG data appeared good through out and matched well with the CTD data.

b.) Comparison of TSG data to CTD data

- The CTD files were thinned to reduce the files to a single point from the downcast at or within 0.5db of 4.5db. These were exported to a spreadsheet.

- The TSG files were averaged over 5 records (~30 seconds). Records were picked out for the times that match the CTD casts and added to the spreadsheet 2024-006-CTDvsTSG.xlsx. Columns were added to calculate the differences between TSG and CTD latitude and longitude. This is used to ensure the matches are good.

- During the first two casts, it was found that both flow rates were zero and significantly different than the ocean values. These TSG values will be removed.

- TSG data were also found at closest times to loops and a the TSG-Loop comparison was done.

- A comparison was made of positions for the CTD and TSG data to check for good matches. The differences in positions are expected to be small despite the averaging because the ship was stopped at these times.

The initial comparison between TSG and CTD data using all but the first two casts were (note the CTD did not have a fluorometer for most of the cruise):

|  | Tintake-Tctd | Tlab-Tctd | TSG Sal-CTD Sal | TSGfl/TSGctd |
| --- | --- | --- | --- | --- |
| average | 0.003837 | 0.340829 | -0.08303 | n/a |
| stdev | 0.00925 | 0.34 | -0.0986 | n/a |
| median | 0.165005 | 0.127636 | 0.17707 | n/a |
| min | 0.7839 | 0.8206 | 1.5642 | n/a |
| max | -1.2856 | -0.6152 | -0.2659 | n/a |

The median and average differences are reasonably close together. There were a few spikes in the data which will be removed manually.

Heating in the Loop

The average difference between the intake temperature and lab temperature was 0.34 deg C with a standard deviation of 0.06 deg C.

f.) Editing

TSG data were edited using CTDEDIT.

- At the beginning, while flow rate was zero data was removed.

- On 23 May 2024, a short chunk of data was removed where flow rate was zero

g.) CALIBRATE, REMOVE and CLEAN

- Based on the small salinity difference between the bottle and TSG CALIBRATE was deemed unnecessary.

- Sigma-T was derived earlier. First, Pressure needed to be added to enable the calculation. ADD CHANNEL was run to add pressure = 4.5db. Then DERIVE was run to derive sigma-T.

- REMOVE was run to remove channels Temperature:Difference, Record # and pressure. The flow rate channels were left in the files.

h) Preparing Final Files

- HEADER EDIT was used to change the DATA DESCRIPTION to THERMOSALINOGRAPH and add the depth of sampling to the header and to change channel names to standard names and formats and to add comments.

- The TSG sensor history was updated.

- As a final check plots were made of the cruise track and time-series and all look fine.

- A cross-reference list was prepared.

##### Loop File

- Loop file was created from CTD rosette surface bottle and loop samples; the end time of casts were used when loops were during CTD casts.

- The sampling method column indicates ROS or USW for loops.

- 6-line headers were added and the usual headers removed.

- The columns were rearranged into the usual order of the 6linehdr files.

- Derived Quantities was run to derive Sigma-T for the CHE files. Sigma-T, event # and sample # were exported to a spreadsheet, ordered on pressure and all P>7db were removed. Sample #s from the 6linehdr file were added and the file was reorganized so that sample #s matched and blanks were entered for loop samples. Sigma-T data were added to the 6linehdr.

- The file was saved as 2024-019-surface-6linehdr.csv.

- The file break column was filled with value 1 so all data to ensure only a single file is created in conversion.

- CONVERT was run to produce an IOS Header file.

- CLEAN was run to get start and stop times and to add flag 0 to empty flag cells.

- Edit Headers was used to add header comments about the cruise, CTD data and sample analysis.

| CTD# | Make | Model | Serial# |  | Used with Rosette? |  | CTD Calibration Sheet Competed? |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0443 |  | Yes |  | Yes |  |  |
| Calibration Information - 0443 |  |  |  |  |  |  |  |  |  |  |
| Sensor |  |  |  | Pre-Cruise |  |  |  | Post Cruise |  |  |
| Name |  | S/N |  | Date |  | Location |  | Date | Location |  |
| Temperature |  | 2106 |  | 16Feb2023 |  | Factory |  |  |  |  |
| Conductivity |  | 2754 |  | 18Jan2023 |  | Factory |  |  |  |  |
| Secondary Temp. |  | 5130 |  | 23Feb2023 |  | Factory |  |  |  |  |
| Secondary Cond. |  | 1766 |  | 18Jan2023 |  | Factory |  |  |  |  |
| Transmissometer |  | 1185DR |  | 11Jan2023 |  | Factory |  |  |  |  |
| Transmissometer |  | 1883DG |  | 11Jan2023 |  | Factory |  |  |  |  |
| SBE 43 DO sensor |  | 4372 |  | 10Feb2023 |  | Factory |  |  |  |  |
| PAR sensor |  | 70613 |  | 24Feb2021 |  | Factory |  |  |  |  |
| SeaPoint Fluor. |  | 4186 |  | May 2023 |  | Factory |  |  |  |  |
| WetLabs Fluor. |  | 2216 |  | 27Nov2018 |  | Factory |  |  |  |  |
| Pressure Sensor |  | 0443 |  | 22Feb2023 |  | Factory |  |  |  |  |
| Valeport Altimeter |  | 76341 |  | 23Nov2021 |  | Factory |  |  |  |  |

# *SeaPoint Fluorometer dummied off before cast 10

# TSG Make/Model/Serial#:	SEABIRD/45/0620

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature-SBE45 | 45-0620 | 12Jan22 | Factory |  |  |
| Conductivity-SBE45 | 45-0789 | 2Feb24 | Factory |  |  |
| Wetlabs WETStar Fluor. For depths deeper than, and including, 125 dbar, we would wait 30 seconds before closing a bottle. For depths shallower than, and including, 100 dbar, we would wait 60 seconds before closing a bottle. | 1656 | 12Mar2021 | Factory |  |  |
| Temperature-SBE38 | 603 | 4Jan 2024 | Factory |  |  |