REVISION NOTICE TABLE

	DATE
	DESCRIPTION OF REVISION

	
	


PROCESSING NOTES
Cruise: 2024-024




Agency: IOS, Ocean Sciences Division, Sidney, BC; Freshwater Institute, Central & Arctic Region, Winnipeg, MB.
Location:  Beaufort Sea 

Project: Canadian Beaufort Sea Marine Ecosystem Assessment
Chief Scientist: Majewski A.
Platform: Frosti
Date: 1 August 2024 – 30 August 2024
Processed by: Germaine Gatien

Date of Processing: 28 August 2025  -  24 September 2025
Number of original HEX files: 55 (2 test casts & 1 upcast only) 
Number of original CTD files: 55

Number of CTD casts processed:  51

 

Number of bottle casts:   38 (1 test & 10 eDNA)
Number of bottle casts processed: 27
INSTRUMENT SUMMARY
SeaBird Model SBE 911+ CTD (Arctic #1189 with pressure sensor 130015) was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#SCST1666DR), a SBE 43 DO sensor (#1489), a SeaPoint Fluorometer (#SCF2859), a Biospherical PAR sensor (#70501), a Surface PAR (#20279), a SeaPoint turbidity meter (#11074), an Optode Dissolved Oxygen sensor (16) and an altimeter (#85484). 
The data acquisition program was SeaSave V 7.26.7.107. 

SUMMARY OF QUALITY AND CONCERNS
The same configuration file used for the whole cruise. There was some confusion in checking the parameters since there was a note in the OSDCommon CONFIG folder indicating that the pre-cruise data for conductivity sensor #4327 were missing. If fact, they were not missing and a note in Arctic calibration spreadsheet did establish that the calibration data from August 31, 2023 were from a factory calibration. The confusion likely arose because the primary sensor originally on the CTD had failed a test and was replaced just before the cruise.
It would be good to place the latest spreadsheet of sensor calibrations with the raw data in OSDCommon/Raw when Arctic cruises are being processed by the Data Products group. Personnel are often out of reach when processing is done, and navigating the Arctic drive is difficult. The calibration parameters are currently in folder:   \Arctic\SHARE\UTILITIES\Instruments. 
There was a bridge log with times (BE,BO and EN), water depths and comments. This was very helpful since there were some errors in the rosette log sheets. The log did not include a list of sensors or details such as the name of the chief scientist or project name.
A list was provided of the first scan number that should be included in downcast files for each event. This made removal of the soak data easy.
Water depths were not in the headers but the information was available from the digital log, so they were added to all files.
The deployment method varied somewhat through the cruise. For the first 2 weeks there were 5m soaks and the CTD was then returned to the surface and the full cast run. For the rest of the cruise some soaks were just under the surface and some were deeper with a return to the surface unless conditions were bad. 
At the time of processing only salinity and carbon sample data were available. 10 casts had only eDNA sampling; CHE files were not needed by the eDNA analyst, so they were not prepared for those events. CTD files were prepared for those events.

There was no dissolved oxygen sampling, but surface saturation values and post-cruise calibrations suggest that calibration was within 1%. 

Often 2 bottles were fired at the chlorophyll maximum but only 1 bottle was sampled.
One salinity sample was mislabelled and missed until late in processing. A CHE file was prepared but the comparison with CTD data was not rerun. 
Extra dissolved oxygen sensors were used during the cruise but that channel is not included in the files to be archived. Special files including that channel were prepared and delivered to Sarah Zimmermann.
PROCESSING SUMMARY

1 Seasave
This step was completed at sea; the raw data files have extension HEX.

2 Preliminary Steps

The same calibration control file was used throughout the cruise. A few problems were encountered in confirming that all parameters were correct:
· The configuration file used at sea contained pre-cruise calibrations, but a note with the raw data said that the pre-cruise calibration for the primary conductivity sensor was missing. In fact, it was available and was used in acquisition. Due to a problem noted previously, the factory check was done many months earlier than for the other T and C sensors. Getting confirmation about this was difficult due to personnel being out of touch, but a spreadsheet was found with clear information. It would be helpful if the calibration spreadsheet for sensors used in the arctic is placed with the Raw data in OSDCommon/Raw for any arctic cruise. Finding that spreadsheet on the Arctic group drive is difficult and future processors may not even know that it exists.
· Parameters from the file were checked and all were from the most recent factory calibrations except for the PAR sensor. The PAR parameters match those for sensor #70507 which was used on the Frosti in 2023 but the serial number entered is 4601. A choice of which numbers to use was unclear. A first attempt used the correct parameters for 4601 but that produced values higher than ever seen before and about 4 times the SPAR values. So the sensor # was fixed in the xmlcon file.

· The configuration was saved as 2024-024-ctd.xmlcon. 
Post-cruise calibrations were available; a post-cruise calibration control file was prepared with those parameters and saved as 2024-024-post.xmlcon.

The rosette log sheets were obtained as well as the salinity analysis spreadsheet. Mid-way through processing carbon data were provided.
The cruise summary sheet was completed.
There was no history available for the pressure, temperature, conductivity and DO sensors since their last factory calibration. This is believed to be the first use since factory calibration.
3 Conversion of Full Files from Raw Data

The deployment method varied. For most casts there was a soak at about 5m, with a few at about 2m and at least one at 10m. Usually there was a return to the surface before beginning the full cast with the exception of events #216, 218, 226, 232, 243, 246, 249, 255, 264, 274, 318 and 334. 
A list was provided of the first scan that should be included in the downcast files. This was extremely helpful.
All hex files were converted using 2023-024-ctd.xmlcon to create CNV files.

A few casts were examined and all expected channels are present. A few features were noted:

· A few small spikes were seen in the secondary conductivity.

· The temperature and conductivity channels are reasonably close during downcasts, though there was considerable noise in both channels in the shallower casts late in the cruise.
· Comparison of an early and late cast showed little change in the differences between channels.
· All profiles look reasonable but the SPAR data have many spikes that were not removed by running WILDEDIT.
· Altimetry looks good.

· The PAR signal was very noisy at the surface, but had the usual profile.
4 BOTTLE FILE PREPARATION 
The ROS files were created using file 2024-024-ctd.xmlcon. 
File #2024-024-0112.ROS was renamed as 2024-024-0111.ROS as it is the upcast corresponding to downcast 2024-024-0011.IOS.

The files were converted to IOS format. 

CLEAN was run to create BOT files. 
Temperature and salinity were plotted for all BOT files. The only problem noted was in event #216. CTDEDIT was used to clean 3 points in the primary salinity at the bottom of the cast. The output file was copied to *.BOT.
A preliminary header check turned up no problems and the track plot looked fine.
The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. 
There were many bottles that were not sampled. Sometimes they were assigned a sample number and sometimes not. So care was needed in adding sample numbers to the addsamp file. Pad values were entered where no sample data are available to add to the files. 
10 casts had eDNA sampling only and the researchers do not need CHE files for those casts.
The addsamp.csv file was converted to CST files, which will form the framework for the bottle files. 
SAM files were created using the Add Sample Number routine and those files were then bin-averaged. Output was SAMAVG.
Next, the salinity analysis spreadsheet was examined to see what comments should be included in the header file. These were used to create file 2024-024-bot-hdr.txt. (Later, carbon data became available and comments were added to the header text file.)
SALINITY

Salinity analysis was obtained in spreadsheet QF 2024-024 SAL_2024-11-13.xlsx. 
The sheet with final data for merging with CTD data was saved as 2023-024SAL.csv. 
A line was added for sample #308 so a flag and comment could be added to explain why it was planned but not taken.
The analysis was 2 to 3 months after collection.
The file was converted into SAL files.
The SAL files were merged with CST files with extensions MRG1. 

DIC-Alkalinity

These data arrived late in processing, so a return was made to this stage. 
The analyst’s file was not of the usual format, missing analysis comments. The comments from a previous cruise from the same analyst were used in the header. 
The file was converted into individual DIC files which were merged with the MRG1 files and named as MRGDIC. 

The MRGDIC files were put through CLEAN to reduce the headers to File and Comment sections only.
These files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so the MRGCLN1 files were reordered on Bottle_Number and saved as *. MRGCLN1s. 

The MRGCLN1s files were then merged with SAMAVG files using merge channel Bottle_Number with output MRG. 

There are some records with sample #-99 since there were no samples taken. This generally happened when 2 bottles were fired at the CHL maximum but only one Niskin sampled. Both records will be kept in the files as there may be some value in seeing variability in the CTD data.
The MRG files were put through CLEAN to produce MRGCLN2 files; 0s were entered into any empty flag channels 
5 Compare  
Salinity  

Compare was run with pressure as reference channel. There was a lot of salinity sampling. While most samples were from the top 50m where gradients are high, there were 129 samples below 100m, and when outliers among those samples were removed based on standard deviation of the CTD salinity in the 10s-window being >0.001psu or differences >0.05psu for the primary and >0.03 for the secondary there were 66 points of comparison. .The deepest was 635m. 
The primary salinity was found to be high by 0.0011psu and the secondary low by 0.0002psu with standard deviations of 0.0055psu and 0.0057psu, respectively. 

There are several competing effects to be considered: pressure dependence, time dependence, local gradients, flushing of Niskin bottles and time awaiting analysis of samples. So the bottles were grouped in a variety of ways to try to understand which reflect calibration best. 
Flushing errors tend to have a different sign for bottles fired at the bottom as samples in the Niskin bottle will come from above the CTD firing level. The opposite happens on upcasts, but errors tend to be higher near the sea bed. When only bottom bottles were included with the same method for excluding outliers, the primary was high by 0.0022psu and the secondary was high by 0.0004psu. So using bottom bottles changes the differences slightly. We could exclude them but then we lose many samples from depths where gradients tend to be lowest. The CTD salinity is likely reading a little lower than the results suggest when bottom bottles are included, but the difference is no more than 0.001psu. Both sensors are likely reading within 0.001psu.

Plots against time (using file pair numbers) shows a slight decrease over time of about 0.001psu. There are not enough data to read too much into this result, but possibly it indicates better flushing in the later parts of the cruise. The deeper casts are fairly well distributed with time.

The analysis was somewhat delayed which can raise the salinity level in samples, thus making the CTD look like it is reading lower than it really is. If that were a factor we would expect the early samples to have suffered the most so that the CTD would read lower early and higher later. This does not appear to be a significant factor in this data set.
For details see file 2022-024-sal-comp1.xls.

6 WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).  

Parameters used were: 
Pass 1    Std Dev = 2 
Pass 2    Std Dev = 5 
Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

7 ALIGN DO

Based on results from previous years using this CTD in the same area a setting of +3s was selected to advance the dissolved oxygen channel. Plots examined afterwards showed this was a good choice to bring the offset between downcast and upcast DO into alignment with that of the temperature channels. 
ALIGNCTD was run on all casts using a setting of +3s to the dissolved oxygen channel.
8 CELLTM

The usual tests for the best settings for this routine are difficult to interpret because of the noisy upcast data, so the default settings was used.
CELLTM was run using (α = 0.0245, β = 9.5) for the primary and secondary conductivity.
9 DERIVE and Channel Comparisons
Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

Tests were run to ensure the previous steps worked properly. 

WILDEDIT- Few spikes present, but WILDEDIT did remove some spikes in conductivity.
ALIGNCTD- Offsets between downcast and upcast DO profiles look similar to temperature profiles. 

T-S plots show that the CELLTM step worked well near the surface, but in deeper water where temperature was slowly increasing with depth, it did not work well. Tests were run using different parameters for CELLTM. No other choice of parameters was found that worked better at the surface. At depth there are better choices, but those don’t work well in higher gradients; any error caused in the deep values were insignificant (<0.001psu).  See Celltm_Study.xlsx.
DERIVE was run a second time on 3 of the deeper casts to find the differences between the channel pairs. 

	Cast #
	Press
	T1-T0 
	C1-C0
	S1-S0
	Descent Rate

	2024-024-0091
	300
	+0.0014
	-0.00009
	-0.0028
	Mod, Steady

	“
	430
	+0.0012
	-0.00010
	-0.0028
	Mod, Steady

	2024-024-0252
	300
	+0.0015
	-0.00011
	-0.0029
	Low, Noisy

	“
	430
	+0.0011
	-0.00010
	-0.0030
	Low, Noisy

	“
	630
	+0.0011
	-0.00010
	-0.0029
	V.Low, V.Noisy

	2024-024-0297
	300
	+0.0013
	-0.00010
	-0.0026
	Mod, Steady

	“
	430
	+0.0013
	-0.00008
	-0.0026
	Mod, Steady


The salinity differences are larger than found in the comparison against bottles, when the secondary was lower than the secondary by 0.0013psu, on average. The bottle comparison includes data from 100db down and there is some pressure dependence of opposite signs in the two channels and more in the primary salinity than in the secondary. When only deep bottles are examined the differences are closer to 0.0025psu. The primary pressure dependence is more severe than that in the secondary, so choosing the secondary sensors for archiving looks like the best choice.
Cast #40 was converted using the post-cruise configuration file and data from around 400db were compared. The post-cruise values show that the primary temperature had risen by 0.0004 while the secondary had dropped by 0.0010. The combination of those fairly small drifts does match the T1-T0 values found in the current data set quite well given only 1 record was compared.

For conductivity both sensors had drifted down by about 0.0005 leading to a drop in salinity of 0.0065 for the primary and 0.0055psu for the secondary. The differences using the pre-cruise calibration is close to the differences found in the study above, as expected. It is smaller using the post-cruise calibration. 
The COMPARE plots against time suggest a slight drop in CTD salinity with time when compared to bottles but that is a very rough judgment. There is no obvious large drop so most of the calibration drift likely occurred after this cruise.
The post-cruise comparison suggests that the primary and secondary salinity are lower than found in the comparison with salinity samples. This can’t be explained as being due to errors in the comparison due to incomplete flushing of Niskin bottles or evaporation/adsorption of samples. Pressure dependence complicates this analysis, but it is likely that the drift suggested in the post-cruise calibrations occurred after this cruise. 
10 Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers. 
CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number. 
A file was available that indicated the scan number for the end of the soak period. This was used in running CLIP to remove the data collected before and during the soak. 

However, for many casts the CTD rose after the soak and those data are not removed. A test run of DELETE showed any data from the upcast were removed but some of the shallowest data are likely to be affected by the rise from the soak level.
11 Checking Headers

An initial cross-reference list was examined and position entries were missing for one cast. Those were added based on the log record.
A header check was done. There are obvious spiky values likely from the surface but no significant problems were noted; no negative fluorescence values were found. The minimum pressure reading was ‑0.08db.
Track plots – The cruise tracks were plotted and look ok; they were added to the end of this report. 
The surface check was run on the files before CLIP was applied. The average value was +0.47db with a minimum value of +0.26db. These measurements appear to all be “in water”. Entries in the log indicate pressures at the beginning and end of casts vary from 0.2db to 0.6db. 
All casts were put through REVERSE and Surface Check to find values at the end of casts; those all appear to be in-water values. The average value was 0.4db. 

So pressures appear to be within specifications.
12 CLIP

A list was provided of the # of scans that needed to be removed from each cast to remove the soak data; 
CLIP was run to remove the soak period for each cast. 
The bottom depth had not been entered in the headers of files. The data are available in the log and were used to create file 2024-024_hdr_mrg.csv. “Merge CSV_Files to Headers” was run to add that information to the CLIP files.
The altimeter readings from the headers of the CLIP files were exported to spreadsheets. A check was made by calculating: 
Check Value = Max Depth Sampled + Altimetry header – Water depth. 
Only 5 casts were found with the check value was >5m. The altimetry is quite noisy at the bottom and the bottom depth when the CTD was near the bottom was a little different from what was recorded at the beginning of the casts. So the larger check values are likely due to multiple causes and none are large. No changes were made to depth entries. 

13 Shift
Fluorescence

SHIFT was run on the fluorescence channel in all casts using the usual advance of +24 records. Plots made after this step were hard to interpret due to noisy fluorescence but there is clearly some improvement in matching the fluorescence offset to the temperature offset and for some casts the match in offset is excellent.
Dissolved Oxygen 

Channel Oxygen:Voltage:SBE was aligned earlier. A few casts were checked to see if further alignment is needed for the DO concentration channel; it is difficult to judge where there are bottle stops, but for casts without such stops the alignment looked good.
Conductivity
During 2023 with the same equipment the choices made were -0.8s for the primary and -0.4s for the secondary. 

Those settings looked excellent for these data as well.

SHIFT was run twice using -0.8s for the primary and -0.4s for the secondary.
14 DELETE

The following DELETE parameters were used: 

Surface Record Removal: Last Press Min
Maximum Surface Pressure (relative): 10.00
Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure 
Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warnings concerns cast #6 which is extremely shallow.
That cast was a test cast and will not be processed further.

One cast was missed in the IOS SHELL steps (#334), so they were repeated for just that cast.

The output files were copied to *.EDT.

These data were not sent to the CTD-QC program as it is not available for this region.

Header Check was rerun and shows that the major spikes have disappeared.

15 Other Comparisons

Previous experience with these sensors – 

There was no known use of the pressure, temperature, conductivity and DO sensors since their latest factory calibration. 

Historic ranges – Local climatology was not available.
Repeat Casts – There were no suitable casts to test repeatability. 
Post-Cruise Calibration – Available at the time of processing. Shows drift in conductivity that appears to have happened after this cruise.
16 DETAILED EDITING
The secondary temperature and salinity channels were edited for all casts since they looked smoother overall in T-S space than the secondary channels and had less pressure-dependence in COMPARE. 
CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and misalignment of T and C and records corrupted by shed wakes. Most casts required editing near the top and bottom of casts.
During the 2nd half of the cruise sea conditions often led to the CTD not being returned to the surface after the soak period. Such casts began at greater depth than usual.
All files were edited except #137.
The output files were copied to EDT. 
T-S plots were examined and the results looked ok; there are some small unstable features in a few casts near the surface. There was no obvious instrumental cause so they may be real. 
17 Initial Recalibration
The pressure does not need recalibration.
There does not appear to have been significant drift in salinity through the cruise and both channels look to be within ±0.002psu.
There was no Dissolved Oxygen calibration sampling, so the surface saturation steps were run early to see if there was any evidence of low DO values.
CHANGE UNITS was used to derive dissolved oxygen in mass units.

DERIVED QUANTITIES was used to calculate DO surface saturation. 

Plots were made of DO surface saturation and most values were between 98% and 102% with just a few a little lower. 

Two casts were then converted using pre-cruise and post-cruise DO calibrations from early and late parts of this cruise. They both indicate that DO calibration had drifted lower by about 1% by November. Some of that drift likely occurred after this cruise.  So values may be reading a little low, but there is insufficient evidence to justify recalibration. Any error is likely <1%.
CALIBRATE was not run. 

18 Fluorescence Processing 
A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)
19 BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure
Averaging interval = 1.000
Minimum bin value =   .000

Average value will be used.
Interpolated values are NOT used for empty bins.

20 Final CTD File Steps (REMOVE and HEADEDIT)
REMOVE was run after REORDER in order to prepare special files with all channels for use of the program lead.
REORDER was run to get the two DO channels together.

REMOVE was run to remove the following channels: Scan_Number, Temperature:Primary, Conductivity:Primary, Salinity:T0:C0, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag from all casts

HEADER EDIT was run on the REM files to add a special header and files were given output extensions CTDSPEC. These files are for the use of the Chief Scientist and contain the extra DO sensor channels. 
REMOVE was run again to remove channels Voltage:6 and Voltage:7. These files have extensions REM2.

HEADER EDIT was run on the REM2 files to fix formats and channel names and to add comments.
These files have extension CTD and are intended for the OSD data archive.
The Standards Check routine was run and no problems were found. 
The cross-reference list was produced for the CTD files and no problems were found.
A header check was run on the final files and no problems were found. 
Plots were examined on-screen and profiles and T-S plots looked fine. 
PAR is often much lower than the Reference PAR but that be related to minimum cast depths and/or the presence of ice.
21 Final Bottle Files
No recalibration was required.

The MRGCLN2 files were put through SORT to order on increasing pressure. 
REMOVE was run to remove the same channels as for the profile files plus Voltage:6 and Voltage:7.
A second SBE DO channel was added to the CTD DO with mass units and REORDER was run to get the 2 SBE DO channels together.
“Merge CSV_Files to Headers” was run to add water depth. Output MRGMRH. 
HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.
The vessel name was corrected.

The standards check suggests that the DIC and ALK data should have format F7.1, but the data were provided with 2 decimal places which is the usual standard now. So F8.2 is was added to the standards file ahead of F7.1.

A header check was run on the final files and showed that only one cast had DIC sampling but not salinity. This was checked and it turned out that one salinity sample was mislabeled as coming from cast #213. The merge process for cast #216 was repeated and a new CHE file prepared.
No further problems were found. 

A cross-reference list was produced and no problems were found.

Plots were examined and no problems were found.

The bottle data were exported to a spreadsheet and compared with rosette sheets and no problems were found.
Particulars – 
Split casts: 111/112 

eDNA casts: 18, 23, 42, 114, 137, 218. 226, 300, 318.
No return to surface after soak: 216, 218, 226, 232, 243, 246, 249, 255, 264, 274, 318, 334
6. Test cast, no sampling.
9. On upcast went back down to get 15m bottle; moderate contact with ship during deployment.
18. CHL max at 10m, so skipped bottles 5, 6, 7 eDNA.

21. Bottom time approximate; stopped ~250m and 100m on upcast to fix wire wrap.

31. No samples 23 or24 – did not actually reach sal=33.1 as planned.

40. No samples 37 or 45 (CHL max at 60m)

52. No samples 53 or 62 (CHL max at 50m); surface layer not very mixed – very scattered.

59. CTD only; Press start ~0.2, end 0.3187

61. No samples 69 or 83B or C (forgot triplicates)

67. No sampling

79. CTD only  P_start 0.2496; P_end 0.3339

84. No samples 97 or 104 or 108
107. CTD only P_start 0.39; P_end 0.449

111. New file started at bottom. 
112 Upcast of previous file. Renamed bottle file as 111.
124. P_start 0.5027; P_end_0.4848
126. wavy )2 profile below 60m

133, CTD only P_start 0.5151

144. CTD only P_start 0.5084; P_end_0.5321

148. CTD only P_start 0.6249; P_end 0.4411
189. Assigned sample numbers 189-205 but no IOS samples, only eDNA.

232. Northward alongshore current? Not strong.

243. Missed a block of sample #s (226-233)

288. 2m soak.

CRUISE SUMMARY – CTD

	CTD#
	Make
	Model
	Serial#
	Used with Rosette?
	CTD Calibration Sheet Competed?

	1
	SEABIRD
	911+
	1189
	Yes
	Yes

	Calibration Information - 1189

	Sensor
	Pre-Cruise
	Post Cruise

	Name
	S/N
	Date
	Location
	Date
	Location

	Temperature


	5830
	4Jan2024
	Factory
	14Nov2024
	Factory

	Conductivity
	2984
	31 Aug2023*
	Factory


	14Nov2024
	Factory

	Secondary Temp.


	
5831
	14Feb2024
	Factory


	14Nov2024
	Factory

	Secondary Cond.


	4339
	12Dec2023
	Factory


	18Dec2024
	Factory

	Transmissometer


	1666
	10Jun2022
	IOS
	
	

	SBE 43 DO sensor
	1489
	22Feb2024
	Factory
	13Nov2024
	Factory

	PAR
	70501
	4Apr2016
	Factory
	
	

	SeaPoint Fluor.
	2859
	21May2020
	
	
	

	Surface PAR
	20279
	4Apr2016
	Factory
	
	

	Pressure Sensor
	130015
	23Dec2013
	Factory
	
	

	Altimeter
	85484
	27Jan2023
	Factory
	
	

	Turbidity Meter
	11074
	23Mar2007
	Factory
	
	

	
	
	
	
	
	


[image: image1.png]PLOTTED: 2075/89/16 13:98:33

2024-024 Event #s

135,00 130,00 125.00 120.00 115.00
72,50 L L L 72,50
72,00 F72.00
71,50+ 71,50
5]
=
-]
i=
=
<
— 7100 F71.00
=
=
=}
Z
70,50+ 70,50
70,00 - 70.00
£9.50 = £9.50
135.08 B 1 115.00

West Longitude

START TIME: UTC 2024/08/@6 21:58:38  END TIME: o




[image: image2.png]PLOTTED: 2075/89/16 14:01:85

2024-024 Stn Name

135,00 130,00 125.00 120.00 115.00
72,50 L L L 72,50
72,00 F72.00
71,50+ 71,50
5]
=
-]
i=
=
<
— 71.00 F71.00
=
=
=}
Z
70,50+ 70,50
70,00 - 70.00
£9.50 o £9.50
135.08 B 125,00 115.00

West Longitude

START TIME: UTC 2024/08/@6 21:58:38  END TIME: o




PAGE  
1

