## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2024-026				Agency: OSD

Location: Vancouver Island Coast		Project: Integrated Pelagic Ecosystem Survey

Chief Scientist: King J.				Platform: Sir John Franklin

Dates: 	Leg 1: 24 June 2024 – 8 July 2024       Leg 2: 11 July 2024 – 22 July 2024

Processed by: Germaine Gatien			Date of Processing: June 2, 2026 – July 28, 2026

Number of original HEX files: 65		Number of processed CTD files: 64

Number of rosette casts: 65			Number of processed CHE files: 64

Number of original TSG hex files: 1		Number of processed TOB files: 0

# INSTRUMENT SUMMARY

CTD #506 was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#983DR), a SBE 43 DO sensor (#0997), a SeaPoint Fluorometer(#3949) and  a Valeport altimeter (#75321).

Seasave version 7.26.7.121 was used for acquisition.

The Deck Unit was SBE11P #1043.

The logging computer was Lenovo Thinkcentre #103.

A thermosalinograph (SeaBird 21  S/N 3363) was mounted with a Wetlabs WETStar fluorometer (#1663). The sampling interval of the TSG was 30s.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log Book and Sampling Logs contained all required information. Other information was included in a text document.

There were 2 legs for this cruise; the CTD and TSG equipment did not change.

Salinity and dissolved oxygen calibration sampling came from only 2 casts and was limited to fairly shallow sampling.

There were sample number errors in CTD files for events #47 and #115.

The primary conductivity data were bad during 2 previous cruises when this CTD was used. This continued to be the case during this cruise until the data suddenly improved. This is assumed to have been due to debris being flushed out of the system. The primary salinity continued to be quite noisy and problems with the sensor cannot be assumed to have resolved completely.

During cast #99 the rosette hit the deck; a log entry indicates no damage was found and CTD data looked ok. However, comparisons with bottles show that both channels were affected. The differences between channels differ from what the post-cruise calibration would suggest.

The secondary salinity was selected for both sections with no correction to events #1-99 and +0.0087psu for events #103-260.  This correction does mean raising the secondary salinity by 0.0057 more than the post-cruise calibration suggests, but is seems clear that something was going on beyond calibration drift. The primary salinity might be ok for the later casts, but there is insufficient evidence to support this.

Salinity values are considered less reliable than usual, so are reported with 3 decimal places only.

It was noted at sea that SeaPoint fluorescence values were very high, but comparison with extracted chlorophyll samples shows that the fluorometry compared well with chlorophyll values. The correspondence was better than we usually see from these instruments.

No estimate was made of dissolved oxygen data accuracy as there was insufficient sampling for that, but no particular problems were noted.

There was no flow meter on the thermosalinograph. Based on reports during other Franklin cruises it is believed that flow rates were likely very low and possibly variable. These differences and reports from later cruises suggest that there is a serious problem with flow in the loop. There was a 12-day gap in the TSG record that is not mentioned in the log.

All TSG variables differ significantly from CTD data. The differences in temperature and salinity are moderate for the first file but still larger and noisier than expected. For the second part of the cruise they vary from being lower than expected to much larger than expected.

The TSG data reliability is considered low, so it was not prepared for archiving.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The configuration files did not change through the cruise. No errors were found.

The Log Book and rosette log sheets were obtained.

- Nutrients, extracted chlorophyll, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

- The cruise summary sheet was completed.

- The history of use of the pressure sensor, conductivity and dissolved oxygen sensors was found.

##### BOTTLE FILE PREPARATION

ROS files were created using the xmlcon files created in acquisition. The tau correction was selected, but not the hysteresis correction since there were no very deep casts.

65 ROS files were produced but for event #260 the bottle didn’t close, so no CHE file will be prepared.

Note that the log gives event 47 for station 11151 while the CTD file and rosette log have event 46. Since event 46 was a SET cast, the bottle files were renamed as 2024-026-0047 and the event number was changed in the CHEPRIM file.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create files 2024-025-bot-hdr.txt which will be updated as needed during processing.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2024-025_OXY*.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and saved as 2024-026oxy.csv.

The only sample flagged “ALL” was for sample 78 for which no other samples were reported.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2024-025_CHL QF*.xlsx. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared and saved as 2024-026chl.csv. The csv file was then converted to individual CHL files.

SALINITY

Salinity analysis was obtained in file QF2024-025_SAL.xlsx which included a precision study. The analyses were carried out in a temperature-controlled lab 44-71 days after collection. The files were simplified and saved as 2024-026sal.csv.

NUTRIENTS  -

The nutrient data were obtained in spreadsheet QF2024-025_NUTS*.xlsx. This includes a precision

There were no event #s in the fille. Those were added.

study. The file was simplified, saved as 2024-026nuts.csv.

The CSV files and ROS files were sent put through a routine that produces CHEPRELIM files; those contain combined CTD and sample data. The output files are to be used to produce CHE files.

The output file includes an error log. This turned up 3 problems that were investigated:

- Event 46 – This file was known to be misnamed but both chemistry and ROS files had the same name, so the CHEPRELIM file simply needed to be renamed as event 47.

- Event 67 – The error reported was a difference in time between log and ROS file header. This is due to difference between UTC and System Upload time. No correction is needed.

- Event 115 – This file was misnamed but the chemistry files were correct, so no chemistry was added. The ROS file was renamed as 116 and run through the Python routine again.

Events 46 and 115 were changed to 47 and 116 in file 2024-026 Rosette Log.xlsx.

Another set of files – SAM – were created only for checking for outliers and running COMPARE to see what recalibrations are needed. They are not used in creating CHE files.

The ROS files were converted to IOS format. CLEAN was run to create BOT files.

Temperature and salinity were plotted from BOT files to check for significant outliers.  None were found.

Add Sample # was used to create SAM files.

A preliminary header check turned up an error in the cruise # but this is correct in the CHEPRELIM files. No correction is necessary for the SAM files.

All salinity values were >25psu in CTD salinity and salinity samples. NO silicate correction is needed.

##### Compare

Salinity

Compare was run with pressure as reference channel.

There were not many bottles to work with.

There were only 2 events with salinity sampling and the differences between the two were striking.

- Event 67

Primary salinity was lower than bottles by an average of 0.0294psu. When data above 100db were excluded it was low by an average of 0.0279psu and a median of 0.0291psu.

Secondary salinity was higher than bottles by an average of 0.0029psu and a median of 0psu..

The primary salinity was lower than the secondary by an average of 0.032psu.

- Event 256

Primary salinity was lower than bottles by an average of 0.0021psu. When data above 100db were excluded it was lower by an average of 0.0016psu and a median of 0.0020psu.

Secondary salinity was lower than bottles by an average of 0.0086psu and a median of 0.0087psu.

The primary salinity was higher than the secondary by an average of 0.007psu.

There are two issues that arise:

- The primary salinity is much lower than bottles for one cast and higher for the other. The primary conductivity sensor had been reported as malfunctioning on previous cruises. The differences between CTD salinity channels changed abruptly mid-way through the cruise, so this is may have been a sudden clearing in flow to the primary conductivity sensor.

- The secondary salinity is close to bottles for one cast and much lower for the other. This is not explained by the fact that the samples from event #67 waited 65 days before being analyzed and event 256 waited only 16 days. We would expect the samples that waited a long time might have suffered some increase due to evaporation or desorption of glass particles, increasing bottle values so that the CTD salinity would appear to be reading lower than it really is. Instead, it looks higher. The later cast has a lower vertical salinity gradient which should also reduce errors due to incomplete flushing.

Clearly, there was something out of line with the primary salinity during event #67. Examination of CTD profile data shows that all casts up to #99 had a large difference between the two channels. Event #103 had a lot of variability. From event #107 onwards the primary data are much closer to the secondary. Both channels are noisy at times; there are too few data to establish if there is pressure dependence.

During previous cruises this primary conductivity sensor had been found to be bad, so results up to cast #103 are what we expect. The sudden resolution of the problem is not noted in the log. But there is a note that the rosette hit the deck during event #99.

There are few bottles and many outliers. It is clear that the primary salinity is not useful for events #1-99. #103 is probably ok. After that it is not clear which channel is more useful and what recalibration should be applied. The downcast data suggest that the secondary is lower than the primary by about 0.0065psu. (See table in § 9). This is in good agreement with the results from bottles during event #256. It does not match the result during event #67.

The history of the CTD shows the secondary salinity being close to bottles during previous cruises.

There was a post-cruise calibration available for conductivity sensors; the temperature sensor was not serviced at that time. The same CTD was used subsequently for one cruise, 2024-027 but with different conductivity sensors. So the factory check should be useful.

The two casts with bottles were converted using the post-cruise conductivity calibration (December 18, 2024)  for the secondary conductivity. Comparisons were made using the same record #s at 4 depths for each cast.

| event 67 |  |  |  |  |
| --- | --- | --- | --- | --- |
| Pressure | 450 | 300 | 100 | 10 |
| Pre-cruise | 34.067 | 33.967 | 33.3063 | 31.7779 |
| Post-cruise | 34.0695 | 33.9697 | 33.3092 | 31.7814 |
| Pre-Post | -0.0025 | -0.0027 | -0.0029 | -0.0035 |
| event 256 |  |  |  |  |
| Pressure | 400 | 300 | 100 | 10 |
| Pre-cruise | 34.0088 | 33.9809 | 33.8145 | 32.7826 |
| Post-cruise | 34.0115 | 33.9837 | 33.8177 | 32.7862 |
| Pre-Post | -0.0027 | -0.0028 | -0.0032 | -0.0036 |

There is no evidence of significant calibration change through the cruise. The conductivity sensors were removed from CTD #0506 before autumn cruises 2024-027 and 2024-028, so it is unlikely they were used after this cruise.

There is a little pressure dependence so a linear correction may be appropriate. The fit found was:

During event #67 the secondary salinity is very close to bottles, but during #256 it is lower by an average of 0.0086psu. If the calibration didn’t change, then what could cause the difference?

A study of differences between the two salinity channels during downcasts shows little change from cast 148 to the end of the cruise. Casts from 103 to 146 were studied to see if there was a steady change, but the data are very noisy and differences clearly pressure dependent. So alignment may be a factor. During upcast bottle stops (mostly at 10m) the differences do look similar to the later casts.

So the comparison with post-cruise calibrations was run again to see how both channels compared. For this run an average of differences during the 10s window around bottle firings was calculated.

| event 265 | primary |  |  |  |
| --- | --- | --- | --- | --- |
| Pressure | 400 | 250 | 100 | 10 |
| Pre-cruise | 34.0154 | 33.9558 | 33.8225 | 32.79 |
| Post-cruise | 34.0156 | 33.9560 | 33.8227 | 32.7902 |
| Pre-Post | -0.0002 | -0.0002 | -0.0002 | -0.0002 |
| event 265 | secondary |  |  |  |
| Pressure | 400 | 250 | 100 | 10 |
| Pre-cruise | 34.0087 | 33.9487 | 33.8154 | 32.7810 |
| Post-cruise | 34.0115 | 33.9516 | 33.8186 | 32.7846 |
| Pre-Post | -0.0028 | -0.0029 | -0.0032 | -0.0036 |

The primary appears to have changed very little while the secondary is low by between 0.0028psu and 0.0036psu with a little pressure dependence.

The post-cruise calibration seems like the most reliable evidence; some change could happen between the cruise and factory check, but that would generally involve an increase in calibration drift, not a decrease.

So why is the difference from bottles larger than the post-cruise calibration suggests?  And why is the difference between channels larger than the post-cruise calibration suggests?

- The post-cruise calibration assumes good flow to the sensor and that may not have been the case. That would not seem to be an issue during a bottle stop.

- The differences between the 2 CTD channels are consistent at ~0.007psu during the bottle stops for event #256. A quick check of many casts suggests a difference of roughly 0.0055psu between channels during downcasts in fairly low gradients. The secondary are low by ~0.003psu due to calibration drift. So the difference of about 0.0025psu to 0.0035psu is due to something else.

- Alignment of the 2 conductivity channels could be a factor during downcasts, but should not be during bottle stops.

- The bottles for event #256 were analyzed fairly quickly so values should be reliable.

- Differences between salinity channels were examined for casts between #67 and #256 and the differences change dramatically between events #99 and #0103. The rosette hit the edge of the deck before event #99. The change follows that cast so it seems likely that something in the configuration was adjusted slightly as the science crew examined the rosette looking for damage. They found none and thought the data looked ok. But they may have accidentally made a change that affected flow.

- There is a possibility that the calibration drift of 0.003psu started due to the accident. The calibration of the secondary sensor may have been better before.

The wisest step would seem to be to use the secondary conductivity, but with no recalibration for the salinity up to event #99 and apply a correction based on bottles from event #103 onwards. Given this uncertainty it will be wise to only quote salinity to 3 decimal points.

Salinity flags were checked and the only doubtful value came from a sample that did not seem out of line.

For full details for the COMPARE run see file 2024-026-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

During this cruise there were only 2 casts with DO sampling. It is known that there were problems with the primary conductivity early in the cruise and the DO sensor is usually mounted on the primary pump. The problem with the conductivity appears to be in the sensor/plumbing rather than a pump issue; temperature data showed no problem so the DO sensor was likely not affected.

DO #997  was used for 3 cruises since it was last serviced with sampling for 2 of them, with few data from the 2nd of those. The offset in previous cruises was 0.032 and that is not expected to vary between cruises. It looks like a reasonable choice for these data.

An initial fit using data from the 2 casts and excluding 3 outliers produced the following fit:

CTD DO Corrected = CTD DO *1.042 + 0.032

The two casts were then plotted separately and produced the following fits excluding the same outliers:

CTD DO Corrected = CTD DO *1.040 + 0.032	(Events #67)  	 (14 samples in fit)

CTD DO Corrected = CTD DO *1.045 + 0.032	(Events #256) 	 (9 samples in fit)

Due to the differences noted in the salinity comparison between the two events, it is possible that DO was affected, but the results are reasonably close, so the combined fit looks like a reasonable choice for recalibration.

Plots were made of SBE DO and Titrated DO values versus salinity to check for outliers. None were identified, and the SBE DO are consistently lower than the Titrated values.

Fluorescence –

COMPARE was run with extracted chlorophyll and CTD Fluorescence from a SeaPoint fluorometer using pressure as the reference variable.

There are a lot of near-surface samples where local gradients vary and the effects of incomplete flushing lead to high variability in the comparison. The SeaPoint Fluorometer values are about 85% of the extracted CHL results but there is great variability.

The following display shows that fluorescence values tended to be too high at low CHL but closer than usual when CHL is between 10 and 25ug/L. That is typical for these fluorometers though the scatter is higher than usual.

Because of the problems noted in salinity between events 99 and 103, the CHL data were compared. Unfortunately, the range of CHL in the two sections was so different that no reasonable comparison could be seen, but there was no obvious effect on fluorescence.

For more detail see file 2024-026-chl-comp.xls.

##### Conversion of Full Files from Raw Data

Hex files were converted using matching configuration files.

All expected channels were found and profiles look reasonable.

Descent rate, oxygen concentration and SBE salinity will be derived later.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, depth, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

ALIGNCTD was run on all casts to advance the oxygen voltage by +2.5s, a setting which has worked well in the past for this type of sensor. Plots were examined and alignment does look improved although the upcast data are very noisy so comparing different settings is difficult.

##### CELLTM

The noise in the upcast data makes tests for the best parameters for this routine very difficult to interpret. In the past when upcast data were not so noisy, the default setting of (α = 0.0245, β=9.5) was generally found to be the best choice.

CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity, dissolved oxygen concentration and descent rate.

Checks were made that each step worked properly:

- Spikes in temperature and conductivity channels were smoothed by the WFILTER step. There were more spikes in the primary conductivity than in the secondary before this step.

- The dissolved oxygen alignment is definitely improved and downcast features are in better agreement, so ALIGNCTD appears to  have worked well.

- CELLTM appears to have worked as expected though noisy upcast data makes the judgment difficult. But in general it has the effect of lowering salinity where temperature is decreasing (cell warms water raising conductivity but temperature sensor reading is correct, so salinity is overestimated). Conversely, it raises salinity where temperature is increasing (conductivity too low). This brings downcast and upcast traces closer together, so T-S downcast and upcast plots are in better agreement after this step.

DERIVE was run a second time on the 2 deepest casts to find the differences between the pairs of temperature, conductivity and salinity channels. Results from a few casts from 2024-002 and 2024-032 are included for comparison. The primary conductivity and salinity channels were found to be bad during cruises 2024-032 and 2024-025. For this cruise the differences suggest they were also bad at event #67 but look good for #0256.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2024-002-0050 | 1000 | -0.0003 | -0.00006 | -0.0006 | High, V Noisy |
| 2024-002-0074 | 1000 | +0.0004 | -0.00015 | -0.0023 | High, X Noisy |
| 2024-002-0090 | 1000 | +0.0003 | -0.00018 | -0.0024 | High, X Noisy |
| 2024-032-0058 | 700 | 0 | +0.00278 | +0.0283 | High, Steady |
| 2024-032-0067 | 560 | -0.0001 | +0.00272 | +0.0280 | High, Steady |
| 2024-032-0084 | 650 | 0 | +0.00260 | +0.0265 | High, Steady |
| 2024-025-0149 | 500 | -0.0001 | +0.00285 | +0.0292 | High, V.Steady |
| 2024-026-0067 | 450 | -0.0002VN | +0.00267 | +0.0302 | High, Noisy |
| 2024-026-0256 | 390 | -0.0003 | -0.0006 | -0.0065 | High, Moderate |

There is no note in the log of any change to the CTD set-up, so it is assumed that some debris was flushing out of the system. Based on the 2024-032 and 2024-025 experience, we can assume that the change concerns the primary system, so the secondary T & C will be selected for all casts. However, the primary channels may be useful for later casts if problems are discovered in the secondary.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert Sea-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

Surface check found the average minimum pressure of 2.6db which looks reasonable.

Most casts were not well mixed at the surface.

There was a single report of a deck pressure of +0.7db, so pressures may be reading low. However, deck measurements do tend to vary so a single reading is insufficient evidence to justify recalibration.

The cross-reference check was run.

There was one error in station name – Event 201 was changed from station 3325 to 3357 in both *.IOS and *.CLN files as well as the CHEPRELIM file.

The report also had a -99 for water depth. A list of water depths showed 2 entries were blank in the SeaBird headers which was changed to -99 in conversion to IOS Headers; the data were entered based on the log sounder record.

The header check routine was run and no errors were found.

The altimeter and water depth readings from the headers of the CLN files were exported to a spreadsheet. A “Check Value” was calculated as:

Check Value =  absolute value (Max. Depth Sampled + Altimetry header – Header depth entry)

Where that number was > 6m checks were made to see if the log entry for water depth differed from the header entry for the file and whether the altimetry signal at the bottom provided a good header value.

- In all cases the altimetry signal looked clear at the bottom.

- For file #63 the water depth entry matched the one from the previous cast, not #63, so that was replaced in the header. The check value was then very small.

- For file #67 there were multiple water depth entries suggesting either noisy sounder data or significant drifting. The value recorded at the bottom of the cast is where the information is most useful and produced a low check value. So that value was entered in the header.

- For the other cases (148, 156, 158) there was no problem found in altimetry or header entries and the check values were all <10m so no changes were made to file headers.

The changes to #63 and #67 were made in the headers of the IOS and CLN files and the CHEPRELIM files. Details may be found in file 2024-026-altimeter-ctd.xlsx.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel for all casts using the usual advance of +24 records. Plots show that the fluorescence offset is reasonably close to the temperature offset after this step.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it did. No further alignment is needed for the DO concentration channel.

Conductivity

The primary salinity data were examined from the latter part of the cast and while the data are much closer to the secondary than before cast #119, they still much noisier than the secondary. Unless significant problems are found later in the secondary data, only the secondary will be processed from this point onwards.

The last time CTD #0506 was used a setting of -0.6 records was used for the secondary conductivity. Tests suggest a lower shift worked better and -0.3 seemed best.

SHIFT was run on all casts using -0.3 records and salinity was recalculated for the secondary channel.

##### DELETE

DELETE was run on the SHFC1 files.

The following DELETE parameters were used for casts 1-39:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 9 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

T-S plots were made and show the data need some editing, especially at the top of casts, but they are reasonably smooth for the most part.

The DEL files were submitted to the CTD-QC File Processor routine to add editing predictions.

The DELPRED files returned were copied to *.EDTFINAL. After graphical editing done the EDT files will be copied to EDTFINAL a well to ensure full set of files even if some were not edited.

Before sending to QC product – copy to EDT. Those files can be deleted once sent.

##### Other Comparisons

Experience with these sensors since last factory service –

2024-002 – Used for most of the cruise. Primary Salinity estimated high by 0.002 and secondary high by 0.0003. Dissolved oxygen was recalculated with scale/offset 1.0232/+ 0.032.

2024-032 – Used for only a few casts and malfunctioned Results from 2024-002 used for DO calibration.

2024-025 – Used for a few casts only – problems with primary conductivity. No corrections applied to salinity or pressure.

Historic ranges –Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed.

All temperature data fell within the climatology except for cast #156 which had temperatures slightly above the maximum at 20m in water of depth about 100m.

The only excursions in salinity were to high values between 20 and 45m at a few sites off the West Coast near the 100m isobath.

These excursions do appear to be indicative of calibration problems.

Post-Cruise Calibration – Pressure was recalibrated in January 2025 but after repairs.

Repeat Casts –There were no repeat casts and nearby casts were too shallow and in too active environments to provide reasonable tests of repeatability.

##### DETAILED EDITING

The *.DEL files were zipped (2024-025-del.zip) and submitted to the CTD-QC-File Processor. DELPRED files were returned.

The secondary channels were chosen for editing. This is the only choice for the early casts. After event #99 this is not obvious, but given the early problems with the primary conductivity and more noise in the primary salinity channel, this seems best.

In any case, the problems noted will limit reliability of the salinity data.

The IOS SHELL version of CTDEDIT was not available at the time of processing due to technical issues. A prototype python editing program was used. Input files were DELPRED.

Because there was no embedded T-S plot, there is less confidence about which features were unstable.

CTDEDIT was used to fine-tune editing predictions from CTD_QC File Processor output. There tend to be too many data identified for removal in the AI prediction, so most editing involved saving records that the prediction would have removed.

The output files from the editor were named EDT.

They contain some records with only pad values so could not be plotted. Data were put through CLEAN to remove those records (output EDTCLN).

Then the files were metre-averaged (output EDTAVG). These are temporary files – after filtering fluorescence the binning process will be rerun.

T-S plots were examined for all casts and further editing was found necessary for 2 casts. They python editor allows the EDT files to be reopened, previously deleted data to be reinstated as needed and further editing to be applied.

The 2nd round improved the data somewhat but the usual balance between removing all data that appears doubtful leaves inadequate coverage of the profile. The unstable features that remain are near the surface and could reflect real conditions.

The EDT files were zipped and sent to the CTD-QC file processor, but since the style of editing is not our usual method, Lee Croft was warned that they may not be good for training the model.

The file names in EDTCLN files for events 46 and 115 were changed to 47 and 116. The event #s in the file headers were also updated.

##### Corrections to Pressure, Salinity and Dissolved Oxygen Concentration

Silicate does not need correction since there was no CTD salinity or sample salinity <25psu in bottle files.

Pressure will not be recalibrated. It may be slightly low, but there is insufficient data to judge.

For salinity the COMPARE results are confusing as it is clear both channels were affected after cast #99.

The primary appeared to be reading low by 0.03psu before #99 and low by 0.002psu after #99 based on comparison with bottles and the post-cruise calibration showed it to be low by 0.0002psu, but that would be with perfect flow to the sensor, which cannot be assumed given the problems in the early part of the cruise.

The secondary appeared to be reading low by 0.0004psu before #99 and low by 0.0087psu after #99.

The differences between channels after cast #99 was about 0.0067psu while stopped and about 0.0065 while underway in low gradients.

The safest approach for recalibration appears to be using secondary salinity for both sections with no correction to events #1-99 and +0.0087psu for events #103-260.  This correction does mean raising the secondary salinity by 0.0057 more than the post-cruise calibration suggests, but is seems clear that something was going on beyond calibration drift.

The choice of secondary salinity for archiving looks appropriate since the primary was very clearly unreliable up to #99. After that point it may have been ok but we have little evidence about that.

A linear correction based on a study of the effect at various depths was selected for recalibrations:

Salinity:T1:C1 Corrected = Salinity:T1:C1 +0.0087

The correction for dissolved oxygen is based on the COMPARE result.

CTD DO Corrected = CTD DO *1.042 + 0.032

COMPARE was not rerun on salinity since the fit was not based on the original fits. Checks were made that the corrections were done properly.

COMPARE was re-run on dissolved oxygen after recalibration. When the same points were included as in the pre-calibration fits, the SBE DO was low by an average of 0.003mL/L. When forced through the origin the fit of differences against SBE DO was very flat.

For further details see file 2024-025-dox-comp2.xls.

CALIBRATE was then run on the EDT files

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined and no problems were noted.

##### Final Calibration of DO

There are insufficient bottle data to do a comparison of downcast SBE DO with upcast samples.

##### Final CTD File Steps (REMOVE and HEADEDIT)

For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Primary, Conductivity:Primary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T0:C0 and Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add comments about processing.

Only 3 decimal places are shown for CTD Salinity.

The Standards Check routine only noted the non-standard CTD salinity format.

The Header Check was run; no problems were found.

Profile and T-S plots were examined and look ok.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Values at 2 to 3m ranged between ~85% and 185%. The lowest values were near shore off the west coast at about 100m depth. The highest values were mostly in the southern part of the project area and inshore of 100m depth.

The majority of values were between 95% and 115%.

Values are reasonable for the area and season.

##### Final Bottle Files

CALIBRATE was run earlier on these files.

COMPARE was rerun earlier to ensure the calibration of DO was done properly and it was.

SORT was run to get files increasing pressure order

For all casts REMOVE was run to remove the following channels: Temperature:Primary, Conductivity:Primary and Salinity:T0:C0.

A second SBE DO channel with mass units was added for both the CTD DO and titrated DO and REORDER was run to get the pairs of DO channels together.

EDIT HEADERS was run to fix formats and channel names and to add comments about analyses and CTD processing.

Data were exported from the CHE files to file 2024-025-bottles-final.xlsx. The entries were compared with the rosette log sheets. One problem was found.

Event #201 was intended to have 1 bottle fired but it didn’t fire so there were no samples. But the rosette sheet exists so CHEPRELIM was created.  The analysts often report such intended samples but did not in this case. So a bottle file was created for event 201 but has no sampling data. ADD CHANNELS was used to add the missing samples with pad values and flag 9 and a comment. This will ensure that no one in future will question what happened.

Standards check was run. No problems were found.

The track plot looks fine.

A cross-reference listing and header check were produced for the CHE files.

##### Thermosalinograph Data

There were 2 thermosalinograph hex files. After conversion the one named 2024-026 was renamed 2024-026-0001 and 2024-026_Leg2 was renamed 2024-026-0002.

The configuration file created in acquisition was used to convert the hex file to CNV format.

There is a break between June 29 18:29 and July 11 2:26 with no TSG data..

There is little history for TSGs on the Franklin. During 2024-025 when it was last used the temperatures in the lab were unexpectedly high and the salinity low with large variability. Differences from the CTD were large with sudden changes.

There were also problems found with TSG data during subsequent Franklin cruises.

The files were converted to CNV format.

The CNV files were converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels.

A track plot was produced; the track is complicated but this is expected for this program and is similar to the plot of CTD stations occupied. The plot was added to the end of this report.

A plot of variables looks reasonable for both intake and lab temperatures and salinity. But the fluorescence looks very flat for most of the cruise. There was no flow meter, but the temperature differences can be used to establish if the flow was well established.

The difference between the two temperature channels is very noisy but in quiet sections is a little higher than -0.4C° early in the cruise and a little lower than that later in the cruise. The noise may be partly due to the time-difference between intake and arrival at the lab and there are also many parts of the track where there is great variability in surface waters. Fluctuating flow rate may also be a factor.

##### Checking Time Channel

Due to technical problems editing of CTD files was delayed, so the comparisons were done using unedited data. These data are unlikely to need extensive editing but surface and bottom data are likely less reliable than usual.

- The DELPRED files were thinned to reduce the files to a single point from the downcast within 0.5db of 5.5db. These were exported to a spreadsheet which was saved as 2024-026-tsg-ctd-loops.xlsx.  39 CTD casts overlapped with TSG records, and had data within 0.5m of 5.5m.

- The TSG ATC file was opened in EXCEL and reduced to times of CTD data in the spreadsheet. Those data were added to 2024-026-tsg-ctd-loops-comp.xls.

- To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. The median differences were 0.0000º and 0.0001º for latitude and longitude but the maxima and minima differences were high due to values from the first cast. During that cast there was a long time spent trying to deal with a leak, so the ship probably did move significantly making the match of positions poor. That cast was eliminated for further study and the largest difference amongst the other casts were 0.003º. The 2 clocks appear to be in reasonable agreement.

##### Comparison of Temperature, Salinity and Fluorescence from TSG and CTD data

Differences between TSG and CTD records:

|  | Tint-Tctd | Tlab-Tint | Tlab-Tctd |  | SALtsg-SALctd | TSG FL / CTD FL |
| --- | --- | --- | --- | --- | --- | --- |
| median | 0.0743 | -0.3698 | 0.4377 | -0.4486 |  | 0.0784 |
| average | 0.1883 | -0.3759 | 0.5921 | -0.4188 |  | 0.1064 |
| stdev | 0.4624 | 0.2771 | 0.5270 | 0.2345 |  | 0.0920 |
| max | 1.8774 | 0.3752 | 2.2491 | 0.0593 |  | 0.4696 |
| min | -0.6688 | -1.4332 | -0.4741 | -0.9032 |  | 0.0276 |

The intake temperature is much closer to the CTD temperature than when this TSG was used during cruise 2024-025. It is still a little higher than we would like, but the intake thermistor is reported to be further from the intake than on other ships so some heating is expected. However, the standard deviation is vary large.

A plot of differences versus event # shows fairly small differences in the early part of the cruise but great variability in the second section.

There is no mention of problems in the log and the operator thought it was working well. There is no mention of the 12 day gap in recording.

All variables differ significantly from the CTD. The differences in temperature and salinity are moderate for the first file but still larger and noisier than expected. For the second part of the cruise they vary from being lower than expected to much larger than expected.

Fluorescence is given in volts so calibration is needed to compare with the CTD. Since early in the cruise there appears to be no signal the value of 0.37volts was assumed to be an offset and the last calibration in 2021 suggests a slope of 13.9. Those values were used for the comparison but may need fine-tuning if these data are to be of use. An attempt using 0 offset made little difference.

The early fluorescence data does not look useful and the later data are suspect, especially given the noise in the other channels. It is likely that the flow in the loop is slow and irregular. A few results from Niskin bottles were used for comparison and the TSG fluorescence appears to be low. The loop samples did not overlap with the TSG record.

These differences and reports from later cruises suggest that there is a serious problem with flow in the loop. This might include variable rates.

Given all the problems with data from these files, the TSG data will not be archived.

##### Conclusions

The data are not suitable for the OSD Archive.

##### Preparing Final Files

Not prepared for archive.

PARTICULARS  - Includes notes from Daily Science Log and Rosette logs

CTD

Spikes expected in primary conductivity based on 2024-0025 experience but most were small and fairly deep.

1-99. Primary conductivity looks bad.

17 & 31. Took bottle for Fukishima study

46. Saved as event 46 but really 47.

63. Log depth in file 152 looks like from previous cast. Changed to 146 based on log entry.

67. Water depths in log highly variable. Altimetry looks good. Header entry looks high – changed to 146.

99. While extending boom, winch went slack and ROS hit edge of deck. Was able to recover and redeploy with CE monitoring. No apparent damage to sensors - data seems fine.

103. Primary conductivity looks better.

119. Primary conductivity looks ok –presume some debris flushed.

201. Bottle didn’t fire. No Sample

253. Deck pressure = 0.7db.

253. Salinity spikes showed up late >200m deep. Seemed approximately 50m apart and coincided with strong scattering layers

256. Salinity Spikes

260. Bottle did not fire, no samples.

TSG

Loop samples taken directly from the TSG outflow pipes.

Loop sample names non-standard – used 9 instead of 5.

2024-025 CRUISE SUMMARY – CTD

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | Yes | Yes |

| Calibration Information - 506 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2374 | 16Feb2023 | Factory |  |  |
| Conductivity | 3184 | 18Jan23 | Factory | 22Nov2024 |  |
| Secondary Temp. | 4883 | 23Feb2023 | Factory |  |  |
| Secondary Cond. | 4395 | 18Jan2023 | Factory | 18Dec2024 |  |
| Transmissometer | 983DR | 13Feb2024 | Factory |  |  |
| SBE 43 DO sensor | 0997 | 25Jan2023 | Factory | 3Dec2024 |  |
| PAR sensor | 70613 | 24Feb2021 | Factory |  |  |
| SeaPoint Fluor. | 3949 |  |  |  |  |
| Pressure Sensor | 0506 | 22Feb2023 | Factory |  |  |
| Valeport Altimeter | 75321 | 23Sept2020 | Factory |  |  |

# TSG Make/Model/Serial#:	SEABIRD/21/3363

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature-SBE21 | 3363 | 1Sept2021 | Factory | 21Nov2024 | Factory |
| Conductivity-SBE21 | 3363 | 1Sept2021 | Factory | 21Nov2024 | Factory |
| Wetlabs WETStar Fluor. For depths deeper than, and including, 125 dbar, we would wait 30 seconds before closing a bottle. For depths shallower than, and including, 100 dbar, we would wait 60 seconds before closing a bottle. | 1663 |  |  |  |  |