## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2024-027				Agency: OSD

Location: Strait of Georgia			Project: SoG Juvenile Salmon

Chief Scientist: King J.				Platform: Sir John Franklin

Dates: 	16 September 2024 – 1 October 2024

Processed by: Germaine Gatien			Date of Processing: June 11, 2026 – 6 August, 20026

Number of original HEX files: 39 (2 test casts)	Number of processed CTD files:  39

Number of rosette casts: 37			Number of processed CHE files:  37

Number of original TSG hex files: 1		Number of processed TOB files: 0

# INSTRUMENT SUMMARY

CTD #506 was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#983DR), a SBE 43 DO sensor (#0997), a SeaPoint Fluorometer(#3949) and  a Valeport altimeter (#75321).

Seasave version 7.26.7.121 was used for acquisition.

The Deck Unit was SBE11P #1043.

The logging computer was Lenovo Thinkcentre #103.

A thermosalinograph (SeaBird 21  S/N 2844) was mounted with a Wetlabs WETStar fluorometer (#1663). The sampling interval of the TSG was 30s.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log Book and Sampling Logs contained required information. Further information about problems encountered was provided in a text document.

Event #648 was a test cast run between Events #48 and #49 to study spiking n deep water. The data are from  an intermediate site and appear useful so were processed.

All events in the Strait of Georgia had a one minute soak at 10m. Archiving did not begin until the soak was completed and the CTD brought back to surface. For events #131-139 there was a soak at the surface.

Previous cruises using this CTD have encountered serious problems with conductivity sensors but those were replace prior to this cruise. The ones used during this cruise performed well in the past when they were on a different CTD. Post-cruise factory service 17 months after this cruise found small calibration drifts only. Calibration sampling for salinity was limited to 2 casts, but the comparison between CTD and samples is consistent with only slight calibration drift in both channels..

Dissolved oxygen sampling was also limited to 2 casts. The comparison of data from the CTD and titrated samples was consistent with results when used in the past. No estimate is made of data accuracy as there was insufficient sampling for that, but no particular problems were noted.

This project has some limitations in how event numbers are recorded so that loops are not recognized as events in the log. The loop samples were given event #s in order of collection. Notes included in a document describing sampling methods state that the sample numbers were assigned in a way to relate them to the last rosette sample taken. This worked for the first few samples but after that the samples numbers don’t seem to make sense. Fortunately, time stamps did enable comparison with TSG data.

Using sample numbers by using the nearest event number of any type would be an easier way to handle loop identification, but recording the time of sampling is the most reliable identification method.

The sampling notes describe attempts to increase flow rate which was believed to be much less than 1L/second. It was noted at sea that the TSG fluorescence signal dropped and a flush of green water occurred when the fluorometer was cleaned. The operator noted that the “lying down” configuration of the instrument might be an issue. Given obvious problems and poor comparisons with loop samples and CTD fluorescence, the TSG fluorescence data are not reliable.

The TSG temperature and salinity data look somewhat better, but comparisons with CTD and loop data suggest they are also affected by flow rates in the loop. TSG salinity is sometimes close to CTD salinity but often it is significantly higher or lower. The differences from loop samples towards the end of the cruise when salinity was highest are the most variable.

The TSG data are not reliable enough to archive. Similar results were obtained on other 2024 Franklin cruises using 2 different TSGs. A flow meter for this thermosalinograph and a switch to an instrument that does not need such a high flow rate will enable better use of TSG data from the Franklin.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The configuration files did not change through the cruise. No errors were found.

The Log Book and rosette log sheets were obtained.

- Nutrients, extracted chlorophyll, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

- The cruise summary sheet was completed.

- The history of use of the pressure sensor, conductivity and dissolved oxygen sensors was found.

##### BOTTLE FILE PREPARATION

ROS files were created using the xmlcon files created in acquisition. The tau correction was selected, but not the hysteresis correction since there were no casts deeper than 1000m.

File name 2024-027-0095.ros was renamed 2024-027-0096.ros.

Water depth was changed in headers for events 31, 49, 139, 145 and 151 based on entries in log. These were changed in the converted cnv files as well.

There were 39 hex files, but for event #58 there was no BL file, so there is no CHEPRELIM file. There was a surface sample taken independently during that cast. Those sample results are available in the QF files for CHL and Nutrient analyses, but there are no CTD data available from the sampling level so no bottle file was produced. There was also a deep water test cast with no sampling, named event 648.

Next, each of the analysis QF spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create files 2024-027-bot-hdr.txt which will be updated as needed during processing.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2024-027_OXY*.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and saved as 2024-027oxy.csv. There were no comments marked ALL.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll and phaeo-pigment data were obtained in file QF2024-027_CHL QF*.xlsx. The file included comments and flags and a precision study. A simplified version of the spreadsheet was prepared and saved as 2024-027chl.csv.

SALINITY

Salinity analysis was obtained in file QF2024-027_SAL.xlsx. There were no duplicate samples taken so there is no precision study. The analyses were carried out in a temperature-controlled lab 24-36 days after collection. The files were simplified and saved as 2024-027sal.csv.

NUTRIENTS  -

The nutrient data were obtained in spreadsheet QF2024-027_NUTS*.xlsx. This includes a precision study. The file was simplified, saved as 2024-027nuts.csv.

The CSV files and ROS files were submitted to a routine that produces CHEPRELIM files; those contain CTD and Sample data. The output files are to be used to produce CHE files.

The output of that routine includes an error log. This turned up a few problems that were investigated and found to be cases where there was no sampling when bottles were fired. Those data will not be processed further.

Another set of files – SAM – were created only for checking for outliers and running COMPARE to see what recalibrations are needed. They are not used in creating CHE files. The steps followed were:

- The ROS files were converted to IOS format. CLEAN was run to create BOT files.

- Temperature and salinity were plotted for all BOT files to check for significant outliers.  None were found.

- Add Sample # was used to create SAM files.

A preliminary header check turned up no obvious errors.

There is 1 salinity sample with a value <25psu. A silicate correction will be needed.

##### Compare

Salinity

Compare was run with pressure as reference channel.

There were only 2 events with salinity sampling. Both showed clear signs of incomplete flushing of Niskin bottles in the presence of high vertical salinity gradients above 200db.

Below that there are 2 major outliers:

- For cast #15 one bottle was fired at the bottom and the CTD salinity looks much higher than the bottle value. This is as expected from bottom bottles as the Niskin may contain water from above the firing level and where the local gradient is fairly high this makes a significant difference. The CTD salinity does not appear to have fully equilibrated and is continuing to increase when the bottle closed.

- For cast #130 the bottle fired at 500db has salinity much higher than the bottle. There was a salinity reversal between 500db and 540db and high variability during the bottle stop.

There were 2 less significant outliers during event #130:

- At 400db the CTD salinity looks low. The salinity was varying throughout the stop.

- At 704db the CTD salinity looks high. This was a bottle fired at the bottom and so likely to contain water from above due to incomplete flushing and possibly some shed wakes bouncing off the bottom.

None of these results suggest a problem in collection or analysis of samples, nor doe the CTD salinity look bad. This is an area where it is difficult to gather useful calibration samples due to poor flushing and few opportunities for deep sampling. No flags are justified as the samples likely represent the Niskin contents well.

Only 6 bottles remain in the comparison after outliers were removed. They suggest that the primary salinity was low by an average of 0.0013psu and the secondary low by 0.0026psu, but the standard deviations are 0.0025psu and 0.0026psu, respectively. The difference between those channels is consistent with the difference between CTD salinity channels as reported in section 9.

Both salinity channels are likely reading a little higher than those averages suggest.

No recalibration is justified.

For full details for the COMPARE run see file 2024-027-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

During this cruise there were only 2 casts with DO sampling.

DO #997  was used for 4 other cruises since it was last serviced; 3 of those included DO sampling. The results of the first 2 were similar. The manufacturer says that the offset is not expected to change with time so that was used for all cruises after the first, which was a Line P cruise with good sampling. (Offset=0.032mL/).

For this cruise the scatter is very high. There are highly variable vertical gradients, both in sign and magnitude. So it is difficult to judge what are to include in the fit.

When just 2 major outliers are excluded the fit was:

CTD DO Corrected = CTD DO *1.039 + 0.032

This is remarkable close to the result when used during 2024-026 which also included sampling from only 2 casts, but they were in more open water where flushing might be better:

CTD DO Corrected = CTD DO *1.042 + 0.032

The 2024-026 result will be used for recalibration. But there is a lot of scatter remaining.

Plots were made of SBE DO and Titrated DO values versus salinity to check for outliers. None were identified. There are a few cases of SBE DO are being higher than the Titrated values as noted in COMPARE.

Fluorescence –

COMPARE was run with extracted chlorophyll and CTD Fluorescence from a SeaPoint fluorometer using pressure as the reference variable.

Most samples are from the top 3m and only 2 from below 8m and the deepest from 25db..

Fluorescence is 85% of extracted CHL when all data are included and 87% when the 2 highest CHL values are excluded.

The usual pattern is seen of CTD fluorescence reading higher than fluorescence when CHL is low and gradually dropping until it is about half of the CHL readings for CHL>5ug/L.

For more detail see file 2024-027-fl-chl-comp1.xls.

##### Conversion of Full Files from Raw Data

Hex files were converted using matching configuration files.

All expected channels were found and profiles look reasonable.

Descent rate, oxygen concentration and SBE salinity will be derived later.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, depth, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

ALIGNCTD was run on all casts to advance the oxygen voltage by +2.5s, a setting which has worked well in the past for this type of sensor. Plots were examined and alignment does look improved although the upcast data are very noisy so comparing different settings is difficult. (Note: a further alignment equivalent to -0.5s was applied in step 12, so the net advance was +2s.)

##### CELLTM

The noise in the upcast data makes tests for the best parameters for this routine very difficult to interpret. In the past when upcast data were not so noisy, the default setting of (α = 0.0245, β=9.5) was generally found to be the best choice.

CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity, dissolved oxygen concentration and descent rate.

Checks were made that each step worked properly:

- Spikes in temperature and conductivity channels were mostly smoothed by the WILDEDIT step with the ones that remaining being mostly during upcast stops. There was one large spike in downcast secondary conductivity that will be better dealt with later in CTDEDIT as it has many points.

- The dissolved oxygen alignment is definitely improved and downcast features are in better agreement, so ALIGNCTD appears to  have worked well but more checks will be done later.

- CELLTM worked as expected bringing upcast and downcast T-S plots closer to each other.

DERIVE was run a second time on the 2 deepest casts to find the differences between the pairs of temperature, conductivity and salinity channels. Results from a few casts from 2024-002, 2024-032 and 2024-026 are included for comparison. The primary conductivity and salinity channels were found to be bad during cruises 2024-032 and 2024-025. During 2024-026 the differences suggest they were also bad at event #67 but look good for #0256.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2024-002-0050 | 1000 | -0.0003 | -0.00006 | -0.0006 | High, V Noisy |
| 2024-002-0074 | 1000 | +0.0004 | -0.00015 | -0.0023 | High, X Noisy |
| 2024-002-0090 | 1000 | +0.0003 | -0.00018 | -0.0024 | High, X Noisy |
| 2024-032-0058 | 700 | 0 | +0.00278 | +0.0283 | High, Steady |
| 2024-032-0067 | 560 | -0.0001 | +0.00272 | +0.0280 | High, Steady |
| 2024-032-0084 | 650 | 0 | +0.00260 | +0.0265 | High, Steady |
| 2024-026-0149 | 500 | -0.0001 | +0.00285 | +0.0292 | High, V.Steady |
| 2024-026-0067 | 450 | -0.0002VN | +0.00267 | +0.0302 | High, Noisy |
| 2024-026-0256 | 390 | -0.0003 | -0.0006 | -0.0065 | High, Moderate |
| 2024-027-0015 | 220 | +0.0001 | -0.00001 | -0.0005 | High, Steady |
| 2024-027-0130 | 220 | +0.0004 | -0.00007 | -0.0012 | High, Steady |
| “ | 500 | +0.0005 | -0.00007 | -0.0012 | “ |
| “ | 700 | +0.0004 | -0.00007 | -0.0012 | “ |
| 2024-027-0135 | 220 | +0.0003 | -0.00006 | -0.0015 |  |
| “ | 500 | +0.0003 | +0.00005 | -0.0009 |  |

The differences are small and in line with cruise 2024-002.

Cruises 2024-032 & 2024-026 had serious problems with primary channels for part or all of the cruise.

Cruise 2024-026 also had a sudden change in both channels mid-way through the cruise.

A test was run converting the deepest cast using post-cruise conductivity calibrations to see if there was significant drift. Data between 100m and 700m, excluding those with pad values were compared. The post-cruise primary salinity was higher than the pre-cruise by a median of  0.0020psu while the secondary post-cruise salinity was lower by a median of 0.0016psu.

|  | C0 Post- Pre | C1 Post-Pre | S0 Post-Pre | S1 Post-pre |
| --- | --- | --- | --- | --- |
| median | 0.0002 | -0.0002 | 0.0020 | -0.0016 |
| std dev | 0.0000 | 0.0000 | 0.0004 | 0.0004 |

The drift in conductivity is small and since the post-cruise service was about 17 months after this cruise some drift likely occurred after 2024-027. The difference between the 2 salinity channels is about 0.001 to 0.0015psu in the current cruise which is slightly less than the 0.002psu difference from the post-cruise service. This suggests that the COMPARE results are showing salinity channels to be reading lower than is actually the case. No recalibration is recommended as calibration drift appears to be very low.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert Sea-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

##### Checking Headers

Surface check found the average minimum pressure of 2.1db which looks reasonable.

No casts were well mixed at the surface.

The header check and cross-reference checks were run.  There is evidence of spikes in channel limit entries. There were a few station format errors. Those were changed in the CNV files and all files were put through conversion and CLEAN again.

Cast #50 had all zero values for fluorescence, as expected, since the sensor was “dummied off” to try to identify source of deep spikes. This channel will be removed from that cast later.

The altimeter and water depth readings from the headers of the CLN files were exported to a spreadsheet. A “Check Value” was calculated as:

Check Value =  absolute value (Max. Depth Sampled + Altimetry header – Header depth entry)

Where that number was > 6m checks were made to see if the log entry for water depth differed from the header entry for the file and whether the altimetry signal at the bottom provided a good header value.

- In all cases the altimetry signal looked clear at the bottom.

- For files 49, 64, 115, 118 and 139 the entries in the log differed from those in the headers and the check value was much lower when the log entries were used. The water depth header entries for CTD and BOTTLE files (*.cln and *.cheprelim) were changed to match the log.

- The remaining 5 casts checked were only slightly out of line and were in areas of narrow channels so some shoaling likely occurred. No changes were made to those files.

The deck pressures shown in the log ranged from 0.25 to 0.5db. No upcasts went through the surface.

##### Shift

Fluorescence

SHIFT was run on the SeaPoint fluorescence channel for all casts using the usual advance of +24 records. The routine did an excellent job of bringing the fluorescence offset close to the temperature offset.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it looked like the 2.5s setting was a bit too high. So SHIFT was run to apply a 12-record shift, equivalent to 0.5s. This bring the net adjustment to +2s.  The results look excellent.

Conductivity

There is no useful previous experience with the primary channel. Tests were run and a setting of -0.7 records looks best overall. The last time CTD #0506 was used a setting of -0.6 records was used for the secondary conductivity. Tests confirm that is a good choice.

SHIFT was run on all casts using -0.7 records for the primary conductivity and -0.6 records for the secondary conductivity.

##### DELETE

DELETE was run on the SHFC1 files.

The following DELETE parameters were used for casts 1-39:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 9 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 042 seconds. (taken from header)

COMMENTS ON WARNINGS: The was only 1 warning and it pertained to upcast data, so is of no concern.

T-S plots were made and show the data need some editing.

The DEL files were submitted to the CTD-QC File Processor routine to add editing predictions.

File 2024-027-0648 was rejected by the routine which deemed it an upcast. No reason could be found for this except that there was a long stop at one level 269db. All reversals were removed by DELETE, but there was one case where pressure did not increase at the 3rd decimal place, so this might be rejected by the file processor.

The DELPRED files returned were copied to *.EDT.

File 2024-027-0648.del was also copied to EDT.

At this point a study was made of whether file 2024-027-0648 was suitable for archiving. The cast was intended as a test of how the CTD performed in deeper water and there was no sampling, but the data look good. A comparison was made with nearby casts and it looks most like event #53. It is closer spatially to event #49 but that cast was much shallower and the T-S plot for #49 is much different from both #648 and #53.

The data from #648 looks good (except for spikes at the bottom) and has dissolved oxygen, salinity and temperature data that are reasonably close to #53, so it does look appropriate to archive this cast.

##### Other Comparisons

Experience with these sensors since last factory service –

The only previous use for the primary conductivity sensor was on a different CTD and with a different temperature sensor; it was low by about 0.004psu but was not selected for archiving.

The secondary conductivity sensor has been used for 3 previous cruises with very small differences from calibration samples.

The pressure sensor has been used for a number of cruises with no corrections applied.

The oxygen sensor was used most recently during 2024-026 on the same CTD; slope/offset used for recal were 1.042/0.032. It was used during part of a Line P cruise  in January 2024 with slope/offset 1.0232/0.32.

Historic ranges –Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. Where local climatology was available all temperature data fell within the limits except for a small excursion to higher temperature around 30m during event #4. There were some high salinity values at the bottom of events #4 and 27.  Salinity was slightly high during events 49, 53, 92 and 105 at various depths between 30 and 75db; those sites are all close to shore. None of the these excursions look systematic and so are not indicative of calibration problems.

Post-Cruise Calibration – The pressure, conductivity and dissolved oxygen sensors were serviced in early  2026. Changes were small.

Repeat Casts –There were no repeat casts and nearby casts were too shallow and in variable environments that are too active to provide reasonable tests of repeatability.

##### DETAILED EDITING

DELPRED files will be edited with the exception of event #648 for which the DEL file will be edited.

CTDEDIT was used to remove records that appear to be corrupted by shed wakes. Salinity was cleaned to remove spikes that appear to be due to small misalignment or instrumental noise. All files required light editing.

Notes about editing applied were added to the files.

After editing, T-S plots were examined for all casts and no further editing was found necessary; some unstable features remain but are may be real given the active mixing in this region.

All EDU files were copied to *.EDT.

The EDT files were zipped and sent to the CTD-QC file processor.

##### Corrections to Pressure, Salinity and Dissolved Oxygen Concentration

No correction is necessary to pressure . The deck pressures ranged from 0.25db to 0.52db, so within specifications and no issues have been noted in previous use.

No correction will be applied to salinity. There was insufficient information from the current bottle comparison, but it does not suggest any large error and post-cruise calibration does not indicate significant drift.

In previous use the secondary salinity was found to be close to bottles and no recalibration was applied.  During 2024-026 there was a change in performance part-way through the cruise that was not consistent with the post-cruise calibration.

There is no evidence of significant drift during this cruise or 2024-028 which followed this one.

There have been problems with the primary conductivity on a few cruises but there was a sudden improvement partway through 2024-026, but insufficient information to determine if it were doing well enough to be of archive quality.

File 2024-027-recal-sil.ccf was prepared to apply a correction to silicate because there was salinity <25psu in one bottle file. Output was *.mrgcorsil.

Next, file 2024-027-recal1.ccf was prepared to apply the following correction to *.mrgcorsil:

CTD DO Corrected = CTD DO *1.042 +0.032

That correction was also applied to the SAM files to enable a comparison.

COMPARE was run again on dissolved oxygen after recalibration. When the same points were included as in the pre-calibration fits, the SBE DO was low by an average of 0.02mL/L. The fit of differences against SBE DO was fairly flat but noisy. We do expect the SBE DO to read a little higher due to incomplete flushing of Niskins and slow response of the SBE sensor. This is not a great result, but there were no many samples and the scatter in the fit is unusually large.

For further details see file 2024-027-dox-comp2.xls.

It is noted that a post-cruise service found a damaged membrane but the time this occurred is not known.

CALIBRATE was then run on the EDT files.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined and no problems were noted.

Files 27, 41 and 648 were examined to see if channel Oxygen:Dissovled:SBE needed editing as spikes were noted around 250db for those casts. A few DO values were padded between 251db and 253db for event 27, below 250.8db for event 260. The bad values for event 648 were removed in the CTDEDIT stage since temperature and salinity also were affected by spikes.

##### Final Calibration of DO

There are insufficient bottle data to do a comparison of downcast SBE DO with upcast samples.

##### Final CTD File Steps (REMOVE and HEADEDIT)

For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Primary, Conductivity:Primary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T0:C0, Prediction_Flag and Flag.

For events #53 Fluorescence was removed.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add comments about processing.

The Standards Check routine and Header Check were run; no problems were found.

Profile and T-S plots were examined and look ok.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Values at 2 to 3m ranged between ~58% in Queen Charlotte Strait (where surface water was very well mixed) to 130% at the head of one inlet. In the central basin values were mostly between from 95% & 110%. There is no evidence of DO sensor calibration problems. This is a weak test in this region and is run only to catch major errors.

##### Final Bottle Files

CALIBRATE was run earlier as well as a second run of COMPARE.

SORT was run to arrange casts in pressure order.

For all casts REMOVE was run to remove the following channels: Temperature:Primary, Conductivity:Primary and Salinity:T0:C0.

A second SBE DO channel with mass units was added for both the CTD DO and titrated DO and REORDER was run to get the pairs of DO channels together.

EDIT HEADERS was run to fix formats and channel names and to add comments about analyses and CTD processing. The project name was changed to “Strait of Georgia Juvenile Salmon”.

Data were exported from the CHE files to file 2024-027-bottles-final.xlsx. The entries were compared with the rosette log sheets.

Profile plots were examined. No outliers were found but most casts had only 1 bottle.

Standards check was run. No problems were found.

The track plot looks fine.

A cross-reference listing and header check were produced for the CHE files.

##### Thermosalinograph Data

There was 1 thermosalinograph hex file.

File 2024-024_tsg_1.hex.was converted using 2024-024_TSG.xmlcon.

After conversion the file was named 2024-027-0001.cnv.

The CNV file was converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels.

A track plot was produced; the track is complicated but this is expected for this program and it is similar to the plot of stations occupied. The plot was added to the end of this report.

A plot of variables looks reasonable for both intake and lab temperatures and salinity.

The fluorescence looks very flat for some parts of the cruise while in others it appears to steadily increase. It dropped at about 18:50 on Sept. 28th which follows some activity to try to clean the fluorometer. The channel is likely not reliable enough to archive.

There was no flow meter and the operator noted that flow was probably not up to the level recommended for the TSG and very low for the fluorometer.

The difference between the two temperature channels is much larger than expected until about 22:00 on September 23rd. It then drops from about 1.6C° to about 0.6C° which is a reasonable value. This change corresponds to a time when the flow was being checked.

##### Checking Time Channel

- At the time these data were processed final CTD files were not available, but bin-averaged DEL files were created, named *.av1 (½m bins).

- The av1 files were thinned to reduce the files to a single point from the downcast within 0.5db of 5.5db. These were exported to a spreadsheet which was saved as 2024-027-tsg-ctd-loop-comp.xlsx.  37 CTD casts overlapped with TSG records, and had data within 0.5m of 6m.

- The TSG ATC file was opened in EXCEL and reduced to times of CTD data in the spreadsheet. Those data were added to 2024-027-tsg-ctd-bot-comp.xls.

- Data were also extracted from the ATC file to match the loop samples.

- To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. The median differences were 0.0000º for both latitude and longitude but one difference was 0.006 º. An exact match is not possible with 30s sampling and the ship may have drifted. Overall, the 2 clocks appear to be in good agreement.

##### Comparison of Temperature, Salinity and Fluorescence from TSG, CTD and Loop data

|  | Tint-Tctd | Tlab-Tint | Tlab-Tctd | TSG Sal- CTD Sal | TSG FL/CTD FL |
| --- | --- | --- | --- | --- | --- |
| max | 2.0456 | 1.7840 | 2.9483 | 0.0713 | 1.88 |
| min | -0.2072 | 0.0986 | 0.0715 | -2.9721 | -0.67 |
| median | 0.1254 | 0.9027 | 1.4346 | -0.2183 | 0.08 |
| average | 0.3250 | 0.9356 | 1.2605 | -0.5006 | 0.26 |
| stdev | 0.4953 | 0.5831 | 0.7528 | 0.7452 | 0.57 |

The intake temperature is higher than the CTD by an average of 0.32C° but the median is 0.125C°  and the standard deviation was very high at 0.50C°. The difference is higher than we usually see, but it is known that the intake thermistor for the TSG is further from the loop intake than seen from other ships and flow in the loop is known to have been variable. During 2024-025 on the same ship but a different TSG, the intake temperature was higher than the CTD by a median of 0.58C°. From other vessels’ TSG systems we do not expect so much variability. It seems likely that the flow rate variability is the cause.

The lab temperature is higher than the intake temperature by an average of 0.94C°. That is similar to the result for 2024-025 when flow rates were also an issue. This seems high especially in early autumn, but may well be accurate if the flow rate was low.

The intake temperature may be accurate but the variability is very high. There is also intake thermistor warming within the ship. This could be accounted for if the variability in the comparison with CTD temperatures were low. For now, the intake temperature cannot be used to estimate ambient ocean conditions.

Salinity is lower than the CTD by an average of 0.40psu. That is a much larger difference than expected and the standard deviation is very high at 0.74psu. Low and variable flow rates in the loop are likely factors and bubbles in the loop water may contribute as well..

The plot of salinity from TSG and CTD versus event # shows the two being in reasonable agreement some of the time, but the TSG values are significantly lower for events  17-35, 125, 131 and 135. Those casts cover several areas, so the only thing likely to create the variations are flow rates. Unfortunately, there was no flow meter so we can’t confirm this.

The fluorescence varies in a strange way, gradually climbing until it suddenly drops to near zero.  The CTD fluorometer also shows a drop late in the cruise but does not gradually increase like TSG fluorescence.

There were 8 salinity and chlorophyll loop samples.

The TSG salinity was lower than loop salinity by a median of 0.07psu, but the differences ranged from - 2.6psu to +2.6psu. The largest differences occurred late in the cruise when the salinity values were highest.

The loop chlorophyll does gradually increase up to Loop Event #38 when it increases a lot. This would be when the ship moved to Johnstone Strait area.. The fluorometer mostly reads significantly lower though as loop salinity rises it sometimes is higher. Usually we find fluorometers read high for low CHL and then drop relative to CHL as CHL increases. This unusual response suggests that the fluorometer was not responding well. Variable flow does not seem responsible as we expect the TSG and Loop were sampling the same water.

The CTD fluorescence compared well based on a comparison to Niskin sampling. Fluorescence from the TSG does not look reliable enough to archive. The data from near the end of the cruise might be ok, but with so few points of comparison this is a weak conclusion.

The salinity comparison displays high variability and that increases with time. The late increase may reflect efforts to improve the flow rate. The TSG salinity looks unreliable.

Overall there are too many problems in these data to send them to the OSD Archive.

Other Franklin cruises using a different TSG had similar problems with flow.

See 2024-027-tsg-ctd-loop-comp.xls for details.

##### Calibration History

During 2024-032 TSG salinity and fluorescence data were bad; there were problems with positions and heating in the loop was very high. A low flow rate likely accounted for this..

During 2024-026 using a different TSG similar problems were found. Data not prepared for IOS archive.

##### Conclusions

The TSG  data are not suitable for the OSD Archive.

##### Preparing Final Files

Not prepared for archive.

PARTICULARS  - Notes from Daily Science Log and Rosette logs

CTD

Deck Pressures recorded: 0.4 (event4), 0.25-0.28db (Event 76), 0.5db (event 89), 0.48-0.52db (Event 135).

1. All bottles fired for test.

4. Station 46 for Kelly Young.

15. Begin started late, see cal cast ???

27. 1 TS spike ~150db & bigger spike ~250db.

31. Shallowed slightly during cast. Use 404m for water depth.

41. Had to terminate cast at250db due to electrical spikes; still able to fire bottles.

648. Test cast post event 41 to test sensors in deep water 263db seems where sensors go crazy.

49. Shallowed after Seasave start. Use 103m for water depth.

53. Fluorometer dummied off at the 3x cable connection. Fluorometer disabled to try to avoid spikes in deep water. Did not work.

58. FL added back in, pylon dummied, no water samples. One large spike but mostly ok. Closed an independent Niskin at surface for sample #24. No CTD data available. No CHE file prepared.

79. T diff bit large?

89. FL at 5.5m was ~2 while TSG reading 0.4. Deck pressure=0.5db.

95/96. The rosette cast was mistakenly saved as 95. It should be 96.

115. Depth dropped after Seasave start?

131-139 – No 10m soak.

139. Depth increased through cast – use 432m for water depth.

145. Bottom depth variable – use 124m. Back to 10m soak.

163. depth shallowing – maybe use 151m.

648. Test cast to check on sensor behaviour in deep water – no sampling. Problems started at 263db. Close to event 49.

TSG – sample #s differ from event #s

The following are loop #s followed by sample #s

5001. Sample # 5004

5002. Sample # 5020 At anchor

5003. Sample # 5025

5004. Sample # 5030 Main line 5.30 (seconds/L) & FL line 24.23 (seconds/L)

5005. Sample # 5038

5006. Sample # 5057 Main line 5.12 (seconds/L) & FL line 24.37 (seconds/L)

5007. Sample # 5058 Main line 5.43 (seconds/L) & FL line 31.92 (seconds/L)

5008. Sample # 5062 Last loop before acquisition stopped.

TSG Field Notes (intake depth 5.5-6m)

In order to maintain sequential event numbers will GFBio, loops are not recorded as part of the sequential event numbers but the sample IDS do follow convention. For example: LOOP 5001 is the first loop sample but the sample ID is 5004 as it would have been the fourth sample ID. This is to provide the user a quick way to know when the sample was taken and coordinate it with the time/date stamp of the last rosette cast taken.

Loops are only retained inside the LOOPS worksheet.

Some cursory measurements suggest we are no where near 1L/second. I had the engineers put on 2 more pumps and we did not get an appreciable change in flow rate and all it did was create pressure on the pump system. We decided to back down to 2 pumps and periodically measure flow rate when taking the loops. I did not notice a change in the graph signal when those pumps were added. We should get a system that requires less flow and also incorporate flow meters that we can control or at least we can digitally determine the flow rate.

A flow rate with a minimum of N=5 will be done on both lines just prior to the sample being taken. A separate excel spreadsheet is in the log recording flow rates for loops.

September 28th at 1731 (UTC) until around 1804, you may find some disruption in the signal. I was trying to re-establish some flow rates and also clean the FL. Not sure why the FL went down but the engineers are going to look at the strainers. I find it odd that our chlorophyll was increasing WHILE we entered and were transiting Johnstone Straight. Upon cleaning the FL and then purging it, there was a flush of green water, so some material must have been clogging it up a bit BUT, I still couldn’t re-establish previous flow rates. After the cleaning, the FL went down considerably. Note: I did clean the FL at the start of the cruise. The lying down configuration may be a problem?

After cleaning,the signal stayed very low for the rest of the cruise but I think it started to come back as we got into Port Hardy. The WCVI Juv Salmon survey after us should be running the TSG and will have more data.

2024-027 CRUISE SUMMARY – CTD

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | Yes | Yes |

| Calibration Information - 506 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2374 | 16Feb2023 | Factory |  |  |
| Conductivity | 4448 | 23Feb2024 | Factory | 21 Jan 2026 | Factory |
| Secondary Temp. | 4883 | 23Feb2023 | Factory |  |  |
| Secondary Cond. | 4434 | 23Feb2024 | Factory | 21 Jan 2026 | Factory |
| Transmissometer | 983DR | 13Feb2024 | Factory |  |  |
| SBE 43 DO sensor | 0997 | 25Jan2023 | Factory | 3Dec2024 | Factory |
| PAR sensor | 70613 | 24Feb2021 | Factory |  |  |
| SeaPoint Fluor. | 3949 |  |  |  |  |
| Pressure Sensor | 0506 | 22Feb2023 | Factory |  |  |
| Valeport Altimeter | 75321 | 23Sept2020 | Factory |  |  |

# TSG Make/Model/Serial#:	SEABIRD/21/2488

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature-SBE21 | 2488 | 2Feb2023 |  |  |  |
| Conductivity-SBE21 | 2488 | 2Feb2023 |  |  |  |
| Temperature-SBE38 | 0881 | 18Feb2023 |  |  |  |
| Wetlabs WETStar Fluor. For depths deeper than, and including, 125 dbar, we would wait 30 seconds before closing a bottle. For depths shallower than, and including, 100 dbar, we would wait 60 seconds before closing a bottle. | 1663 |  |  |  |  |