## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2024-034

Agency: Ocean Sciences Division (OSD)

Location: Broughton Archipelago / Discovery Islands

Project:  Aquaculture Benthic Recovery

Party Chief: Sutherland T.

Platform: Vector

Date:  14 June 2024 – 19 June 2024

Processed by: Germaine Gatien

Date of Processing: 13 August 2025 – 26 August 2025

Number of original HEX files:

Number of CTD files:  15	Number of CTD casts processed:  15

Number of bottle files: 10	Number of bottle casts processed: 9 (test cast had no sampling)

Number of TSG files:   1	Number of TSG files processed:  1

# INSTRUMENT SUMMARY

An SBE911+ CTD #0585 was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1201DR), a SBE 43 DO sensor (#1176), a SeaPoint fluorometer (#3950), an SBE pH sensor (#0691), a Biospherical PAR sensor (#4565), a reference SPAR sensor (#20518)  and an altimeter (#73171).

Seasave version 7.26.7.121 was used for acquisition.

The salinometer used at IOS was a Guildline model 8400B Portasal, serial # 73274.

An IOS rosette with 24 10L bottles was used.

A thermosalinograph (SeaBird 45 S/N 3411) was used.

# SUMMARY OF QUALITY AND CONCERNS

The Science Log and Sampling Logs were available in digital format. The header page had complete information about participants and equipment except that the pH sensor was not in the list. There were some notes in the comments sections of both logs about problems encountered, but some changes made to planned sampling were not noted in either log.

The configuration file was not set to download NMEA positions. Positions were obtained from the log and added to the data files. At the same time the station names were changed to standard syntax for this region with upper-case letters.

The nutrient samples were not logged properly and were stored in the wrong freezer; they were analyzed more than 3 months after collection. All quality flags were set to 3 by the analyst.

The pH:SBE channel was removed from events #1 to #30 because the cap was left on the sensor.

There were very few salinity calibration samples, but information was available from a subsequent cruise using the same equipment, as well as post-cruise calibration information for the primary temperature and conductivity.

Differences between the SBE DO sensor and titrated samples were larger than expect based on the post-cruise factory calibration. Observations during cruise 2024-007 which immediately followed 2024-034 had similar results, though the correction was a little smaller. It is likely that there was some contamination of the sensor membrane that was flushed out before the factory service. The lower correction for 2024-007 may be due to  higher dissolved oxygen vertical gradients so that flushing errors offset some of the differences or possibly some flushing of the membrane with use.

The SBE DO sensor has a fairly long response time so data accuracy is not as high when it is in motion as it is during stops for bottles. To get an estimate of the accuracy of the SBE DO data during downcasts a rough comparison was made between downcast recalibrated SBE DO and upcast titrated samples. Some of the difference will be due to sample collection and analysis errors, and incomplete flushing of Niskin bottles and small mismatches in depth in the comparison, so the following statement likely underestimates SBE DO accuracy. The low vertical dissolved oxygen gradients should limit the errors.

Downcast (CTD files) Oxygen:Dissolved:SBE data for this cruise are considered, very roughly, to be:

±0.25 mL/L from 10 - 100db

±0.10 mL/L below 100db

The Thermosalinograph performed well.  There was no flow meter or intake thermistor.

Salinity was recalibrated based on the results of cruise 2024-007 when more CTD cast data were available for comparison. A proxy for intake temperature was created (Temperature:Primary) by recalibrating lab temperature based on comparisons during 2024-007.

# PROCESSING SUMMARY

GENERAL

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

- During the cruise that followed a deck pressure check was run at the dock. That file was converted and pressures ranged from 0.22db to 0.33db. A factory service in late 2024 also showed that pressure was high by 0.2db.

- The digital daily log and rosette log sheets were obtained and checked for comments.

- Extracted CHL, Salinity, Nutrients and Dissolved Oxygen data were obtained in QF spreadsheets.

- The cruise summary sheet was completed.

- This is the first known cruise to use this equipment since it was last serviced.

- The equipment was used for 2024-007 shortly after this cruise; those data have been processed and the information from that cruise will help with calibration of these data.

- The calibration control files were obtained – as they stayed the same through out the cruise 2024-034-0001.xmlcon was used for processing.

- The calibration control file was not set to record NMEA positions.

##### BOTTLE FILE PREPARATION

The ROS files were created using 2024-034-0001.xmlcon. The hysteresis correction was not selected since there were no casts deeper than 100db; the Tau correction was selected.

The ROS files were converted to IOS format

The files were put through CLEAN to create BOT files. Temperature and salinity were plotted for all BOT files to check for outliers, none were found.

The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. There was no sampling from event #1, so that was removed from the list.

Sample numbers were added based on the rosette logs.

The ADDSAMP file was then sorted on event number & then sample number. It was converted to CST files using Convert Spreadsheet.

It was used to add sample numbers to the BOT files – output *.SAM.

The SAM files were bin-averaged on bottle # and called SAMAVG.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

Next, the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were saved in file 2024-034-bot-hdr.txt.

SALINITY

Salinity analysis was obtained in file QF2024-034_SAL*.xlsx. The analyses were carried out in a temperature-controlled lab within 21 days of collection. The files were simplified and saved as 2024-034SAL.csv. That file was then converted to individual SAL files.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet QF2024-034OXY*.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and saved as 2024-034_OXY.csv. That file was converted into individual *.OXY files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2024_034_NUTS*.xlsx which includes flags, comments and a precision study. The spreadsheet page with the final data was simplified and saved as 2024-034NUTS.csv. All samples were flagged due to improper storage that caused a delay in analysis. That file was converted into individual *.NUTS files.

The SAL, OXY, and NUTS files were merged with CST files in 4 steps. The files were then put through CLEAN to reduce the headers to File and Comment sections only. These files were ordered on sample number, but the SAMAVG files are ordered on bottle number, so the MRGCLN1 files were reordered on Bottle_Number and saved as *. MRGCLN1s. The MRGCLN1s files were then merged with SAMAVG files using merge channel Bottle_Number.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. None were found.

The only salinity <25psu found during rosette sampling events was for a case where the Niskin was not fired, so no recalibration will not be required for silicate.

The headers are missing positions and in one case water depth. Spreadsheet 2024-034-merge-headers.csv was prepared with that information plus station names adjusted to the usual syntax. It was run on the MRGCLN2 files, output MRGHDR.

A header check was run and no problems were found.

##### Compare

Salinity

This cruise had only 1 cast with salinity calibration samples and most samples were in the top 100db. They only deep sample was from the bottom of the cast.

The primary and secondary CTD salinity was higher than the bottles by 0.006psu and 0.007psu, respectively. Six of the other bottles showed the CTD to be reading low. This is indicative of poor flushing since the deepest bottle tends to contain lower salinity from above, while the bottles fired on the upcast tend to contain higher salinity from below.

There is insufficient evidence to guide decisions on recalibration.

For more detail see file 2024-034-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel. The casts were from a coastal environment and most casts were shallow.

The range of DO values was extremely low since dissolved oxygen was extremely well-mixed.

When samples were excluded which had a high standard deviation in the SBE DO (>0.015mL/L) and a few other outliers based on residuals, the group of samples lie in a very tight group.

The SBE DO is low by an average 8.7% and a range of 6.35% to 10.8%.

This is an extremely large difference; we do not expect calibration drift on that order unless the membrane is damaged.

On the cruise that followed (2024-007) the size of errors was on the order of 5.5% while a factory calibration in late 2024 showed drift to be on the order of 2% with no membrane damage noted.

For 2024-007 we might guess that poor DO response time would account for some of the error, but for this cruise the low vertical gradients should make that error insignificant.

There are some distinct profiles that were explored. The first 3 casts in the comparison had a DO minimum around 20m while the rest had DO decreasing throughout. So the data were examined in 2 groups.

Group 1 – casts 2, 8, 32

Group 2 – casts 36, 45, 47, 67, 71, 95

The Group 1 data indicate that SBE DO is low by 11% while for Group 2 it is low by 8%.

It is understandable that the gradient reversal will change how incomplete flushing and slow response affect values. However, the gradients are mostly very low.

Plots of DO profiles from titrated samples and CTD sensor suggest an almost constant difference.

A comparison against file pair number does suggest a gradual change with differences of about -0.41mL/L at the beginning to -0.36mL/L at the end.

A look at 2024-007 to see if there is any further hint of a gradual change shows much lower differences than during 2024-034. There is an apparent small gradual change from -0.22mL/L to -0.21mL/L, but this is small and could well be due to variable gradients in different areas.

The large drop between cruises may be partly due to incomplete flushing being a more important factor during the 2nd cruise, which would reduce the apparent error.

The analyst noted that the quality of the bottle data was low. Samples were analyzed at IOS within 2 of collection and many showed signs of poor seals. But the differences seem remarkably consistent even among unflagged cases, with just the small reduction with time.

This suggests that something got on the membrane of the CTD sensor before or very early in the cruise since Event #2 was affected.. There was no DO sampling during Event #1.

The differences do appear to get smaller with time, but that may be due to the time between collection and analysis increasing with event # or due to variable vertical DO gradients. So applying a time-dependent correction is not justified but adding 8.7% to all values looks appropriate.

See 2024-034-dox-comp1.xls.

SBE Fluorescence

Extracted Chlorophyll data were not available at the time of processing.

##### Conversion of Full CTD Files from Raw Data

All files were converted using files 2024-034-0001.xmlcon. The hysteresis correction was not selected since there were no deep casts; the Tau correction was selected. A few casts were examined and all expected channels are present.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the

full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

ALIGNCTD was run on all casts using +2.5s.

##### CELLTM

CELLTM was run using the default settings (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

At this stage it was discovered that Event #93 as shown in the log had been saved as #94. The file name was changed since the log indicates #94 was a grab. The error probably arose due to a restart of the CTD cast and either the log or the file name needs changing. The file name was changed.

Tests were run to see if the previous steps performed as expected:

- Most downcast spikes noted in conductivity profiles were removed by WILDEDIT but some were not since they contained multiple points. Most spiky data are in the upcasts.

- The dissolved oxygen alignment is definitely improved, though it is so well mixed that it is difficult to judge; ALIGNCTD appears to  have worked well.

- CELLTM worked as expected, lowering salinity where temperature is decreasing (cell warms water raising conductivity but temperature sensor reading is correct, so salinity is overestimated). Conversely, it raises salinity where temperature is increasing (conductivity too low). This brings downcast and upcast traces closer together, so T-S downcast and upcast plots are in better agreement after this step, most obvious in simple casts where temperature decreases with pressure.

DERIVE was run a second time on 2 of the deeper casts to find the differences between the pairs of temperature, conductivity and salinity channels. The differences in salinity channels is slightly larger than the 0.0014psu found in the bottle comparison, but that was based on only 3 bottles.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2024-034-0008 | 250 | -0.0005 | +0.00008 | +0.0010 | F. High, Steady |
| 2024-034-0069 | 300 | -0.0009 | +0.00005 | +0.0012 | F. High, Steady |
| 2024-034-0095 | 250 | -0.0008 | +0.00007 | +0.0013 | High, Steady |
| 2024-007-0005 | 220 | -0.0010 | +0.00004 | +0.0010 | Mod. F.Steady |
| 2024-007-0054 | 310 | -0.0012 | +0.00012 | +0.0022 | High, Steady |
| 2024-007-0075 | 310 | -0.0011 | +0.00010 | +0.0019 | High, Steady |
| 2024-007-0086 | 310 | -0.0011 | +0.00013 | +0.0020 | High, Steady |

Because this is an area where calibration sampling is shallow and more affected by incomplete flushing of Niskin bottles than offshore, there is less confidence in bottle sampling. Since there was a post-cruise calibration available for temperature, conductivity, pressure and dissolved oxygen sensors, there is an extra way to judge calibration. It appears that this CTD was not used between 2024-007 (which immediately followed this cruise) and the December 2024 service. However, some drift is expected even when equipment is sitting on the shelf, so this method is also subject to error.

During the processing of 2024-007 a comparison was done on 2 casts to compare results using the pre-cruise and post-cruise configurations files. The results showed that pressure was low by ~0.2db, primary temperature was high by 0.0014C and the secondary high by 0.0007C, and primary salinity was high by <0.002psu. The secondary salinity was found to have had major drift which must have occurred after these 2 cruises. The difference between the 2 temperature channels during this cruise is similar to that found at post-calibration, so any drift is presumed to be small. The primary salinity was found to be higher than bottles by no more than 0.002psu, so a much smaller difference than found during this cruise; this is not unexpected because there were few bottles and none were deep. The post-cruise data suggests that the dissolved oxygen values are low by about 2%.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SeaBird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number. There are no negative fluorescence values.

##### Checking Headers

No position information was entered in the Seabird headers, so those data were found in the digital log and used to prepare file 2024-034-merge-headers.csv. Also entered were station names to get the usual syntax using upper-case letters. Water depths were entered based on the log as well – they were in the raw files but there were a few errors.

A header check was run and no problems were found; fluorescence values were all >0.

A cross-reference list was produced and no problems were found.

Surface check – The mean surface pressure was 1.8db, with a minimum of 0.7 db which is within the expected range for Vector cruises in protected waters. A surface check of reversed files gave a minimum of 1.4db. Neither check is close enough to the surface to enable assessment of pressure accuracy.

The track plots were added to the end of this report. Event #1 in Saanich Inlet was excluded to better display the project area.

A check of altimeter and water depth entries was made by comparing the water depths entered in the file headers with the sum of the maximum depth sampled plus the altimetry reading when the CTD was at the bottom of the cast. Only 1 cast had a difference >2m. The log shows a considerable difference between the BE and BO sounder readings. When the BO reading was entered the check value went from -8 to +3db, so the BO entry was substituted in the header.

The sounder was obviously working well, with the largest difference likely due to shoaling.

The change was made to the CLN file, MRH, SAM, SAMAVG and MRGCLN2 files for cast #67.

##### Shift

Fluorescence

SHIFT was run on the fluorescence channel with an advance of 24 records. Plots showed this setting was appropriate.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier using Seabird software. No further alignment is needed for the DO concentration channel.

Conductivity

Tests were run on a few casts to determine the best setting to align conductivity and temperature by judging the effect on salinity as seen in T-S space. SHIFT was run on all casts using -0.6 records for the primary conductivity and -0.4 records for the secondary. Salinity was recalculated for both channels.

pH

Tests were not very helpful in assessing the best choice for alignment because the range of pH values was of the same order as noise in the signal, but using the setting chosen for 2024-007 looked like it made a slight improvement.

SHIFT was run on all casts using +60 records.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

There were no warnings.

The DEL files were zipped and sent to the CTD-QC program and DELPRED files were returned.

##### Other Comparisons

Previous experience with these sensors – This was the first use of these sensors since factory service. Cruise 2024-007 occurred after this cruise and the data had been processed. Bottle sampling suggested that the primary was high by 0.0030psu and secondary high by 0.0044psu but this was based on only 3 samples. recalibration was based on the post-cruise factory calibration.

Historic ranges – No local climatology available.

Post-Cruise Calibration – There was a post-cruise conductivity calibration available that showed excessive drift in the secondary sensor but it is assumed this occurred after this cruise. The primary had drifted high, but the largest errors were at very high salinity values. In the range of interest for this cruise the primary salinity was high by about 0.002psu. Temperature calibrations showed little drift.

##### DETAILED EDITING

The primary T-S pair were chosen for editing and eventual archiving.

All DEL files were uploaded into the CTD-QC File Processor to obtain AI generated data quality flags.

CTDEDIT was used to clean salinity and to remove records near the top and bottom of the most casts as well as records with clear evidence of shed wake corruption. Salinity was cleaned lightly.

There are many unstable features in T-S space; many are likely due to real instabilities in the water column as that is typical for this region. Both  salinity channels were displayed in CTDEDIT so that a judgment could be made about whether small spikes were likely real or due to sensor noise.

The edited files were copied to *.EDT.

All casts were edited. Comments about editing were added to the headers.

After editing T-S plots were examined for all casts; there remain many small unstable features but they are in areas when they may well reflect real conditions. No further editing was applied.

##### Recalibration

Calibration control file 2024-034-recal1.ccf was prepared to apply the following corrections:

- Addition of 0.2db to Pressure and 0.2m to Depth. (Based on deck test and post-cruise calibration.)

- Subtraction of 0.002psu from Salinity:T0:C0. (Based on the post-cruise calibration. This is significantly less than the bottle comparison result but that was based on sampling in shallow water where flushing of Niskin bottles may lead to samples coming from deeper than CTD data.)

- Subtraction of 0.0034psu from Salinity:T1:C1 (To bring it into line with the primary salinity based on the post-cruise calibration. While it was not chosen for archiving, it was recalibrated in case those data are needed later.)

- Dissolved oxygen was corrected by multiplying all values by 1.087 based on the bottle comparison. This is higher than the post-cruise calibration, but evidence suggests that there was some contamination of the sensor membrane. The cruise that followed had a similar character but the correction was 1.055, possibly due to the contamination being partly flushed away or because larger flushing errors in the later cruise due to larger vertical DO gradients.

The bottle files did not require a silicate correction given no samples associated with salinity <25psu.

Calibrate was run on MRGHDR and SAM files using 2024-034-recal1.ccf to recalibrate pressure, depth, Salinity:T0:C0, Salinity:T1:C1 and Dissolved Oxygen. Output: *.samcor1 and *.mrgcor1.

COMPARE was rerun on salinity and dissolved oxygen to ensure the corrections were applied correctly.

- For salinity both salinity channels were high by an average of 0.004 at the bottom of the cast. Any flushing error at the bottom would make the CTD appear to be reading high and there are salinity values that match the bottle about 3m above the level of the bottle stop. .

- For dissolved oxygen the correction looks excellent, with the DO sensor being low by an average of <0.001mL/L when the same data were included as in the original comparison.

CALIBRATE was run on the EDT files.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined.

Profile plots were examined to see if there any problems. There are some unstable features, but this is an areas where that may well be real due to active mixing. Overall the data look ok.

A final run of COMPARE was done to determine whether the downcast dissolved oxygen data correspond well to the bottle samples from the upcast. When displayed against pressure differences were found to be large in the top 10m where there are large gradients. Based on this comparison we can make an estimate of accuracy of downcast CTD DO. While this is expected to be a conservative assessment since some of the difference will be due to real change, problems with flushing of Niskin bottles and/or analysis errors and small mismatches in depth in the presence of large DO gradients. For this cruise vertical gradients were low so such errors should be small. Nonetheless, the following statement likely underestimates SBE DO accuracy:

Downcast (CTD files) Oxygen:Dissolved:SBE data for events #27-58 for this cruise are considered, very roughly, to be:

±0.25 mL/L from 10 - 100db

±0.10 mL/L below 100db

##### Final Bottle Files

- The MRGCOR1 files were put through SORT to order on increasing pressure.

- The following channels were removed from the bottle files: Scan_Number, Salinity:T1:C1, Temperature:Secondary, Conductivity:Secondary, Altimeter, Status:Pump, Descent_Rate, Flag, Oxygen:Voltage:SBE and pH:SBE (for casts 1-30 only).

- A second SBE DO channel with mass units was added for both the SBE DO and REORDER was run to get the pair of DO channels together.

- HEADER EDIT was run to ensure formats and units are correct

- Standards check was run and no problems were found.

- Plots were examined and the only outliers found were samples already flagged by the analyst for quality problems.

- Bottle data were exported to a spreadsheet and no problems were found.

- A header check were run. No problems were found.

- A cross-reference listing was produced for the CHE files.

##### Final CTD File Steps (REMOVE and HEADEDIT)

- For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Secondary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T1:C1, Flag and Prediction_Flag and pH:SBE (for casts 1-30 only).

- A second SBE DO channel (with umol/kg units) was derived.

- REORDER was run to arrange channels in usual order.

- HEADER EDIT was used to fix formats and channel names and to add comments to the headers.

- The Standards Check routine was run and no problems were found.

- The Header Check was run; no problems were found.

- Profile and T-S plots were examined and no problems were found.

- The sensor history was updated.

Oxygen saturation was derived and ranged from 68% to 140%. This is typical of this region with low values in the channels, higher in Phillips Arm and very high in Saanich Inlet. Inshore values are not usually helpful in judging oxygen calibration.

##### Thermosalinograph processing

There was 1 thermosalinograph file.

There was no flow meter or intake thermistor. The intake is at about 2m.

There were no loop samples.

a.) Checking calibrations

The configuration file was identical to the one used for 2024-007.

b.) Conversion of Files

The file was converted to CNV format.

The file was converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

Add Time Channel was used to add Time and Date channels.

A time-series plot was produced; a few records at the beginning look as though flow had not been fully established, so the temperature and salinity data should be editing for the first few minutes..

The track plot looks fine and was added to the end of this report.

c.)  Comparing TSG to near-surface CTD cast data

The CTD files were thinned to 2db to file 2024-034-ctd-tsg-comp.xlsx. Some casts had no data from that level.

The TSG files were opened in Excel.

Data were averaged over 2 minutes and then thinned to the times of CTD casts. For a few casts there was no corresponding TSG data.

TSG data were then added to the comparison file.

There were 13 points of comparison.

Latitude and longitude from the 2 data sources were compared to ensure a good match.

All position differences were <0.0001º so the matches are excellent.

The differences were largest in the latter part of the cruise, but this reflects ship movement to other channels and is not likely a sign of calibration drift.

The comparison shows very large differences and variability.

|  | Lat Diff | Long Diff | Ttsg-Tctd | SALtsg-SALctd |
| --- | --- | --- | --- | --- |
| min | -0.00001 | -0.00012 | -0.1175 | -3.1746 |
| max | 0.00009 | 0.00035 | 2.5831 | -0.2223 |
| median | 0.00003 | 0.00001 | 0.3583 | -0.8656 |
| average | 0.00003 | 0.00003 | 0.5829 | -1.0609 |
| std dev | 0.00003 | 0.00011 | 0.6651 | 0.8443 |

These differences are large, but this is an area of large near-surface gradients. The CTD data are generally biased towards coming from deeper in the water column than the loop water, since there are few data above 2m. As well, the data are generally noisy due to the CTD mixing up near-surface waters.

Three casts that had the lowest near-surface temperature gradients showed the TSG temperature to be high by 0.23, 0.24 and 0.27Cº. We expect the lab temperature to be higher than the intake temperature due to heating in the loop, and in June that is likely to be moderately high since ambient ship temperatures are usually well above the 10 Cº -11 Cº found at 2db. Results from the cruise that followed this one, and had more casts for comparison, indicated that temperature was low by about 0.25Cº. That looks like an appropriate choice for these data.

The salinity comparison is so noisy that it looks impossible to make a reasonable recalibration. The cruise that followed had some quiet sections where salinity appeared to be low by about 0.34psu and that was used for recalibration. That is the best available information and will be used for this cruise as well.

Calibration History

This TSG was used on 2024-007 which followed this cruise on the Vector using the same equipment. Corrections applied to those data were the addition of 0.34psu to salinity and subtraction of 0.25Cº  from temperature to estimate intake temperature..

e.) Editing

File 2024-034-001.atc was opened in CTDEDIT; temperature, salinity and conductivity data were padded for first 6 records and to clean 1 spiky point in salinity. The output file was saved as *.ED1.

f.) Addition of intake temperature proxy and Calibration

Add Channel was run to add channel Temperature:Lab, setting it equal to Temperature:Primary and channel Pressure set to 2.0db to enable derivation of sigma-T.

CALIBRATE was run to add 0.34psu to salinity and to subtract 0.25Cº from Temperature:Primary.

Derived Quantities was run to calculate sigma-T using Temperature:Primary and Salinity.

REMOVE was used to remove the following channels: Scan Number, Flag and Pressure.

h.) Preparing Final Files

HEADER EDIT was used to change the DATA DESCRIPTION to THERMOSALINOGRAPH and to change channel names to standard names and formats. Comments about the data were added.

The TSG sensor history was updated.

Particulars

1-30 – Cap left on pH sensor.

CRUISE SUMMARY – CTD & TSG

| CTD# | Make | Model | Serial# |  | Used with Rosette? |  | CTD Calibration Sheet Competed? |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 |  | Yes |  | Yes |  |  |  |
| 3 | SEABIRD | 21 | 3411 |  | n/a |  | Yes |  |  |  |
| Calibration Information - 0585 |  |  |  |  |  |  |  |  |  |  |
| Sensor |  |  |  | Pre-Cruise |  |  |  | Post Cruise |  |  |
| Name |  | S/N |  | Date |  | Location |  | Date | Location |  |
| Temperature |  | 2449 |  | 23Feb2023 |  | Factory |  | 11 Dec 2024 | Factory |  |
| Conductivity |  | 1764 |  | 17Feb2023 |  | Factory |  | 19 Dec 2024 | Factory |  |
| Secondary Temp. |  | 4484 |  | 22Feb2023 |  | Factory |  | 11 Dec 2024 | Factory |  |
| Secondary Cond. |  | 2128 |  | 17Feb2023 |  | Factory |  | 17 Jan 2025 | Factory |  |
| Transmissometer |  | 1201DR |  | 11Jan2024 |  | Factory |  |  |  |  |
| pH:SBE |  | 0691 |  | 9Mar2023 |  | Factory |  |  |  |  |
| SBE 43 DO sensor |  | 1176 |  | 1Mar2023 |  | Factory |  | 13 Dec 2024 | Factory |  |
| PAR |  | 4665 |  | 24Feb2021 |  | Factory |  |  |  |  |
| SPAR |  | 20518 |  | 16Feb2024 |  | Factory |  |  |  |  |
| SeaPoint Fluor. |  | 3950 |  | May 2023 |  | Factory |  |  |  |  |
| Pressure Sensor |  | 0585 |  | 20Feb2023 |  | Factory |  | 9 Dec 2024 | Factory |  |
| Altimeter |  | 73171 |  | 30Jan2024 |  | Factory |  |  |  |  |

| Calibration Information – TSG 3411 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 3411 | 25Jan2024 | Factory |  |  |
| Conductivity | 3411 | 25Jan2024 | Factory |  |  |

Track Plots do not include event #1 in Saanich Inlet