## Main
# Channels= | 3 |  |  |  | COMPARE Version 3.6.1 | Run at 2025/08/14 13:38:04 | Compare Excel Add-in Version 2.2
| Primary | Secondary |  | Comparison Name
Ref | Pressure | Pressure |  | Pressure - Pressure
Comp 1 | Sal_Bot | Sal0 |  | Sal0 - Sal_Bot
Comp 2 | Sal_Bot | Sal1 |  | Sal1 - Sal_Bot
Average Size: | 1
|  |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary
Primary | Secondary | Secondary | Pressure | Pressure | Pressure | (S-P) | Sal_Bot | Sal0 | Sal0 | (S-P) | Sal_Bot | Sal1 | Sal1 | (S-P)
File Name | File Name | Count | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff
2024-034-0067.mrg | 2024-034-0067.sam | 241 | 232.029 | 232.029 | 0.29566 | 0.00010681 | 30.6851 | 30.6913 | 0.0027364 | 0.0062122 | 30.6851 | 30.6925 | 0.0026523 | 0.0073624
2024-034-0067.mrg | 2024-034-0067.sam | 241 | 100.096 | 100.096 | 0.027201 | -6.1035e-05 | 30.461 | 30.4579 | 0.00044472 | -0.0031109 | 30.461 | 30.4588 | 0.00049548 | -0.0022087
2024-034-0067.mrg | 2024-034-0067.sam | 241 | 75.062 | 75.0617 | 0.025427 | -0.0002594 | 30.2833 | 30.2744 | 0.0010106 | -0.0089207 | 30.2833 | 30.2749 | 0.0010423 | -0.0083752
2024-034-0067.mrg | 2024-034-0067.sam | 241 | 50.113 | 50.113 | 0.024973 | -1.1444e-05 | 30.1841 | 30.1695 | 0.0021207 | -0.014637 | 30.1841 | 30.1699 | 0.0015522 | -0.014174
2024-034-0067.mrg | 2024-034-0067.sam | 241 | 25.128 | 25.1276 | 0.021678 | -0.00035286 | 29.8898 | 29.8773 | 0.001556 | -0.012461 | 29.8898 | 29.8786 | 0.001687 | -0.011236
2024-034-0067.mrg | 2024-034-0067.sam | 241 | 10.119 | 10.1193 | 0.022181 | 0.00033188 | 29.6366 | 29.6524 | 0.0033297 | 0.015833 | 29.6366 | 29.653 | 0.002904 | 0.016388
2024-034-0067.mrg | 2024-034-0067.sam | 241 | 4.858 | 4.8581 | 0.026196 | 5.0068e-05 | 29.3369 | 29.0444 | 0.027086 | -0.29254 | 29.3369 | 29.0473 | 0.031609 | -0.28956
2024-034-0067.mrg | 2024-034-0067.sam | 208 | 1.792 | 1.7921 | 0.029891 | 7.689e-05 | 28.3197 | 28.1767 | 0.049427 | -0.14302 | 28.3197 | 28.2045 | 0.038382 | -0.11522
Create Fit Sheet
Plot
Plot
Plot

## File List
Number of Files= |  | 1
Pair # | Primary File | Secondary File | First Row in Main | Last Row in Main | Comparison
1 | 2024-034-0067.mrg | 2024-034-0067.sam | 12 | 19 | 2024-034-0067.mrg & 2024-034-0067.sam

## Fit Index
Index to Fit Worksheets
Fit# | Description | Versus | Number of File Pairs
1 | (Sal0 - Sal_Bot) | Pressure | 1
2 | (Sal1 - Sal_Bot) | Pressure | 1

## Fit_1
Calculation of Linear Least Squares Fit For
(Sal0 - Sal_Bot)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal0 - Sal_Bot), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.91614 |  | X | CALC Y
m= | 9.9279e-05 | 1.7342e-05 |  | Std Dev of Resid |  | 0.002801 |  | 25.1276 | -0.013668
b= | -0.016162 | 0.0020902 |  | Resid Sum of Squares |  | 2.3537e-05 |  | 232.029 | 0.0068732
# of Values Extracted = |  |  | 8
# of Values Used in Fit = |  |  | 5
File | Use | Pressure | Sal0 | Sal0 | Diff (Sal0 - Sal_Bot) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | YES | 232.029 | 30.6913 | 0.0027364 | 0.0062122 | 0.0068732 | 0.00066099 |  | 232.029 | 0.0062122
1 | YES | 100.096 | 30.4579 | 0.00044472 | -0.0031109 | -0.006225 | -0.0031141 |  | 100.096 | -0.0031109
1 | YES | 75.0617 | 30.2744 | 0.0010106 | -0.0089207 | -0.0087103 | 0.00021033 |  | 75.0617 | -0.0089207
1 | YES | 50.113 | 30.1695 | 0.0021207 | -0.014637 | -0.011187 | 0.0034498 |  | 50.113 | -0.014637
1 | YES | 25.1276 | 29.8773 | 0.001556 | -0.012461 | -0.013668 | -0.001207 |  | 25.1276 | -0.012461
1 | NO | 10.1193 | 29.6524 | 0.0033297 | 0.015833 | -0.015158 | -0.030991
1 | NO | 4.8581 | 29.0444 | 0.027086 | -0.29254 | -0.01568 | 0.27686
1 | NO | 1.7921 | 28.1767 | 0.049427 | -0.14302 | -0.015984 | 0.12704
YES
YES
YES
YES
YES
NO
NO
NO
FIT vs. Pressure
FIT vs. Sal0
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Fit_2
Calculation of Linear Least Squares Fit For
(Sal1 - Sal_Bot)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (Sal1 - Sal_Bot), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.69079 |  | X | CALC Y
m= | 0.00013162 | 6.2268e-05 |  | Std Dev of Resid |  | 0.0034788 |  | 25.1276 | -0.01393
b= | -0.017238 | 0.0042684 |  | Resid Sum of Squares |  | 2.4205e-05 |  | 100.096 | -0.0040631
# of Values Extracted = |  |  | 8
# of Values Used in Fit = |  |  | 4
File | Use | Pressure | Sal1 | Sal1 | Diff (Sal1 - Sal_Bot) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 232.029 | 30.6925 | 0.0026523 | 0.0073624 | 0.013302 | 0.0059394 |  | 100.096 | -0.0022087
1 | YES | 100.096 | 30.4588 | 0.00049548 | -0.0022087 | -0.0040631 | -0.0018544 |  | 75.0617 | -0.0083752
1 | YES | 75.0617 | 30.2749 | 0.0010423 | -0.0083752 | -0.0073581 | 0.001017 |  | 50.113 | -0.014174
1 | YES | 50.113 | 30.1699 | 0.0015522 | -0.014174 | -0.010642 | 0.0035316 |  | 25.1276 | -0.011236
1 | YES | 25.1276 | 29.8786 | 0.001687 | -0.011236 | -0.01393 | -0.0026942
1 | NO | 10.1193 | 29.653 | 0.002904 | 0.016388 | -0.015906 | -0.032294
1 | NO | 4.8581 | 29.0473 | 0.031609 | -0.28956 | -0.016598 | 0.27296
1 | NO | 1.7921 | 28.2045 | 0.038382 | -0.11522 | -0.017002 | 0.098223
NO
YES
YES
YES
YES
NO
NO
NO
FIT vs. Pressure
FIT vs. Sal1
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE