
REVISION NOTICE TABLE

	DATE
	DESCRIPTION OF REVISION

	
	

	
	


PROCESSING NOTES
Cruise: 2024-035




Agency: Ocean Sciences Division (OSD)
Location: Broughton Archipelago & Strait of Georgia
Project:  CSRF Aquaculture Benthic Recovery Project
Party Chief: Sutherland T.
Platform: Vector
Date: 7-13 August 2024 
Processed by: Germaine Gatien
Date of Processing: 6 November 2025 – 26 November 2025
Number of original HEX files: 22
Number of CTD files:  22

Number of CTD casts processed: 22
Number of bottle files: 9 (1 test)

Number of bottle casts processed: 8 
INSTRUMENT SUMMARY
SBE911+ CTD 1453 was mounted in a rosette and attached were Wetlabs CSTAR transmissometer 1201DR, SBE 43 DO sensor 1438, SeaPoint fluorometer 3640, pH sensor 0691, PAR sensor 4565, SPAR sensor 20518 and altimeter 73171.  

Seasave version 7.26.7.121 was used for acquisition. 

There was no thermosalinograph data. 

SUMMARY OF QUALITY AND CONCERNS
The Science Log was in digital format. The header page had complete information about participants and equipment. There were good notes in the comments sections of both the CTD and Sampling logs about problems encountered. 
The fluorometer malfunctioned occasionally late in the cruise with sudden shifts to high values that just as suddenly resolved. This affected downcast data in cast #80 and upcast data during a bottle stop during cast #86. Affected data were replaced with pad values. The same fluorometer malfunctioned when later used during 2024-010; since it worked fine in the lab in early 2025, the problem was thought to be due to a faulty connector. 
The pH sensor data was removed from events #2 and 81 because the cap was left on the sensor.
Dissolved oxygen was recalibrated using information from cruise 2024-010. CTD salinity was not recalibrated due to low confidence in comparisons with bottle samples during both this cruise and 2024-010. However, the error in salinity is likely small with values likely within 0.003psu.
PROCESSING SUMMARY

1. Seasave
This step was completed at sea; the raw data files have extension HEX.

2. Preliminary Steps

· The digital daily log and rosette log sheets were obtained and checked for comments. 
· Salinity, Nutrients and Dissolved Oxygen data were obtained in QF spreadsheets. Chlorophyll samples were processed by the chief scientist and are not included in the bottle files.
· The history of the pressure sensor, conductivity and dissolved oxygen sensors were checked. The temperature and conductivity sensors have not been used since they were last serviced in early 2024. The pressure sensor has been used during many cruises with no record of a problem. The last deck pressure measurements were ~0.2db.
· The calibration control files were obtained – as they stayed the same throughout the cruise 2024-035-0001.xmlcon was saved as 2024-035-ctd.xmlson which was used for processing.
3. BOTTLE FILE PREPARATION 
The ROS files were created using 2024-035-ctd.xmlcon. The hysteresis correction was not selected since there were no very deep casts; the Tau correction was selected.
The ROS files were converted to IOS format
The files were put through CLEAN to create BOT files. 
Temperature and salinity were plotted for all BOT files to check for outliers, none were found. 
The BOT files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. Sample numbers were added based on the rosette logs.
The ADDSAMP file was then sorted on event number & then sample number. It was then used to add sample numbers to the BOT files – output *.SAM.
The SAM files were bin-averaged on bottle # and called SAMAVG.  
The addsamp.csv file was converted to CST files, which forms the framework for the bottle files. 

Next, the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were saved in file 2024-035-bot-hdr.txt. 
There is a discrepancy between the sampling log and daily science log. Cast #59 was the CTD cast, but the samples all have event #60. The bottle file names were corrected.
SALINITY 

Salinity analysis was obtained in file QF2024-035_SAL*.xlsx. The analyses were carried out in a temperature-controlled lab within 23-26 days of collection. The files were simplified and saved as 2024-035SAL.csv. That file was then converted to individual SAL files.  

DISSOLVED OXGYEN  

Dissolved oxygen data were provided in spreadsheet QF2024-035_OXY_11Mar2024.xlsx which includes flags, comments and a precision study. Draw temperatures are available. The spreadsheet page with the final data was simplified and saved as 2024-035OXY.csv. That file was converted into individual *.OXY files.
NUTRIENTS 

The nutrient data were obtained in spreadsheet QF2023_035_NUTS.xlsx which includes flags, comments and a precision study. The spreadsheet page with the final data was simplified and saved as 2024-035NUTS.csv. That file was converted into individual *.NUTS files. 
The SAL, OXY, and NUTS files were merged with CST files. The files were then put through CLEAN to reduce the headers to File and Comment sections only. These files were ordered on sample number, but the SAMAVG files are ordered on bottle number, so the MRGCLN1 files were reordered on Bottle_Number and saved as *. MRGCLN1s. The MRGCLN1s files were then merged with SAMAVG files using merge channel Bottle_Number. 
The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions. None were found. 
A header check was run and no problems were found.
4. Compare  

Salinity 
There were 15 salinity samples. Those from above 40m show clear signs of either incomplete flushing of Niskin bottles in the presence of high gradients or mismatches in depth between Niskin and CTD.

First, bottles were removed from the comparison if the standard deviation in the CTD salinity was >0.001psu. The primary salinity was high by an average of 0.0025psu (std dev 0.0048psu) and the secondary salinity was low by 0.0019 (std dev 0.0050). When the average was based on only the 2 deepest samples (excluding one from near the seabed), the primary salinity was found to be high by +0.0028psu and the secondary salinity was low by 0.0015psu. 

During 2024-010 which used the same equipment in October the salinity was found to be high by 0.005psu and the secondary salinity high by 0.002psu, but there were few deep samples. 

The data from this cruise are likely to be more affected by incomplete flushing in higher gradients, which would make the CTD salinity to appear to read lower than it really is. This would also make dissolved oxygen look higher than it really is. Evidence from 2024-010 suggests flushing was somewhat better then and we would expect it to be a little better due to more open waters in the sampling area.

For more detail see file 2024-035-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run between measured and bottle dissolved oxygen. There were few samples and the CTD data were very noisy, not surprising given the cruise location. A rough fit suggests that the SBE DO is low by 2% to 3%.  This is lower than found during 2024-010 when the correction found was:
CTD DO Corrected = CTD DO *1.0324 + 0.007.

The lower slope in these data may well be due to incomplete flushing and/or mismatch in depths in high vertical DO gradients. The DO values are also in a very small range, so choosing a linear fit is unreliable.
A plot of Oxygen:Dissolved and Oxygen:Dissolved:SBE versus Salinity:T0:C0 showed only one outlier and that was already flagged 3 due to a leaky Niskin. The local gradients are large, so the differences may be real; no change to the 3 flag is justified.
For more detail see file 2024-035-oxy-comp1.xls and 2024-010-oxy-compl.xls.
Fluorescence – No extracted CHL data were available for comparison.
5. Conversion of Full CTD Files from Raw Data

All files were converted using files 2024-035-0001.xmlcon. The hysteresis correction was not selected since there were no deep casts; the Tau correction was selected. A few casts were examined and all expected channels are present.  
6. WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the 
full cast files (*.CNV).  

Parameters used were: 
Pass 1    Std Dev = 2 
Pass 2    Std Dev = 5 
Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

7. ALIGN DO

ALIGNCTD was run on all casts using +2.5s.

8. CELLTM

CELLTM was run using the default settings (α = 0.0245, β=9.5) for both the primary and secondary conductivity. 
9. DERIVE 

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.
Tests were run to see if the previous steps performed as expected:

· Some spikes noted in conductivity profiles were removed by WILDEDIT. Some remain but are in the upcast sections.

· The alignment of dissolved oxygen is definitely improved with offsets between downcast and upcast temperature and oxygen profiles similar. Downcast features are in good agreement with temperature traces, so ALIGNCTD appears to  have worked well. 
· CELLTM worked as expected, lowering salinity where temperature is decreasing (cell warms water raising conductivity but temperature sensor reading is correct, so salinity is overestimated). Conversely, it raises salinity where temperature is increasing (conductivity too low). This brings downcast and upcast traces closer together, so T-S downcast and upcast plots are in better agreement after this step.
DERIVE was run a second time for the deepest cast to find the differences between the pairs of temperature, conductivity and salinity channels. The sensors were used during cruise 2024-010 so a few differences from that cruise are included (shaded).
	Cast #
	Press
	T1-T0 
	C1-C0
	S1-S0
	Descent Rate

	2024-035-0083
	340
	-0.0002
	-0.00043
	-0.0042
	Mod, V. Steady

	2024-035-0083
	420
	-0.0002
	-0.00041
	-0.0042
	Mod, V. Steady

	2024-010-0026
	340
	-0.0001
	-0.00061
	-0.0060
	High, V. Steady

	2024-010-0052
	340
	 0.0000
	-0.00053
	-0.0054
	High, Moderate

	2024-001-0057
	300
	-0.0002
	-0.00046
	-0.0044
	High, F. Steady


Temperature differences are small with insignificant variation.

Conductivity and salinity differences vary somewhat but not greatly.
There are too few deep casts to determine if there is temporal drift.
The CTD salinity channel differences underway are reasonably close to the range found during bottle comparison (about 0.0035psu-0.0048psu). Some differences are expected since the CTDs were in motion for one and stopped for the other and sampling was limited.
10. Conversion to IOS Header Format

The IOSSHELL routine was used to convert SeaBird 911+ CNV files to IOS Headers. 
CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number. There are no negative fluorescence values.
11. Checking Headers

A header check and cross-reference list was run. The pH value was very low for one cast indicating that the cap was left on. All were checked and the pH data are bad for events #2 & 81 so will be removed later. There is a note in the log that #12 had the cap on, but those data look ok.
Surface check – The mean surface pressure was 1.5db which is reasonable for the Vector in protected waters. The minimum, 0.37db, occurred during cast #18 with pumps on and “in-water” values for conductivity and salinity. The maximum was 3.97db.
The track plots were added to the end of this report, one with station names and the other with cast number. Event #1, a test cast in Saanich Inlet, was not included to enable better display of project area.
A check was made of the water depths entered in the file headers by adding the maximum depth sampled plus the altimetry reading when the CTD was at the bottom of a cast. 21 of the 22 casts had a check value within ±4db, but for event 89 the check value was 27m. The sounder entry is clearly inappropriate since the CTD got to 83m. The log shows a sounder reading of 61m at the beginning of the cast and 89m at the end. Using 89m gives a check value of 1db. So the water depth entry was changed in the header to match the end value. The same correction was made to bottle files from that cast.
12. Shift
Fluorescence

SHIFT was run on the fluorescence channel with an advance of +24 records. Plots showed this setting was appropriate. The plots also show that there were patches of fluorescence values that shift suddenly to much larger values, then shift back to believable values. Problem casts were:

· Event #80 – This is the only one where the problem was during the downcast, affecting data below 12db. The bad data will be removed later in CTDEDIT.
· Event #81is affected briefly during the 20db stop but not during the 10s window used in the comparison.

· Event #86 is affected during most of the 20db stop including the time when the bottle was fired. That fluorescence value will be padded at the REO stage, late in processing.
Dissolved Oxygen 

The Dissolved Oxygen voltage channel was aligned earlier using Seabird software. No further alignment is needed for the DO concentration channel.
Conductivity
SHIFT was run on all casts using -0.8 records for the primary conductivity and -0.5 records for the secondary—because these values were used on another cruise with these sensors. Salinity was recalculated for both channels. Test plots indicate the settings are reasonably helpful with sometimes smaller or larger shifts seeming slightly better, but nothing clearly “the winner”. 
The primary and secondary T-S plots also differ little.

pH
SHIFT with an advance of 24 scans was run on the pH channel based on tests; this brings the downcast profiles into better correspondence with temperature. Most of the offset in the down/up pH profiles is due to the upcast.

13. DELETE

The following DELETE parameters were used: 

Surface Record Removal: Last Press Min
Maximum Surface Pressure (relative): 10.00
Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure 
Sample interval = 0.042 seconds. (taken from header)
There were no warning in the log file.
14. Other Comparisons

Other experience with these sensors – 
2025-035: This is the first cruise since most of the sensors were serviced in early 2024. The pressure sensor was used on many previous cruises and the last deck pressure measurement was ~0.2db.
2024-010: This cruise occurred after -035 but was partly processed earlier, since it had better calibration sampling for dissolved oxygen. The primary salinity was found to be high by about 0.0048psu and the secondary salinity was found to be very close to bottles but the standard deviation was very high. Deep bottles suggested the secondary salinity was high by 0.002psu. 
Dissolved oxygen had many bottles and an appropriate correction was found to be 
CTD DO Corrected = CTD DO *1.0324 + 0.007. 
The fluorometer used during 2024-035 performed poorly during 2024-010 but the problem appears to have been due to connections rather than the sensor.
Historic ranges – No local climatology was available for this region.
Post-Cruise Calibration – There were no post-cruise calibrations available. 

15. DETAILED EDITING
The secondary T-S pair were chosen for editing and eventual archiving. While either look useful, the secondary calibration looks better, low by 0.002 for this cruise and high by 0.002psu for 2024-010. The 2024-010 data are likely more accurate given that flushing of Niskin bottles is likely better. But neither look as reliable as we would like. Nonetheless, salinity is likely good to ±0.003psu.
All DEL files were uploaded into the CTD-QC File Processor to obtain AI generated data quality flags.
All casts were edited lightly, with salinity being cleaned lightly and records corrupted by shed wakes being removed from near the top and bottom of many casts.

A second round of editing was applied to cast #80 to pad fluorescence values below 13.1db. 

The edited files were copied to *.EDT.
After editing T-S plots were examined for all casts; there remain many unstable features but they are in areas when they may well reflect real conditions. No further editing was applied.
The EDT files were zipped and returned to the CTD-QC AI project.
16. Recalibration
There were no data with Bottle Salinity <25psu, so silicate does not require recalibration.
For the 2024-035 dissolved oxygen, the fit looks unreliable due to the very small range of DO values and a lot of scatter; this is not unexpected given the area of sampling is likely to have poor flushing of Niskin bottles. For 2024-010 there is a better range and flushing is likely a little better. In the absence of any other information, the 2024-010 result will be applied to 2024-035. This may be an underestimate of the error.
For 2024-035 salinity, the fit makes it clear there was poor flushing and there were only 2 samples below 200m. Using only those 2 samples the primary salinity is found to be high by 0.0025 and the secondary low by 0.0015psu. 

For 2024-010 the primary salinity was found to be high by 0.0048psu and the secondary very close to the bottle. However, based on 2 samples near the mouth of Juan de Fuca where flushing is usually better the primary is high by 0.0052psu and the secondary high by 0.0017psu.

Recalibration will not be attempted given the few bottles and large scatter. It is likely that the secondary salinity is within 0.003psu.

CALIBRATE was run using file 2024-035-recal1.ccf to correct channel Oxygen:Dissolved:SBE using: 
CTD DO Corrected = CTD DO *1.0324 + 0.007.
COMPARE was rerun after this step to ensure the correction was applied correctly and it was. The CTD DO looks slightly high, as expected, given the bottles likely did not flush completely.

The step of comparing with CTD downcast data with upcast bottles was skipped as there is too much variability in this region for it to be instructive.

17. Fluorescence Processing
A median filter with a fixed width of 11 was applied to the fluorescence channel.

18. BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure
Averaging interval = 1.000
Minimum bin value =   .000

Average value will be used.
Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined. 
Profile plots were examined to see if there any problems. A few small unstable features remain near the surface; they may be real so no further editing was applied. 
19. Final CTD File Steps (REMOVE and HEADEDIT)
· For all casts REMOVE was run to remove the following channels:

Scan_Number, Temperature:Primary, Conductivity:Primary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T0:C0, Flag and Prediction_Flag.
For casts #2 and #81 channel pH:SBE was also removed.

· A second SBE DO channel (with umol/kg units) was derived. 
· HEADER EDIT was used to fix formats and channel names and to add comments to the headers.
· The Standards Check routine was run and no problems were found. 
· The Header Check was run; no problems were found. 
· Profile and T-S plots were examined and no problems were found. 
· Oxygen saturation was derived and plotted; as is usual for this area the results varied widely, from 60% in tidally-mixed areas up to 170% in Saanich Inlet where DO was >10mL/L. 

· The sensor history was updated. 

20. Final Bottle Files
· The MRGCOR1 files were put through SORT to order on increasing pressure. 

· Scan_Number, Temperature:Primary, Conductivity:Primary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T0:C0 and Flag.

For casts #2 and #81 channel pH:SBE was also removed.
· A second SBE DO channel with mass units was added for both the CTD DO and REORDER was run to get the pair of DO channels together.
· For cast #86 the fluorescence value was padded at 20db and the header range adjusted.
· HEADER EDIT was run to ensure formats and units are correct
· Standards check was run and no problems were found.
· Bottle data were exported to a spreadsheet and no problems were found.
· A header check were run. No problems were found. 
· A cross-reference listing was produced for the CHE files.
Particulars – including notes from log
2 Nutrients filtered about 12 hours after collection.

2 & 81. Forgot pH cap. 
59. Shown as event 60 on Sampling log but 59 in Daily Science log. Changed samples to #59.
60. Log says BUCKET, Sampling log says ROS at 60. Switched 5m and chlmax bottle on the fly. Sample labels have wrong depth. 

81. Niskin 1 has been leaking from spigot.

85. Big fluoro layer at 10m.
91. Bucket sample only – no bottles fired.
CRUISE SUMMARY – CTD
	CTD#
	Make
	Model
	Serial#
	Used with Rosette?
	CTD Calibration Sheet Competed?

	1
	SEABIRD
	911+
	1453
	Yes
	Yes

	Calibration Information - 1453

	Sensor
	Pre-Cruise
	Post Cruise

	Name
	S/N
	Date
	Location
	Date
	Location

	Temperature


	2023
	8Feb2024
	Factory
	
	

	Conductivity
	3500
	23Feb2024
	Factory


	
	

	Secondary Temp.


	6565
	17Feb2024
	Factory


	
	

	Secondary Cond.


	5043
	23Feb2024
	Factory


	
	

	Transmissometer


	1201DR
	11 Jan2024


	IOS
	
	

	PAR
	4615
	24Feb2024
	Factory
	
	

	SPAR
	20518
	16Feb2024
	Factory
	
	

	SBE 43 DO sensor
	1438
	27Feb2024
	Factory
	
	

	SeaPoint Fluor.
	3640
	17Mar2024
	Factory
	
	

	pH
	0691
	9Mar2023
	Factory
	
	

	Pressure Sensor
	1453
	21July2022
	Factory
	
	

	Altimeter
	73171
	30Jan2024


	Factory
	
	


The first cast in Saanich Inlet is not shown in the plots below to enable better display in project area.
[image: image1.png]North Latitude

2024-035 Event #s

127.00 125,00
50,80 50,80
50,70 [s0.70
.60 L 0.50
50.50+ - 50.50
50,40 .40

127.00 12500

‘West Longitude




[image: image2.png]North Latitude

2024-035 Stn Names

127.00 1.5 125.00 15.50 125.00
.00 ‘ i ! 50.00
" Ww
5
50,70 ‘ kS F50.70
0 4
MG
&
" Sinon]:
Q}ﬁq‘?.ﬁ?
s
.50 PR Fs0.60
@ 4 4
O
50.50~ - 50.50
.40 5.4
127.00 125.00

West Longitude




PAGE  
13

