## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2024-061

Agency: IOS, Ocean Sciences Division, Sidney BC

Chief Scientist: Bokvist J.  		Platform: Nita Maria

Location: WCVI / Barkley Sound	Project: Follow the Fish Barkley Sound Purse Seine

Date: 28 May 2024 –30 September 2024

Processed by: Germaine Gatien

Date of Processing: 5 February 2026 –    4 March 2026

Number of original HEX files: 92 	Number of CTD files: 76

# INSTRUMENT SUMMARY

A SeaBird Model SBE-19 CTD (s/n 8241) was mounted with a Wetlabs ECO Fluorometer #FLRT-7495 and dissolved oxygen sensor SBE43 #4354.

# SUMMARY OF QUALITY AND CONCERNS

This program involved repeat trips over 5 months. The same CTD was used throughout.

There were several problems with the naming of the files. The original files were uploaded and saved with standard-style names but the event number part of the file names did not always correspond to event numbers in the event spreadsheet. There was a warning about this in the documentation provided. It is very awkward to change names when the correct file name has already been used for a different cast.

An extra complication is that times were reported as being in PST but it appears that this should be PDT. After much effort to match header times to those in the event spreadsheet, it was clear that there were still problems with events #61 to #65. For those casts the times recorded in the spreadsheet could not be related to the times in the file headers. Since there were only data from the 10m soak period, the effort to establish correct names was abandoned as there are no useful data to archive.

The deployment scheme varied:

- For events #1-50 there was a surface soak followed by a full cast.

- For events #51-55 & #57 there was a drop to 10m for a soak but no return to the surface before running the full cast. For those casts upcast data were processed. Data quality is lower from the upcast but this ensures some data are available from 0 to 10m and when there is no rosette in use the upcast data are not greatly affected by shed wakes.

- For some events the recommended method was used with a drop to 10m for a soak followed by a return to the surface before running the full cast.

Some files contained no useful data and no CTD files were produced:

- For Events 28, 70, 72 & 73 there was only surface data acquired.

- File #33 contained only a little upcast data with many pressure spikes.

- For files #62-67 there was only data from the surface to 10m with no data after the 10m soak.

There was no calibration sampling, but there were post-cruise calibrations from January 2025 for pressure, conductivity and dissolved oxygen sensors, so calibration checks were possible and corrections were applied to CTD salinity and dissolved oxygen. The corrections to salinity were supported by limited calibration sampling during 2024-005 when the same equipment was used. Pressure required no recalibration.

There was also a post-cruise fluorometer calibration but the factory service report indicates that the sensor was tuned before the later calibration, so the results are not suitable for the 2024 data. We consider fluorescence data nominal.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea.

##### Preliminary Steps

The cruise summary sheet was completed.

##### Conversion of Raw Data

The configuration file used at sea was correct. It was saved as 2024-061-ctd.xmlcon.

All casts were converted using that file.

The file for event #60 did not convert because the file contained no data due to low battery.

A sample of casts were plotted and all channels were present and profiles looked ok for most. But some casts contained very few data. All very small files were examined.

- Casts #60 and #61 have only a few surface records. These will not be processed further.

- Casts #62 to #67 contain only data from the drop to the soak level.

- All expected variables are present. Pressure looks quite smooth.

Temperature, conductivity and fluorescence profiles look normal.

The minimum fluorescence was reasonable, ~+0.04ug/L except for spikes to negative values at the bottom of many casts. If this remains after the Delete step then Clean can be re-run to pad such values.

##### WILDEDIT

The only file found with notable spikes was cast #33 which was full of spikes and attempts to remove them left few data with large jumps in pressure. The data appear to be from the upcast only. This cast will not be processed further.

The only other spikes noted in the data occurred at the beginning or end of casts or included many points, and will be removed in the normal course of editing. So WILDEDIT was not run.

##### FILTER

The resolution of this instrument appears to be good, so the pressure does not obviously need filtering. However, a test showed some improvement in salinity and dissolved oxygen if a low-pass filter, size 1s, was run on pressure, so that was done for all casts.

The temperature and conductivity were examined and the usual approach of applying a cosine filter size 8 in routine WFILTER did a good job of removing small reversals.

This step had little effect on Dissolved oxygen data.

##### ALIGNCTD

ALIGNCTD was run on all casts to advance the DO channel by 0.5s based on results from other cruises in the past few years that used this CTD.

##### CELLTM

CELLTM was run on all casts using the SeaBird recommended parameters, (α, 1/β) = (0.04, 8).

##### DERIVE

Program DERIVE was run to calculate salinity and dissolved oxygen concentration (tau correction included).

Plots were examined to see if steps 5, 6 and 7 had worked well.

The low-pass Filter step made a slight change to pressure but had no notable effect on salinity.

The Windows Filter step did improve the salinity, removing small instabilities in T-S space.

The alignment step worked well to bring the offset between downcast and upcast values in dissolved oxygen into good alignment with those in temperature.

The CELLTM step made no noticeable difference likely because the temperature range is small in such shallow casts.

A post-cruise calibration was available, but attempts to convert the raw data using those values failed as the routine chooses calibration information from those listed within the file. When the pre-cruise calibration data were removed from the raw file for event #74, the data could be produced with new calibration data. So both files were opened in EXCEL for comparison.

Data from the 10m soak and upcast were removed. Differences were found.

File 2024-061-post-vs-pre-comp.xlsx shows the comparisons.

| Differences Post-Pre | Pressure | Temperature | Conductivity | Oxy mL/L | Fluor | Depth | Salinity |
| --- | --- | --- | --- | --- | --- | --- | --- |
| average | -0.02624 | -0.00004 | 0.00112 | 0.11930 | 0.68489 | -0.02600 | 0.01128 |
| stdev | 0.00120 | 0.00005 | 0.00001 | 0.06167 | 0.56711 | 0.00126 | 0.00054 |
| max | -0.02400 | 0.00000 | 0.00116 | 0.37020 | 3.46000 | -0.02300 | 0.01180 |
| min | -0.02800 | -0.00010 | 0.00109 | 0.05000 | 0.39820 | -0.02800 | 0.00930 |

The oxygen voltage had 0 change, as expected so is not included. 
Pressures from the post-cruise were lower by 0.026db which is excellent for this type of sensor.

Temperature differed by <0.0001Cº which is extremely small. No recalibration is needed.

Conductivity, salinity and fluorescence differences were more significant.

The pre-cruise salinity values appear to be low by about 0.011psu. There is a little pressure dependence above 6db where differences are about 0.0094psu but they rise quickly to about -0.0110 and then slowly rose to -0.1117psu at the bottom. Applying an average correction of 0.011psu looks appropriate.

Fluorescence values all are lower when using the pre-cruise calibrations. The difference is dependent on fluorescence with a linear fit:  Post-cruise Fluor = 2.0833 * Pre-cruise Fluor + 0.325

This is a very large change, but a check of the service report for this CTD indicates that the fluorometer was repaired before the post-calibration was done, so the result is not relevant for 2024 data.

The dissolved oxygen is low by an average of 0.12mL/L but the standard deviation is very high and we know that the error varies with DO concentration, so the data were investigated in a different way to estimate calibration error.

When calibration samples are available we compare the difference between SBE DO and titrated samples from bottles and plot that against SBE DO. We can approach these data in the same way by finding the difference between the pre-cruise SBE DO data and the data using post-cruise calibrations and plotting that against the pre-cruise DO values. The fit looks very much like what we get from a bottle comparison. The data from close to the surface fit very badly, so data above 2.5db were excluded. 

The fit is similar to those found with bottle data and indicates a correction of +3.6% which is typical of this type of sensor. When only data below 6db are included the fit is tighter, and when the offset is forced to =0, the data are low by 3.555%.  So applying a correction of 1.036 looks appropriate.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert the CNV files to IOS Headers.

A spreadsheet was provided with time (labelled as PST) and positions of casts as well as station names and bottom depths.

There was a note included saying that the names assigned are not in the order they were collected.

At this point a cross-reference list could be produced to try to understand what event numbers belonged to which files. There were 18 files that needed new file names and event numbers in the headers.

One problem with confirming the matches is determining if the spreadsheet information is really in PST and if the file headers are really in UTC. If the UTC time is right then PST times are off by 1 hour. If the header is off by 1 hour, then the PST time looks right. During 2024-005 when the same CTD was used, the start times in the files agree with the log times which are said to be in UTC.  So it is likely that the header times are correct and the log times are in PDT. They also vary somewhat from the header times usually by just a few minutes but occasionally by up to 30 minutes; this is typical of small-boat cruises. For event #46 the time in the spreadsheet was clearly wrong as it matched that from the previous cast.

The file names were adjusted to match the order of acquisition – this required opening files to check the instruments time stamp. The earlier steps were repeated using the renamed files. This took a few attempts as issues were discovered.

Next, spreadsheet 2024-061-mrh.csv was prepared for addition of header positions and other information.

Merge CSV file to Headers was run using that file as input.

Track plots were checked for each leg to make sure the order of sampling looked right and it did.

Some files were then removed from the cast list as they contain no useful data: 28, 33, 61, 62-67, 70, 72 & 73.

CLEAN was run to add event numbers and replace pad values in pressure with interpolated values.

##### Checking Headers

A header check was run and the only problems found likely come from surface and bottom spikes.

A cross-reference list was produced and no problems were found

Surface Check was run and the average was -0.015db with very low salinity, so right at the surface. This is very similar to the result in previous use of this sensor. No recalibration will be applied to pressure.

##### CLIP

Plots were examined to determine how much surface data needs to be removed in order to ensure DELETE chooses data from the final drop from the surface.

- First, casts 58-67 & 74 were examined to see if there is any useful data. These casts had an initial drop to about 10m and no data after that. In some cases where there was enough data at the stopping depth to make a judgment, it appears that the data from the drop are not reliable. After temperature dropped slightly (likely a shed wake passing through) it then increased significantly. In most cases there are not enough data to make a judgment, but we do expect that the pumps would not have been on long enough for temperature and salinity to equilibrate. So those casts will not be processed further.

- Next, casts 51-57 which had a soak at 10m followed by a drop to the bottom were examined. The top 10m of data do not look useful. An attempt will be made to use upcast data. While generally of lower quality than downcast data, the absence of a rosette reduces the wake effects, so they are probably sufficient and the surface data is likely of importance to the users. Those casts were put through REVERSE with output *.REV.

- For other casts plots were made to indicate how many scans need to be removed to ensure soak data are not selected by DERIVE. CLIP was run using a table. Plots were made after this step and the settings were fine-tuned for 2 casts.

##### SHIFT

Conductivity

Tests run on 3 other cruises (2023-010, 2023-064, 2024-005) using this CTD found shifts of -2.2 records to conductivity made the best improvement to stability in T-S space. That setting was applied to all casts and the results look good.

Fluorescence

The fluorometer was not pumped, so a shift in alignment is expected to be small or unnecessary. Profile plots of temperature and fluorescence were examined and confirm that the alignment is ok.

Dissolved Oxygen

This channel was aligned earlier, but checks were made by examining plots of temperature and dissolved oxygen. The alignment looks good.

##### DELETE

DELETE was run on all casts using the following parameters:

Surface Record Removal: Last Press Min.	Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 5 points) was deleted from 10db to 10db above the maximum pressure.

COMMENTS ON WARNINGS: There were no warnings.

The DEL files were zipped and submitted to the CTD QC model. The upcast files were rejected due to scan numbers decreasing with pressure. The file extensions DEL were changed to DELPRED in order to avoid being lost in the processing.

##### DETAILED EDITING

All DEL files were copied to *.EDT and then all DELPRED files were copied to *.EDT so both those that went through CTD QC process and those that did not will be kept.

CTDEDIT was used to edit 28 files; this was limited to removal of a few records near the top and/or bottom corrupted by shed wakes and to clean salinity, mostly due to poor response to temperature gradients. Many other casts had low salinity at the surface that looks odd in T-S space but the features are stable and large near-surface salinity gradients are to be expected in this region, so no editing was applied to such features. Notes of editing details were made in the headers.

The edited files were copied to *.edt.

Plots showed some unbelievable fluorescence near the surface for events #54 (76mg/m^3) and #57 (66mg/m^3). A text editor was used to replace those values with pad values.

The EDT files were zipped and submitted to the CTD QC process.

##### Calibration checks

Sensor History – This equipment was used during 2023-064, 2023-010 and 2024-005. Salinity samples were taken during 2023-010. CHL and Salinity comparisons showed both the fluorometer and salinity sensor were performing within expectations, but data were insufficient to state salinity accuracy. 2024-005 overlapped this program and had some calibration sampling; salinity was found to be low by about 0.01psu and fluorescence was about 60% of CHL when CHL was >5ug/L.

Historic Ranges – There was no local climatology available, but all temperatures fell within a climatology

for a large area around Barkley except that surface temperature were sometimes high in the top few metres in July. Salinity was always low near the surface at Sproat Narrows and generally low at the surface for the sites to the east especially in the earlier months, as expected close to shore due to runoff. The more westerly sites (Lyall, Henderson, Turtle Island and Village Reef) were fully within the salinity climatology throughout the program.

The climatology in use covered the whole year and a large area, so it is not very informative.

Post-cruise calibrations – January 2025 calibrations were available for pressure, conductivity, dissolved oxygen and fluorescence sensors, but the fluorometer was repaired before that calibration.

##### CALIBRATE

Pressure does not require recalibration. Salinity and dissolved oxygen do need correction.

Based on the comparison described in §8 CALIBRATE was run to apply the following corrections:

Salinity:T0:C0 Corrected = Salinity:T0:C0 +0.011

CTD DO Corrected = CTD DO *1.036

##### Fluorescence Filter

The fluorescence data do not require filtering.

##### Bin Average, Remove, Derive DO in mass units, Reorder

The COR1 files were bin averaged using 1db bins.

REMOVE was run to remove Scan_Number, Oxygen:Voltage, Descent Rate, Flag and Prediction_Flag channels.

Dissolved Oxygen was derived in mass units.

Oxygen saturation was calculated. Plots of near-surface saturation show a range of 90% to 165%. This matches the results from the 2023 Follow the Fish program.

The highest values (>140%) were seen at Lyall Point, Nanat and Chup Point in August. The values were lowest at Sproat Narrows ranging from 92% to 110%.

A final check of plots showed a little more editing was needed for 2 events – 43 and 44, so those were put through steps 14-18 again.

##### HEADER EDIT and final checks of CTD files.

Header Edit was used to fix headers, fix formats and to add comments about processing.

A cross-reference listing was produced.

A header check and standards check were run on the CTD files and a few corrections made until all errors were corrected.

Plots of CTD casts were examined; there a few slightly unstable surface features but they are small and may be real, so no further editing was applied.

Particulars

1-50. CTD soak at surface.

18. No data acquired. Replaced with repeat cast 19.

28. Surface data only.

29. Replacement for cast 28.

33. Full of pressure spikes. De-spiking leaves little data with huge jumps in pressure.

51-57. CTD soak done at about 10m for 2-minute, then dropped right to bottom and brought back up. Used upcast data.

56. Surface data only. Replaced with repeat cast #57.

58. Used fishing set GPS not oceanographic GPS. Data only to 10m.

59. Data only to 10m.

60. Low battery. No data acquired. Lyall Point.

61. Low battery. No data acquired – no start time recorded. Hand Island

62-67. Data to 10m only with no surface soak. Data not useful –  Not processed.

66. Replaced with repeat cast 67. Only 10m soak data in both. Not processed.

70, 72 & 73. Surface data only. Not processed.

CRUISE SUMMARY

| Cruise ID#: 2024-061 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Dates: Start: 28 May 2024 End: 130 September 2024 |  |  |  |  |  |
| Location: Barkley Sound |  |  |  |  |  |
| Chief Scientist: Bokvist J. |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 19+ | 8241 | No | Yes |  |

CTD CALIBRATION INFORMATION

Make/Model/Serial#:	SEABIRD/SBE19+/8241	Cruise ID#:		2024-061

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 8241 | 9Dec2022 | Factory | 10 Jan 2024 | Factory |
| Conductivity | 4345 | 9Dec2022 | Factory | 10 Jan 2024 | Factory |
| ECO Fluorometer | 7495 | 4Dec2022 | Factory | 23 Jan 2025 | Factory |
| SBE43 Oxygen | 4354 | 20Dec2022 | Factory | 16 Jan 2025 | Factory |
| Press | 8241 | 3Dec2022 | Factory | 2 Jan 2025 | Factory |

Station positions are approximate as they varied among different legs of the program.