## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2025-005					Agency: OSD

Location: Salish Sea					Project: Salish Sea Biophysical Survey

Chief Scientist: Jackson J..				Platform: Vector

Date: 8 April 2025 – 15 April 2025

Processed by: Germaine Gatien		Date of Processing: 27 August 2026 – 29 September 2026

Number of original HEX files: 	70	Number of processed CTD files: 69 (11file upcast only)

Number of rosette casts: 29		Number of processed CHE files: 27 (1 test & 2 split casts joined)

# INSTRUMENT SUMMARY

CTD #585 was mounted in a rosette and attached were 2 Wetlabs CSTAR transmissometer (498DR & #1883DG), a SBE 43 DO sensor on the primary pump (#1176), a WetLabs ECO Fluorometer -unpumped (#2216), a Biospherical QSP_200L4S sensor (#4615), a Biophysical QSR_2200 SPAR sensor (#20518) and a Valeport altimeter (#73171).   (The log has the wrong serial # for the red transmissometer.)

Seasave version 7.26.7.121 was used for acquisition. The data logging computer was Lenovo ThinkCentre. The deck unit was a Seabird model 11+. An IOS rosette with 24 10L General Oceanics, Ocean Test Equipment bottles was used.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log Book was in good order with some comments about problems encountered and a detailed list of equipment. There was an error in the red transmissometer serial #. The sampling logs were in good order except that there were 2 sheets identified as event #1. One was a test, but having an event number confuses the program used to prepare bottle files. The log also indicated that there was a TSG in use, but no TSG data were found.

The standard deployment procedure for was used for most casts with data acquisition starting after the CTD returned to the surface after a soak at 10m. For events 1-4 acquisition started at the beginning of the cast so data from the soak needed to be removed. For event #78 the inlet protocol was followed with only a surface soak of 1 minute.

There were 2 WetLabs CStar transmissometers in use during this cruise:

Channel Transmissometer refers to sensor #1185DR (650nm - red)

Channel Transmissometer:Green refers to sensor #1883DG (530nm - green)

For comparison with other Institute of Ocean Sciences cruises, note that the transmissometer wavelength is 650nm unless otherwise stated.

Events #50 and 51 occurred at the same site because the first cast was interrupted a power loss. All downcast data came from #50. A single bottle file (2025-005-0050.che) was produced by combining the 2 casts.

The SBE DO sensor has a fairly long response time so data accuracy is not as high when it is in motion as it is during stops for bottles. This will be especially true when vertical DO gradients are large. To get an estimate of the accuracy of the SBE DO data during downcasts (after recalibration) a rough comparison was made between downcast SBE DO and upcast titrated samples. Some of the difference will be due to problems with flushing of Niskin bottles and/or analysis errors and small mismatches in depth in the presence of large DO gradients, so the following statement likely underestimates SBE DO accuracy.  There were also frequent DO gradient reversals complicating the comparison.

Downcast (CTD files) Oxygen:Dissolved:SBE data for both CTDs for this cruise are considered, very roughly, to be:

±0.50 mL/L from 10-50db (generally too high in mixed layer)

±0.30 mL/L from 50db to 100db

±0.10 mL/L from 100db to 200db

±0.05 mL/L below 200db.

For details see files 2025-005-dox-comp3.xls

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained.

- Nutrients, extracted chlorophyll, dissolved oxygen and salinity data were obtained in QF spreadsheet format from the analysts.

- The cruise summary sheet was completed.

- This was the first know use of the conductivity sensors since their most recent factory service. The pressure and dissolved oxygen sensors were used during 2025-005.

##### BOTTLE FILE PREPARATION

ROS files were created using 2025-005-ctd*.xmlcon. The tau corrections were selected.

The ROS files were converted to IOS format and put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files to check for significant outliers.

In each of casts #27 and #43 there were 3 consecutive outliers around 75db in channel Salinity:T0:C0 that were replaced with pad values using a text editor.

The BOT files for #50 and 51 were renamed as 50.bota and 50.botb and joined with output 2025-005-0050.BOT.

Ultraedit was used to fix the bottle number and firing order number for the bottle fired last (from *.botb).

Add Sample Number was used to add sample #s to the BOT files with output SAM.

The SAM files were bin-averaged on bottle #s for use in COMPARE, but they are not needed for CHE prep.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2025-005-bot-hdr.txt which will be updated as needed during processing.

DISSOLVED OXGYEN

Oxygen samples were analyzed at sea using an automated Winkler titration System (Metrohm Dosimat model 876 and a UV light source and detector with a 365nm filter controlled by LV02_876 software designed and constructed by Scripps Institution of Oceanography) with modifications based on Carpenter (1965) and adhering to WOCE protocols (Culberson 1991).

For details including a precision analysis, see document QF2025-005_OXY*.xls.

EXTRACTED CHLOROPHYLL

Chlorophyll samples were filtered onto 25mm GF/F filters and immediately frozen and stored in a -80C freezer until analysis.

All samples were returned to the Institute of Ocean Sciences and analyzed 3-10 days after the cruise on a Turner Trilogy 7200-000 fluorometer, using methods described in Holm-Hansen et al (1965).

Fluorescence readings taken before and after sample acidification were used to calculate chlorophyll and phaeopigment concentrations.

For further details on methodology, see document QF2025-005_CHL*.xlsx Methodology sheet.

NOTE: Flags and comments apply to chlorophyll values only.

No flags or comments are assigned for Phaeo-Pigment values.

NUTRIENTS

Nutrient samples were collected in plastic tubes and quick frozen in aluminum blocks stored in a -20C freezer. For samples 400m and deeper a second set of samples are collected and stored cool for silicate analysis. All samples were returned to the Institute of Ocean Sciences and they were analyzed using an

Astoria-Pacific Analyzer following methods described in IOS Nutrient Methods (1996) Barwell-Clarke and Whitney.

For details including a duplicate analysis, see document QF2025-005nuts*.xlsx.

SALINITY

Salinity samples were collected in 200 mL type ll glass bottles with disposable

nylon inserts and screw caps supplied by Ocean Scientific International Limited.

They were analyzed in a temperature-controlled lab on a Guildline 8400B Salinometer

standardized with IAPSO standard seawater 8 - 13 days after collection.

For details including a duplicate analysis, see document QF2025-005_SAL*.xlsx.

The 4 chemistry files were simplified as CSV files. A note was made of cases where the oxygen sample has a comment “ALL:” with flag 3 due to do bottle integrity issues. Flag 3 was added to affected sample data in the CSV files for nutrient, salinity and extracted chlorophyl.

The chemistry CSV files and  ROS files were submitted for preparation of preliminary bottle files.

One problems were found – there were 2 rosette sheets with event #1 – one was a test, so was removed.

The Python routine was run to produce files CHEPRELIM.

Error messages were found and checked against logs and converted files. A few needed investigation:

Event 1 — Bottle Position mismatch between CTD and rosette – Most bottles were not sampled and are not noted in the sampling log. OK.

Event 51 — Bottle Position mismatch between CTD and rosette – Only 1 bottle fired after a power cut. This file needs to be joined to file #50.

Event 57 was misnamed in raw files. They were renamed as 58.

The data from those files were exported to a spreadsheet and compared to the sampling logs. See file 2025-005-bottles.xlsx.

Problems were found in a few files, mostly concerning CHL samples:

- Event 11 – CHL sample 50 missing from QF file. Intended duplicate. I found 50A in raw data in CHL spreadsheet. No problem noted in logs. (Later added by analyst)

- Event 19 – CHL data confusing – Log indicates samples 60, 61(dups), 62, 63. QF has 56, 60, 62, 63 with no dups. Raw data has samples 56, 60, 61a,  62, 63.  Sample 56 has very low value so could be from 75db as indicated in QF file. (Later fixed by analyst.)

- Event 50 – Sample 173 entered as event 51. This was from a single bottle fired after a power cut. Spreadsheets for DO and NUTS were adjusted to show event 50 for the samples originally entered as event 51. JOIN used to combine 50 and 51.cheprelim files. A text editor was used to insert the CTD values based on data in the SAMAVG files.

- Event 55 – CHL Sample 186 missing from QF sheet but is in raw data. (Later added by analyst.)

- Event 58 was missing due to the wrong event # in the file name and in the 2nd line in the file, so the python routine was rerun.

The problems with CHL were referred to the analyst and a corrected spreadsheet was provided.

Flag “3” was added to CHL samples #24, 50, 51, 97, 168, 186, 262, 263 since the comment assigned by the dissolved oxygen analyst applies to all samples.

File 2025-005-bottles .xlsx was updated with the new CHL data.

A preliminary header check was run and no problems found.

CTD salinity was checked and there are values <25psu, so silicate correction is needed.

##### Compare

Salinity

Compare was run with pressure as reference channel. The file was saved in XLS format immediately. The routine does not work well unless this is done.

Outliers were removed where the standard deviation in the SBE salinity was >0.001psu and pressure was <52db. Two other outliers were excluded.

The primary salinity was found to be low by 0.0005psu (std dev 0.0031psu) and the secondary were high by 0.0036psu (std dev 0.0031psu).

Bottles from the bottom were checked to see if they were out of line, but the only 2 that were out of line had been excluded due to high standard deviations.

The differences would be higher by about 0.001psu if more outliers were removed until the fit was quite flat with pressure as recommended by Seabird.

Differences versus file pair # were very flat for the primary salinity and fairly flat for the secondary.

None of the outliers look enough out of line to conclude there is a problem with sampling or analysis.

For full details for the COMPARE run see file 2025-005-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run using the converted SBE data from the SAM files and the samples from the CHEPRELIM files.  Pressure was the reference channel.

For the initial fit the offset was forced to -0.0138 as that was established in the previous cruise when there were many samples with DO < 1mL/L. For this cruise there are only 2 bottles from Saanich Inlet with low DO from this cruise, and those were both very low. The bottle at 150db shows the SBE DO to be low by 0.014mL/L and the 172db bottle at the bottom of the cast shows SBE DO to be low by 0.048mL/L. The bottom bottles tend to contain water from higher in the water column so we expect the SBE DO to read somewhat lower. The 150db sample being from the upcast has the opposite but usually small effect especially since the local DO gradient is very low. So it would seem that the SBE may be reading slightly high in very low DO.

However, using offsets based on those 2 bottles leads to fits that do not look good overall, whereas the fit is better if we use the result from 2025-001 when the SBE DO appeared to be low by 0.0138mL/L for very low DO. The difference is likely due to the fact that the DO sensor is slow to respond in very low DO as is found in Saanich Inlet. It is also likely that the bottles do not flush well in the quiet waters of Saanich Inlet, so the results are difficult to interpret.

Seabird do say that the offset is not expected to change with time.

The correction applied for 2025-001 was :

CTD DO Corrected = CTD DO *1.0199 + 0.0138

The fit for this cruise if we leave the offset free is:

CTD DO Corrected = CTD DO *1.0166 + 0.0431

The fit for this cruise using the 2025-001 offset is:

CTD DO Corrected = CTD DO *1.0223 + 0.0138

The larger slope may be due to calibration drift, but there are possible explanations. So a preliminary study was done of cruise 2025-006 where there was more DO sampling and it showed an even larger slope (1.028) when the same offset was used. So there does appear to be drift occurring with this sensor.

The only outlier that looks suspicious is for sample #227 which was already flagged 46. One sample looks closer to sample #228 and is enough out of line to conclude there is a problem with sampling or analysis.

For full details for the COMPARE run see file 2025-005-dox-comp1.xls.

Fluorescence –

COMPARE was run with extracted chlorophyll and CTD Fluorescence from an ECO fluorometer using pressure as the reference variable.

The CHL and fluorescence values were very close during event #1 in Saanich Inlet where both values were very high. This may be partly due to extreme vertical gradients. As CHL samples likely come from a little lower in the water column, the ratio will be less reliable and the true ratio is likely somewhat lower than 1.

At other sites the pattern is what we usually see with the fluorometer reading higher at low CHL values. As CHL values rise, the ratio ECO/CHL falls. The two values were close when CHL was about 1ug/L. These patterns are typical of these fluorometers.

For details see file 2025-005-chl-fl-comp1.xls.

##### Conversion of Full Files from Raw Data

Hex files were converted using configuration files 2025-005-ctd*.xmlcon.

One file did not convert due to missing header information, but was converted using the matching configuration file.

All expected channels were found and profiles look reasonable.

Depth, descent rate and SBE salinity were derived later in processing.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, depth, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

This step worked well at removing some large spikes in both conductivity channels. Temperature profiles checked had no significant spikes to remove.

##### ALIGN DO

ALIGNCTD was run on all casts to advance the oxygen voltage by +2.5s, a setting which has worked well in the past for this type of sensor. The results were examined during step #9 and found to be good.

##### CELLTM

The noise in the upcast data makes tests for the best parameters for this routine very difficult to interpret. In the past when upcast data were not so noisy, the default setting of (α = 0.0245, β=9.5) was generally found to be the best choice.

A few casts were checked for this cruise and the default setting does improve the data. CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity, dissolved oxygen concentration and descent rate.

Plots were examined and the alignment of dissolved oxygen looks good.

T-S plots show that the CELLTM step worked reasonably well although the profiles are often too noisy to judge.

DERIVE was run a second time on the 2 deepest casts to find the differences between the pairs of temperature, conductivity and salinity channels.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2025-005-0027 | 350 | +0.0005 | +0.00039 | +0.0039 | High, Moderate |
| 2025-005-0077 | 350 | +0.0005 | +0.00039 | +0.0039 | High, V Steady |
| “ | 500 | +0.0004 | +0.00039 | +0.0039 | “ |
| “ | 650 | +0.0004 | +0.00039 | +0.0038 | “ |

The differences are small and consistent with time and pressure.

The salinity differences are a little higher than expected in newly calibrated sensors, but it is possible they were used on a previous cruise that has not yet been processed.

There may be slight increases with time but the later casts are in shallow water, so this is not clear.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert Sea-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

All CLN files were copied to *.CLIP

Most casts used the normal soak procedure of soaking at 10m, returning to the surface and then starting archiving. CLIP was used to remove data from a few casts where that procedure was not used with output CLIP1:

1. Only a surface soak. Removed 580 records.

3. Records include the 10m soak. Removed 4600 records.

4. Records include the 10m soak. Removed 2400 records.

78. Used inlet procedure of surface soak. The soak only lasted about 45s. Salinity and temperature had just begun to settle down, so it is not clear if equilibration had been achieved. Dissolved oxygen values tend to be a little high in the top 10m when there is no 10m soak due to bubbles. It is recommended that waits be 2 minutes when there is no 10m soak. 1050 records were removed.

The CLIP1 files were copied to CLIP.

##### Checking Headers –

- The cross-reference check and header check were run. Some problems were detected by looking at the ranges of some variables::

  - A large spike in fluorescence that included both positive and negative values was found in event #62 during the upcast. Fortunately, it did not occur during the 10s-window selected for the bottle comparison. No action was required.

  - During event #60 there were some negative pressure values. They occurred at the beginning of the cast as the CTD rose from about 2db up to the surface. The negative pressures correspond to times when red transmissivity was 0 and when pressure moved from -0.00db to +0.139db the transmissivity rose sharply. This suggests that the pressure accuracy is excellent.

  - No water depth was entered in the header for event #1. The log indicates that the sounder depth was 180m. This was entered in the header of the CLIP file.

- Surface check was run. The average was 2.1db, as expected since acquisition generally started after the 10m soak.

- Cruise tracks were plotted and added to the end of this report.

The altimeter and water depth readings from the headers of the CLN files were exported to a spreadsheet. A check value was calculated by subtracting water depth from maximum depth sampled plus altimetry header entry. Where that number was > 6m checks were made to see if the log entry differed from the header entry and whether the altimetry signal at the bottom provided a reliable header value.

Casts 13, 26, 43 and 51 were investigated. For 13, 26 and 43 the log entry looks excellent, so the headers were edited to use that value in both CHEPRELIM and CLIP files.

Cast #51 is from the same site as cast #50. Plots show very similar depth so it makes sense that 51 has same water depth of 168m but the cast was much shallower at about 75m. In this case the problem is with the altimetry header. The altimetry was extremely spiky at the bottom of the cast and this was misinterpreted by the algorithm. The altimetry header was removed from this cast.

##### Shift

Fluorescence

The ECO fluorometer was not pumped and the data look well aligned with temperature though low gradients and jitter in the ECO signal limit the ability to check on this.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it looked excellent. No further alignment is needed for that channel.

Conductivity –

Tests were run on a selection of casts to find the alignment shift best for the conductivity sensors as judged by noise in T-S space. The best choice was -0.8 for the primary and -0.5 records for the secondary channel.

SHIFT was run on both channels using those settings with output SHFC0 and SHFC1.

##### DELETE

DELETE was run on the SHFC1 files.

The following DELETE parameters were used for all casts:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 9 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Experience with these sensors since last factory service –

2025-001- This was 1st use of conductivity sensors since recal. Pressure sensor was used on 2025-001 with no problem noted. Dissolved oxygen sensor was used on 2025-001 and was recalibrated using slope/offset=1.0199/0.0138.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed where available. No local climatology is available for events 1 and 69 – 78, and the climatology does not go all the way to the surface. All salinity values fell within the climatology except for a few low salinity values at the bottom of cast #83 in the eastern Juan de Fuca Strait where such excursions are common. Temperature values were all within the climatology. There  is no evidence of calibration problems.

Post-Cruise Calibration –A December 2025 post-cruise report on dissolved oxygen sensor found it low by about 2%.

Repeat Casts –There were no casts deep enough or close enough to enable test of repeatability.

##### DETAILED EDITING

The *.DEL files were zipped and submitted to the CTD-QC-File Processor. DELPRED files were returned.  The DELPRED files were copied to *.EDT.

The secondary channels were chosen for editing and archiving.

CTDEDIT was used to remove records that appear to be corrupted by shed wakes. Salinity was cleaned to remove spikes that appear to be due to small misalignment or instrumental noise. All files required some editing.

Notes about editing applied were added to the files. The output files *.edu were copied to *.edt.

After editing, T-S plots were examined for all casts. There are some small unstable features mostly during the top and bottom of casts. This is an area of active mixing and unstable features may be real, so editing was limited to cases where there was evidence of shed wake corruption or salinity spikes caused by small misalignments of T and C.

The edited files were submitted via the CTD-QC File Processor.

##### Corrections to Salinity and Dissolved Oxygen Concentration

File 2025-005-recal-sil.ccf was used to corrected silicate where salinity in bottle files was <25psu.

The primary salinity does not need recalibration. The secondary was recalibrated to bring it into correspondence with the primary in case it is later found to be more appropriate for some casts.

File 2025-005-recal1.ccf was prepared to subtract 0.0041psu from channel Salinity:T1:C1 and to apply the following correction to dissolved oxygen:

CTD DO Corrected = CTD DO *1.0199 + 0.0138

CALIBRATE was run on the CHEPRELIM files and SAM files using 2025-005-recal1.ccf.

COMPARE was rerun using CHEPRELIMCOR1and SAMCOR1 files after recalibration.

A check was made that the salinity recalibration was done correctly and it was.

Next, the dissolved oxygen data were examined by running COMPARE. When the same data were included as in the original fit, the results were excellent. The SBE dissolved oxygen values were found to be low by an average of 0.012mL/L. If the flushing is poor, the SBE values may be lower by more than that. However, there were many local reversals in DO gradient, so flushing errors would be of variable sign which would reduce that error. The vertical gradients were only very high at the surface.

See file 2025-005-dox-comp2.xls.

CALIBRATE was run on the EDT files. A quick check showed that it worked as expected.

##### Fluorescence Processing

No filtering was applied since the ECO fluorometer was not pumped.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen T-S plots were examined There remain some unstable features but those are expected in this region. A few casts were opened a second time in CTDEDIT to see if further editing would be useful, but this did not prove to be the case.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and delayed response and noise in CTD data.

Downcast files (bin-averaged to 0.5m-bins) were thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles. The average difference was 0mL/L with a standard deviation of 0.6mL/L. Differences below 40db were all within 0.3mL/L. From 200db down all differences were slightly negative. This suggests that bottles were quite well flushed and/or the SBE DO was reading a little low. The frequent gradient reversals are another source of variability.

Plots were made against pressure using the same data to enable an estimate of differences versus pressure.

Downcast (CTD files) Oxygen:Dissolved:SBE data for both CTDs for this cruise are considered, very roughly, to be:

±0.50 mL/L from 10-50db (generally too high in mixed layer)

±0.30 mL/L from 50db to 100db

±0.10 mL/L from 100db to 200db

±0.05 mL/L below 200db.

For details see files 2025-005-dox-comp3.xls

No further recalibration of channel Oxygen:Dissolved:SBE is suggested by these results.

##### Final CTD File Steps (REMOVE and HEADEDIT)

For all casts REMOVE was run on the AVG files to remove the following channels:

Scan_Number, Temperature:Secondary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T1:C1 & Flag and Prediction_Flag.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channel pairs together.

HEADER EDIT was used to fix formats and channel names and to add comments about processing.

The Standards Check routine and Header Check were run; it was found that the edited version of cast #63 had been accidentally deleted, so the processing of that cast was repeated from the editing phase to the end. No problems were found.

Profile and T-S plots were examined and look ok.

PAR and SPAR sensors worked throughout the cruise; the values differed greatly with SPAR generally higher though not always.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. As is usual for this areas values varied greatly, ranging from 80% to 140% with the lowest values in the southern part of the Strait of Georgia, Haro Strait and Juan de Fuca Strait. They gradually increased from south to north until all values north of the Gulf Islands were >110%.  Saanich Inlet was very high at 140%. These values are in the expected ranges for these areas in spring.

##### Final Bottle Files

The CALIBRATE steps were run earlier in order to do COMPARE runs, but they need to be repeated due to some CHL corrections done later to bottle files; those did not affect the DO and SAL comparisons.

CALIBRATE was run on CHEPRELIM files using file 2025-005-recal-SIL to adjust silicate where salinity was <0. (Output CHEPRELIMSIL).

CALIBRATE was run on CHEPRELIMSIL to correct SBE dissolved oxygen and secondary salinity – output CHEPRELIMCOR1.

The file was sorted on pressure – output MRGSORT.

For all casts REMOVE was run to remove the following channels: Scan_Number, Temperature:Secondary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter, Salinity:T1:C1, Flag.

A second SBE DO channel with mass units was added for both the CTD DO and titrated DO and REORDER was run to get the pairs of DO channels together.

EDIT HEADERS was run to fix formats, remove one header entry concerning winch location, fix channel names and add comments about analyses and CTD processing.

Data were exported from the CHE files to file 2025-005-bottles-final.xlsx. The entries were compared with the rosette log sheets and no problems were found.

Standards check was run. No problems were found.

The track plot looks fine.

Plots of each file were examined and no problems were found.

A cross-reference listing and header check were produced for the CHE files.

Particulars  - Notes from Daily Science Log and Rosette logs

1. Bottles 1-7 sampled; all bottles fired to check integrity.

1-4. Archiving started as soon as rosette went into water.

5-end. Archiving after soak.

12. NMEA data missing; entries made after conversion to CNV based on log entry and System UTC.

17. Swapped out bottle 13 as it has a chip missing on top

27. Lots of current, drifted off station during cast

50. Stn 9. Lost power of ship, cut off everything 75m of upcast - recovered CTD.

51. Stn 9b. Bottle 5 fired at 75m.  Bottle data from 50 & 51 combined in processing. Named 50.

57/58. Bottle file named as event 57 should be 58. File name fixed.

78. Followed Inlet protocol - 1min soak at surface, 5m off bottom.

85. Fighting current and line angle so some of the depths may be slightly off.

CRUISE SUMMARY

Cruise ID#: 2025-005

Dates: Start: 8 April 2025 End: 15 April 2025

Location: Salish Sea

Chief Scientist: Jackson J.

CTD CALIBRATION INFORMATION

Make/Model/Serial#:	SEABIRD/SBE911/0585	Cruise ID#:		2025-001

| CTD# | Make | Model | Serial# |  | Used with Rosette? |  | CTD Calibration Sheet Competed? |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 |  | Yes |  | Yes |  |  |
| Calibration Information - 0585 |  |  |  |  |  |  |  |  |  |
| Sensor |  |  |  | Pre-Cruise |  |  |  | Post Cruise |  |
| Name |  | S/N |  | Date |  | Location |  | Date | Location |
| Temperature |  | 2449 |  | 11Dec2024 |  | Factory |  | 22Jan2026 | Factory |
| Conductivity |  | 3321 |  | 18Dec2024 |  | Factory |  |  |  |
| Secondary Temp. |  | 4484 |  | 11Dec2024 |  | Factory |  | 22Jan2026 | Factory |
| Secondary Cond. |  | 2754 |  | 21Jan2025 |  | Factory |  |  |  |
| Transmissometer |  | 498DR |  | 25Mar2025 |  | Factory |  |  |  |
| Transmissometer |  | 1883DG |  | 11Jan2024 |  | Factory |  | 1May2026 | Factory |
| SBE 43 DO sensor |  | 1176 |  | 13Dec2024 |  | Factory |  | Dec 2025 | Factory |
| PAR sensor |  | 4615 |  | 24Feb2021 |  | Factory |  |  |  |
| SPAR sensor |  | 20518 |  | 16Feb2024 |  | Factory |  |  |  |
| WetLabs Fluor. |  | 2216 |  | 8 Mar2017 |  | Factory |  |  |  |
| Pressure Sensor |  | 0585 |  | 9Dec2024 |  | Factory |  |  |  |
| Valeport Altimeter |  | 73171 |  | 30Jan2024 |  | Factory |  |  |  |