RBR CTD DATA PROCESSING NOTES

Cruise: 2025-097

Agency: OSD

Locations: Clayoquot Sound

Project: Clayoquot Sound Biological Oceanography  (Ha-oom Fisheries Society)

Party Chief: Kelly Young

Platform: Unknown

Date: January 28, 2025 – December 10, 2025

Processed by: Samantha Huntington

Date of Processing: May 6, 2026 – May 11, 2026

Number of Raw files:34 		Number of Processed Files34

Instrument Summary

Equipment:

For casts 1-24, an RBR Maestro3 CTD (s/n 203655) with a Marine CT Conductivity and Temperature sensor, an RBR Pressure sensor, an RBRcoda T.ODO Fast Temperature and Dissolved Oxygen concentration sensor, and a Li-cor PAR sensor, as well as a Seapoint Tu turbidity sensor.

For casts 25-34 an RBR Maestro3 CTD (s/n 231803) with a Marine CT Conductivity and Temperature sensor, an RBR Pressure sensor, an RBRcoda T.ODO Fast Temperature and Dissolved Oxygen concentration sensor, and a Li-cor PAR sensor, as well as a Seapoint Tu turbidity sensor and an RBR pH sensor.

The CTDs belongs to the Nuu-chah-nulth Tribal Council.

All sampling was done at 8Hz.

Summary of Quality and Concerns

A cast list of times and locations was provided“2025 Haoom_CTD_metadata.xlsx”.

The data overall look good with some bad data visible during the soak and at the bottom of most casts.

Times and dates were matched to the log provided, bottom depth from the log did not always match the maximum pressure reached.

The PAR and pH data were bad on all casts and these channels will be removed from the files.

CTD casts were conducted as the sun sets using the following protocols as provided by the Ha-oom Fishery Society.

- Depth Measurement: Determine the depth at the site in feet using a depth sounder. Convert it to meters and subtract 5 meters to establish the maximum CTD depth.

- Avoiding Depth Contours: Ensure CTD sampling avoids depth contours and potential obstructions to prevent snags or entanglements.

- Cap Management: Remove and securely store all caps in a safe location.

- Initiate Recording: Twist the knobs to initiate data recording.

- Calibration Procedure: Allow the CTD to calibrate for 2 minutes near the water's surface before lowering it to the desired depth.

- Controlled Descent: Allow the CTD to freefall at the desired recording rate, eliminating the need for consistent manual adjustments.

- Data Upload Protocol: After cast is done, stop recording by twisting the knob. Upload data on the same day.

Processing Summary

- Conversion to IOS Headers

Eleven files were provided:

2025-097 203655_20250129_0904.rsk with profiles 1 and 2.

2025-097 203655_20250211_1838..rsk with profiles 3 and 4.

2025-097 203655_20250324_1836.rsk with profiles 5-7.

2025-097 203655_20250423_1931.rsk with profiles 8-10.

2025-097 203655_20250521_1517.rsk with profiles 11 and 12.

2025-097 203655_20250626_1053.rsk with profiles 13-16.

2025-097 203655_20250731_1332.rsk with profiles 17-20.

2025-097 203655_20250814_0924.rsk with profiles 21-24.

2025-097 231803_20250930_1348.rsk with profiles 25-28.

2025-097 231803_20251124_1733.rsk with profiles 29-31.

2025-097 231803_20251210_1318.rsk with profiles 32-34.

34 profiles were found in the files and they were extracted using python function READ_EXCELrsk().  T

All processes for IOS Shell conversion were run twice to account for the different variables associated with each CTD.

Two files (2025-097_CTD_Data_1-24.csv and 2025-097_CTD_Data_25-34.csv ) with all the data including event numbers and a single line of headers was prepared using python function MERGE_FILES().

A 6-line header was inserted using python function Add_6Lineheader_2().

File “2025-097_header-merge.csv” was created, based on the information provided by the chief scientist.

- Column “File_Name”: entries were derived from the event number.

- Column “LOC:LATITUDE”: latitude was provided and reformatted to “XX XX.XXXX N !(deg min)”.

- Column “LOC:LONGITUDE”: longitude was provided  and reformatted to “XX XX.XXXX W !(deg min)”.

- Column “LOC: Event Number”: entries were event numbers.

- Column “LOC: STATION”: all stations were set according to the information provided.

- Column “LOC: Water depth”: all depths were set according to the information provided.

The sampling site was mapped (Figure 1) using from “2025-097_header-merge.csv” using python function Plot_Track_Location() to check the location of all casts.

Prior to conversion to IOS header format, the presence of zero-order holds were checked using Python function Plot_Pressure_Diff().  Zero-order holds were found (Figure 2.) and these values were replaced with Nan using the python function Correct_Hold().

A new csv file was created “2025-097_CSV_DATA-6Linedr_corr_hold.csv”  (for each set of profiles) and the corrected values were checked in python function Plot_Pressure_Diff().  Zero-order holds were found to be removed (Figure 3.).

CONVERT Spreadsheet Files was run to produce files with IOS Header format. Header entries of “Administration”, “File” and “Instrument” were filled in this step.

The routine “Merge:CSV Files to headers” was run to add location headers to the IOS files.

Raw data were plotted and examined and there is some bad data at the top and bottom of many casts, and a few spikes in Salinity and Conductivity.  Dissolved Oxygen Concentration is a measured rather than derived channel for this CTD and a shift will be needed for better alignment.  Dissolved Oxygen Saturation is messy and will be removed later in processing. PAR and pH will be removed as well.

Turbidity and PAR have some spikes and some negative values which will be addressed later.

The pH sensor must have been malfunctioning as the pH values were 0.5 at all points in all casts measured with the CTD with a pH sensor.

All bad data and spikes will be examined again after DELETE.

Next CLEAN was run to add a start time and event numbers to headers.

- Data processing

- Correction to Pressure: a few negative pressures were found at the end of some upcasts, corresponding Conductivity was lo and Pressure was not corrected.

- Time was provided in local time, this was adjusted to UTC used ADD TIME CHANNEL, 8 hours were added to casts 1-4, and 29-35. 7 hours were added to casts 5-28.

- CLIP: Pressure is steady for a variable number of scans. Initial records were removed until the downcast began using file “2025-097_CLIP.csv”. Pressure is very steady for a while below 0.5 m, and it appears to be a soak of about 2 minutes at 1m.

- Filter: a Gull-winged filter, size 5, was applied to temperature, conductivity, dissolved oxygen,  and pressure. Salinity will be calculated in the next step.

- SHIFT: Based on suggested values in document “Guidelines for processing RBR CTD profiles”, the alignment of temperature and conductivity was corrected by applying a shift of -2 scans in conductivity.

- SHIFT: Better alignment with Oxygen profiles was found by advancing it by 11 scans. The advice given in document “Guidelines for processing RBR CTD Profiles” was that an advance between 2 and 3 seconds is appropriate. T-O plots before and after alignment were compared.  Oxygen alignment is improved.

- Delete was run to remove records with a descent rate lower than 0.3m/s over 8 points. This was not applied in the top 10m to avoid loss of surface records as the CTD began its descent.

- Profile plots were examined after DELETE and confirm that plots show reasonable values for salinity and conductivity.

- Bad data was removed from the bottom of casts 9, 16, 17, 18 and 25.

- A second channel for Dissolved Oxygen Concentration was added with units umol/kg.

- REORDER was run to put the channels in the correct order.

- Final checks and header editing

- REMOVE was run to remove the following channels from all casts: Date, Time:UTC and Event.  Dissolved Oxygen Saturation, PAR and pH were also removed.

- BIN AVERAGE was used to metre-average data.

- CALIBRATE was run to convert conductivity units to S/m using file 2025-097-recal2.ccf.

- Header Edit was used to fix channel names and format as listed below:

- Salinity:CTD  ==> Salinity

- Turbidity:Seapoint [FTU] ==> [NTU]

- Oxygen:Dissolved ==> Oxygen:Dissolved:RBR [umol/L]

- Oxygen:Dissolved:NEW ==> Oxygen:Dissolved:RBR [umol/kg]

CLEAN was run to reset the Maximum and Minimum values in the Header.

- Header Check was run and no problems were found.

Figure 1 – location of casts.

Figure 2 – Zero Order holds

Figure 3 – zero order holds removed