RBR CTD DATA PROCESSING NOTES

Cruise: 2026-035

Agency: OSD

Locations: Vancouver Harbour

Project: OPP Vancouver Harbour Oceanography Survey

Party Chief:  Blanken H.

Platform: Doug Anderson

Date: February 16, 2026 – February 20, 2026

Processed by: Samantha Huntington

Date of Processing: May 22, 2026 - ???

Number of Raw files: 126 	Number of Processed Files: 126

Instrument Summary

Equipment: RBR Concerto3 CTD (s/n 204694) with Turbidity sensor (RBRcoda Tu) and oxygen sensor (JFE Advantech Rinko III B s/n ???)

Sampling frequency was at 8Hz.

Summary of Quality and Concerns

A cast list of times and locations was provided, “Tap Log Export 20260219.csv”, this contained the log data from a few previous trips and this data was ignored.  The information from the csv file was used to create  “2026-035CTDLogSheet.xlsx”.  The event numbers skipped 71, so this was added in and the event numbers were changed accordingly.  A total of 126 casts were recorded.

Events 15 and 16 were both at station EB 42, this cast was repeated because the Oxygen sensor cap was left on during the first cast.  Dissolved Oxygen will be removed from this cast later in processing.

Four Ruskin files were provided but the final one , “241289_20260220_1609.rsk”, contained all casts including two test casts from a few days before this field work began. Those two casts will be skipped.

The data overall look good, there are some bad data at the top of some of the upcasts and a few spikes in a few of the soaks, these will be examined again after the DELETE step.

Processing Summary

- Conversion to IOS Headers

File 241289_20260220_1609.rsk contained 128 profiles, of which the first two were skipped and the remaining 126  were extracted using python function READ_RSK().

A single file (2026-035_CTD_Data.csv) with all the data including event numbers and a single line of headers was prepared using python function MERGE_FILES().

A 6-line header was inserted using python function Add_6Lineheader_2().

File “2026-035_header-merge.csv” was created, based on the information provided by the chief scientist.

- Column “File_Name”: entries were derived from the event number.

- Column “LOC:LATITUDE”: latitude was provided and reformatted to “XX XX.XXXX N !(deg min)”.

- Column “LOC:LONGITUDE”: longitude was provided  and reformatted to “XX XX.XXXX W !(deg min)”.

- Column “LOC: Event Number”: entries were event numbers.

- Column “LOC: STATION”: entries were taken from the Log file.

- Column “LOC:WATER DEPTH: entries were taken from the Log file

The sampling site was mapped (Figure 1) using from “2026-035_header-merge.csv” using python function Plot_Track_Location() to check the location of all casts.

Prior to conversion to IOS header format, the presence of zero-order holds were checked using Python function Plot_Pressure_Diff() and python function check_for_zoh().  This function returned a need for a correction and this was done using python function READ_RSK() but this time using the pyRSKTools internal correct hold function. A new file, “2026-035_CTD_DATA_6linehdr_corr_hold.csv

CONVERT Spreadsheet Files was run to produce files with IOS Header format. Header entries of “Administration”, “File”,  “Instrument” and “Geographic Area” were filled in this step.

Next CLEAN was run to add a start time and event numbers to headers.

The routine “Merge:CSV Files to headers” was run to add location headers to the IOS files.

Raw data were plotted and examined:

- Salinity looks good.

- Temperature looks good.

- Conductivity looks good.

- Oxygen looks good with some bad data at top and bottom of some casts which will  be examined after Shift and Delete.

- Turbidity has some spikes at the surface and some very high values during the soak, this will be examined again after DELETE.

A record number was  added to each record using the routine Add time channel.

- Data processing

- Correction to Pressure: Negative pressures were not seen in all casts but they were present in many with corresponding Conductivity values inconsistent with the RBR being on deck, a pressure correction of 0.1 was added to Pressure and Depth.

- CLIP: Pressure is steady for a variable number of scans. Initial records were removed until the downcast began using file “2026-035_CLIP.csv”.

- Filter: a Gull-winged filter, size 3, was applied to temperature, conductivity, and pressure. Salinity will be calculated in the next step.

- SHIFT: Based on suggested values in document “Guidelines for processing RBR CTD profiles”, the alignment of temperature and conductivity was corrected by applying a shift of -2 scans in conductivity.

- SHIFT: Better alignment with Oxygen profiles was found by advancing it by 11 scans. The advice given in document “Guidelines for processing RBR CTD Profiles” was that an advance between 2 and 3 seconds is appropriate. T-O plots before and after alignment were compared.

- Derive Oxygen Concentration:

- Data from the shift files was extracted using python function get_rbr_data().  Oxygen:Dissolved:Concentration was derived using python function convert_oxy().  Both of these functions can be found in the python file rbr_oxy.py.  A file containing the derived oxygen “2026-035_Oxygen_for_IOSshell.csv” was created so that it could be converted into the IOS shell files.

- MERGE was used to merge the derived Oxygen:Dissolved:Concentration back into the IOS Shell file.

- Delete was run to remove records with a descent rate lower than 0.3m/s over 8 points. This was not applied in the top 10m to avoid loss of surface records as the CTD began its descent.

- Profile plots were examined after DELETE and confirm that plots show reasonable values for salinity and conductivity.

- Bad Data Removal: Bad data was removed from the bottom of casts 1, 3, 5, 7, 10-16, 18-20, 25, 30, 38, 43, 46, 49, 51, 64-68, 71-72, 77, 80-81, 83, 88, 89, 91-93, 107 and 109.

- Final checks and header editing

- REMOVE was run to remove the following channels from all casts: Date, Time:UTC, Event and Record Number.

- Dissolved Oxygen was removed from cast 15.

- REORDER was run to order the channels.

- BIN AVERAGE was used to metre-average data.

- CALIBRATE was run to convert conductivity units to S/m using file 2026-035-recal2.ccf.

- Header Edit was used to fix channel names and format as listed below:

- Temperature: deg C(ITS90) ==> deg C (ITS90)

- Oxygen:Dissolved:Saturation:RBR [%]:

- Oxygen:Dissolved:Saturation ==> Oxygen:Dissolved:Saturation:RBR

- Format: F11.4 ==> F8.2

- Pressure [decibar]

- Format: F11.2 ==> F7.1

- Depth [metres]

- Units:  meters ==> metres

Salinity:CTD ==> Salinity

- Oxygen:Dissolved:Rinko [mL/L]

- Oxygen_mL_L ==> Oxygen:Dissolved:Rinko

- Format: F10.6 ==> F7.2

- Oxygen:Dissolved:Rinko [umol/kg]

- Name:   Oxygen_umol_kg ==> Oxygen:Dissolved:Rinko

- Units: ==> umol/kg

- Format: F10.6 ==> F6.1

- Turbidity:Seapoint ==> Turbidity:RBR:Coda

CLEAN was run to reset the Maximum and Minimum values in the Header.

- Header Check and Standards Check were run and no problems were found.

Figure 1 – location of casts – basemap resolution does not cover the extent of the inlets.

Figure 2 – Zero order holds check

Figure 3 – Zero order holds removed