Study of Transmissivity Calculation

The following table explains the different variables used in calculations of transmissivity.

| Measurements | In equation from SeaBird | Factory |
| --- | --- | --- |
| Transmissivity – volts |  | Tr |
| Transmissivity - %/m | Light Transmission |  |
| Voltage in Air – Factory | A0 | Vair |
| Voltage in Air – Lab (current) | A1 | n/a |
| Blank Voltage – Factory | Y0 | Vd |
| Blank Voltage – Lab (current) | Y1 | n/a |
| Voltage in pure water – factory | W0 | Vref |
| Transmission % in pure water relative to air | 100% | n/a |
| Transmission % in pure water relative to water | 90.2% * - for red wavelength | n/a |
| Slope | M | n/a |
| Offset | B | n/a |

The values that go into a SeaBird configuration file are M and B.

Up until about 2000 the slope and offset were calculated as follows:

M=20*(A0-Y0)/(A1-Y1)

B=-M*Y1

From 2000 to 2011 calculations were either done by Doug Anderson, or after his death, were done following his method by Germaine Gatien. The method changed soon after Doug took over. His method was as follows:

M=(Tw/W0)*(A0-Y0)/(A1-Y1)

B=-M*Y1

Doug had a list of values for Tw and W0 for each instrument. He always used 90.2% for Tw.

From SeaBird Application Note No. 91 (2011) the calculation should be done as follows:

M=(Tw/(W0-Y0))*(A0-Y0)/(A1-Y1)

B=-M*Y1

So Doug’s used W0 where we are now told to use W0-Y0. Perhaps this is a change in advice, or perhaps he just made an error in setting up the spreadsheet.

| TR s/n | W0-Y0 | W0 | Difference |
| --- | --- | --- | --- |
| 333 | 4.791 | 4.753 | 0.038 |
| 498 | 4.829 | 4.765 | 0.064 |
| 732 | 4.700 | 4.637 | 0.063 |
| 953 | 4.768 | 4.719 | 0.049 |
| 983 | 4.727 | 4.674 | 0.053 |
| 1005 | 4.707 | 4.674 | 0.033 |
| 1185 | 4.742 | 4.719 | 0.023 |

For sensor #1185 the following are the initial calculations of M and B using the 3 methods.

Old Method: 20.0000 / -1.2400

Doug’s Method: 19.1142 / 1.1851

SeaBird Method: 19.0215 / 1.1793

And the factory and lab values for initial calibration and 2 years later were:

| Values from factory and lab tests |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| 1185 | A0 | A1 | Y0 | Y1 | W0 |
| Oct-08 | 4.804 | 4.804 | 0.062 | 0.062 | 4.719 |
| Aug-10 | 4.804 | 4.769 | 0.062 | 0.057 | 4.719 |

The resulting calculations of M and B using these values and 3 methods are:

| Slope/Offset Results for Different Methods |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Old Method |  | DA method |  | SeaBird method |  |
| M | B | M | B | M | B |
| 20.0000 | -1.2400 | 19.1142 | -1.1851 | 19.3687 | -1.2009 |
| 20.1273 | -1.1473 | 19.2359 | -1.0965 | 19.4920 | -1.1110 |

And to see how big a difference there is in calculated transmissivity in %/m. the following are the results for the August 2010 calibration and the differences between the SeaBird results and the two other methods for a range of input voltages:

| Voltage Measured | Old Method | Doug A's Method | SB Method | SeaBird - Doug A.'s method | SeaBird-Old Method |
| --- | --- | --- | --- | --- | --- |
| 1 | 18.9801 | 18.1395 | 18.3810 | +0.2415 | -0.5991 |
| 2 | 39.1074 | 37.3754 | 37.8730 | +0.4976 | -1.2344 |
| 3 | 59.2347 | 56.6113 | 57.3650 | +0.7537 | -1.8697 |
| 4 | 79.3621 | 75.8472 | 76.8570 | +1.0098 | -2.5051 |

If we now start using the SeaBird method then values will be somewhat higher, but we can make a note in the headers that all transmissivity conversions done for cruises from January 2011 onwards will use the method outlined in SeaBird Application Note 91.

Germaine Gatien

June 20, 2012