# 2014 spring bloom at Egmont: chlorophyll fluorometer deployment/recovery report

June 2014

Stephanie King, Sea This Consulting

king@seathisconsulting.com

# Background

A chlorophyll fluorometer has been deployed annually since 2010 at the town of Egmont (49.5N, 123.93W; Figure 1) for the purpose of monitoring potential seeding of the spring bloom in the Strait of Georgia from inlets.  The fluorometer was hung from the private dock at the Bathgate Resort and Marina.

Figure 1. Location of the Egmont fluorometer deployment 2014.  The fluorometer is hung from the Bathgate Resort and Marina’s dock at the first float house.

# Sensor and settings

The Wetlabs ECO Fluorometer FLNTUSB-2236 used in this deployment measures chlorophyll fluorescence emission at 695nm, turbidity from scattered light at 700nm and temperature. A copper wiper covers the sampling window between groups of measurements to reduce fouling, and the sensor was housed in an aluminum case painted with an anti-fouling paint (Figure 2).

For this deployment, the sensor was programmed to make a group of 5 measurements, each about 1-2 seconds apart, every 30 minutes.  The sensor has internal data storage and was powered by six 9-volt batteries.

Figure 2. Sensor and protective aluminum case.  The lower half of the aluminum case which is submerged on deployment is painted in black  antifouling paint.

# Deployment

FLNTUSB-2236 was deployed on January 29, 2014 by Jim and Ann Gower.  The senor’s sampling window was located at a depth of about 1m.

# Recovery

FLNTUSB-2236 was recovered on April 23, 2014 by Jim and Ann Gower.

# Data output

The raw output downloaded from the instrument as a text file, but the text file had unexpected characters in it.  This has been a problem before, and as before, the instrument was sent to Wetlabs for data recovery and servicing.

The raw output recovered from Wetlabs was processed with a program written in IDL to output the average and standard deviation from each group of measurements.  Measurements in air from prior to deployment and after recovery have been removed and scaling factors in Table 1 applied using the following equations:

Chlorophyll = Scale Factor*(Output-Dark Counts)

Turbidity = Scale Factor*(Output-Dark Counts)

Temperature = (output x slope) + intercept

Table 1. Calibration factors provided for FTNTUSB-2236 from Wetlabs calibration on Apr. 26, 2011.

| Chlorophyll |  | Turbidity |  | Thermistor |  |
| --- | --- | --- | --- | --- | --- |
| Dark counts | Scaling factor | Dark counts | Scaling factor | slope | intercept |
| 49 | 0.0182 | 50 | 0.0485 | -0.0056 | 70.8083 |

The following files have been provided for the IOS data archive:

-	An excel file containing the scaled, unprocessed (i.e. not averaged) data

-	A text file with the metadata

-	This deployment summary

# Additional comments and/or issues

The time series for chlorophyll fluorescence, turbidity and temperature are shown in Figure 3.   Fouling does not appear to be a problem.

The 2014 spring bloom at Egmont begins on March 2, although with relatively low concentrations of around 10 mg/m3 at its peak on March 9.  The bloom at Egmont was several weeks earlier and less intense than the chlorophyll fluorescence observed at the Halibut Bank buoy.

Figure 3. Averaged time series for chlorophyll fluorescence, turbidity and temperature. The standard deviation (red series) is for each burst of 5 measurements.