# CDAQ data from W.E. Ricker

31 March 2000

## Wind Speed and Direction:

Manufacturer:	R.M. Young Company

Model:	05106

Serial number	WM23045

Calibration date:	23 Jan 1996

Wind speed:

Range	0 to 60 m/s (130 mph), gust survival 100 m/s (220 mph)

Sensor:	18 cm diameter 4-blade helicoid propeller

Threshold sensitivity:	1.1 m/s (2.5 mph)

Wind Direction:

Range:	360° mechanical, 355° electrical (5° open)

Sensor:	balanced vane, 38 cm turning radius

Damping ratio:	.25

Threshold sensitivity	1.3 m/s (2.9 mph) at 10° displacement

## Barometer

Model:	SBF270

Serial number	640327

Calibration date	18 October 1995

Pressure range	800 to 100 millibars

Accuracy:	+/- .2 mbars

Zero pressure output	0 mbars, internally adjustable

Power requirements:	9.6 to 16 VDC

Output	0 to 5 volts

## Radiation Sensor:

Model:	Li-COR quantum sensor, LI-190SZ

Serial number:	Q20852

Calibration date:	12 July 1995

Resistance:	604 ohms

Output:	5.35 microamps per 1000 umol s-1 m-2

Maximum output	5.0 mv/1000umol s-1 m-2

Sensitivity:	8 uA per 1000 umol s-1 m-2

Detector:	high stability silicon photovoltaic detector (blue enhanced).

## Relative Humidity / Temperature Probe:

Sensor type:	1000ohm platinum RTD

Humidity:	capacitive polymer

Accuracy:	.3°C at 0°C (standard)

3% from 0 to 100%RH

output signal:	0 to 1 VDC

## Thermosalinograph (Sea temperature / conductivity Sensor):

Model:	Seabird seacat

Serial number	2488

temperature:

Range:	-5 to 35°C

accuracy:	.01°C

resolution:	.001°C

conductivity:

range:	0 to 7 S/m

accuracy:	.001 S/m

resolution:	.0001 S/m

Calibrations:

For data prior to 20 April 1999

temperature:	conductivity:

g	4.16616136e-03	-4.09360610e00

h	5.96058571e-04	4.89478043e-01

i	3.63943645e-06	7.77660256e-04

j	-1.97344952e-06	-8.36424908e-06

for data from 28 April 1999 to 28 January 2000

temperature:	conductivity:

g	4.16597834e-03	-4.08510707e00

h	5.95571522e-04	4.88671381e-01

i	3.24339437e-06	7.06323936e-04

J	-2.07705489e-06	-4.10309868e-06

For data after 9 February 2000

Temperature:	conductivity:

g	4.16687828e-03	-4.07884131e00

h	5.97614889e-04	4.87202706e-01

i	4.71276877e-06	1.00393708e-03

j	-1.74864349e-06	-1.90889973e-05