StnP_README.DOC

A mooring was put at station P (50 N, 145 W) in support of JGOFS.

Data were obtained from September 1995 to October 1998.  Usually the mooring was picked up and serviced in a day or two but there are two gaps.  One was from 26 April to 14 June 1997 because the mooring could not be recovered due to bad weather in February 1997.  The other was from 4 to 12 September 1997 when there was some delay between recovery and deployment.  The last deployment was in February 1998.  Data were collected until 30 October 1998 when the S4 memory filled up.  This mooring was to be serviced in September 1998 but the release failed.  It was finally recovered in September 1999.

The S4 was at about 30 m for summer deployments and at about 50 m for winter deployments. Depths for the various periods need to be checked.

Three 1 minute averages of velocity in row were taken once per hour.  After each velocity average spot readings of conductivity and temperature were recorded.  Having three readings allows for some check for bad readings.  A time line was recorded after every 18 readings (6 hours).  The S4’s time lines are given at the end of a sampling interval; the scalars are at these times but the velocities being an average over 1 minute represent an average centered at ½ minute earlier.

In the list that follows the last two digits of the serial number are used.  The full serial numbers are in the headers and the .CFG files to be used with APPNEW to set the velocity range (350 cm/s in this case), the resolutions (HI or 14 bit for C and T) and provide calibration corrections to C and T.

STN P DATA FILES

Data file       Inst. #	Dates						Depth		Clock at end

STP01.S4B	  45	3 September 95 to 29 February 96		  50 m		185 sec. slow

STP02.S4B	  46	29 February 96 to 20 May 96			  30 m		110 sec. slow

STP03.S4B	  45	20 May 96 to 27 August 96			  30 m		110 sec. slow

STP04.S4B	  46	27 August 96 to 26 April 97			  50 m		340 sec slow

STP05.S4B	  45	14 June 97 to 4 September 97			  30 m		74 sec. slow

STP06.S4B	  46	12 September 97 to 26 February 98		  50 m		192 sec. slow

STP07.S4B	  45	26 February 98 to 30 October 98		  30 m		250 sec. slow

At the start the clocks were correct within about five seconds.  The clocks run slower in the field than the lab because the temperature is lower.  The clocks were read somewhat after recovery and the slowness at the end was estimated using the rate of loss in the lab (1 second in 21 days for #45 and 8 seconds in 45 days for #46).  The rate of time loss for each instrument is similar except for record 4 when 46 lost more time than would be expected from the other records.  The time check was done about 2 months after recovery.  The time lost during deployment is unlikely to be less than about 280 seconds.  The error in the average time correction would be about 30 seconds.  A 30 second error shifts the phase of a semi-diurnal tide by ¼ degree.

Velocity zeros were adjusted before each deployment to be zero within 1 bit (0.2 cm/s).  They were checked after each deployment and found to be zero within about 1 bit except for # 45 in April 1996 which had one component with a zero of 0.5 cm/s.  Compasses were checked before and after deployment.  Directions were good to 3 degrees or better.  Because the S4’s rotate freely about their mooring rod small velocity and direction errors will tend to zero when a few readings are averaged.

Calibration corrections are based on calibrations (see S4README.DOC page 2 for methods) at salinity’s of about 12 and 24 over a range of temperature from about 5 to 20 degrees in January and July 1995.  Temperature corrections were consistent to 0.01 for # 45 and 0.015 for # 46.  Salinity was consistent to 0.15 for both.  Checks were made between deployments.  Temperature was consistent to 0.01 for # 45.  Temperatures for # 46 were about 0.02 low in August 97 and May 98.  Thus T’s should be good to 0.03 or better for 45 and 0.04 or better for 46.  Salinity for 45 was 0 to 0.3 low; for 46 salinity was 0.2 low to 0.1 high.  In the data when changing from one instrument to the other salinity matches to 0.2 or better and temperature to 0.1 or better.

Generally the data quality seems good.  There seems to a bad temperature reading in each of records 4 and 6.  There are rapid T and C changes when the instrument is near the bottom of the mixed layer and there are probably some C spikes as well.  The mean velocities are small (< 5 cm/s) but there are the high frequency oscillations with amplitudes up to 40 cm/s or so that vary with time.  A few expanded plots show that there are inertial oscillations (period about 18 hours) and semi-diurnal tides.  There could be velocity spikes too but they cannot be distinguished from the high frequency variations on a full record plot.  Despiking and expanded plots need to be done before using these data.